- 11D-8.0035 Approval of Alcohol Reference Solution and Sources.

(1) The Department shall approve a source of alcohol reference solution for use by agencies in the State of Florida. The source approved by the Department shall be an entity that manufactures alcohol reference solutions and meets the following requirements:

(a) The source must prepare alcohol reference solution, and be capable of producing a minimum batch volume of 800 bottles, each containing at least 500 milliliters, to produce the following vapor alcohol concentrations: 0.05 g/210L, 0.08 g/210L, and 0.20g/210L;

(b) The source must have performed and documented tests that demonstrate that the alcohol reference solutions are reliable for at least two years from the date of manufacture.

(2) The Department shall approve each lot of alcohol reference solution prior to distribution for use in Florida.

(a) The Department shall determine the alcohol concentration in a minimum of ten (10) sample bottles of each lot of alcohol reference solution using gas chromatography or other scientifically accepted method. Duplicate analyses will be performed on each sample bottle of alcohol reference solution. All analysis results shall fall within the alcohol reference solution acceptable range.

(b) The Department shall notify the source that the approved lots may be distributed for use in Florida, and shall issue a Certificate of Assurance, FDLE/ATP Form 32 – Rev. March 2001.

(3) Alcohol reference solution lots approved by the Department shall be used in agency or Department inspections within two (2) years of the date of manufacture.

(4) Approval analyses of alcohol reference solution lots shall be based on requirements and procedures in effect at the time such lots are submitted for approval. No post-approval analysis is required for previously approved alcohol reference solution lots.

Rulemaking Authority 316.1932(1)(a)2., (f)1., 322.63(3)(a), 327.352(1)(b)3., (d) FS. Law Implemented 316.1932(1)(b)2., 316.1934(3), 322.63(3)(b), 327.352(1)(e), 327.354(3) FS. History–New 7-6-99, Amended 7-29-01, 12-9-04.