64E-5.624 Decay in Storage.

(1) A licensee shall hold radioactive material with a physical half life of less than 90 days for decay in storage before disposal as ordinary trash. A licensee is exempt from the requirements of Rule 64E-5.328, F.A.C., of these regulations if:

(a) The radioactive material is held for decay a minimum of 10 half-lives;

(b) The radioactive material is monitored at the container surface before disposal as ordinary trash and its radioactivity cannot be distinguished from the background radiation level in a low background radiation area with a radiation survey instrument set on its most sensitive scale and with no interposed shielding;

(c) All radiation labels are removed or obliterated; and

(d) Each generator column is separated and monitored individually with all radiation shielding removed to ensure that its contents have decayed to background radiation levels before disposal.

(2) The licensee shall retain a record of each disposal for 3 years. The record shall include:

(a) The date of the disposal;

(b) The date on which the radioactive material was placed in storage;

(c) The radionuclides disposed;

(d) The model and serial number of the radiation survey instrument used;

(e) The background dose rate;

(f) The radiation dose rate measured at the surface of each waste container; and

(g) The name of the individual who performed the disposal.

Specific Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.732.