# CHAPTER 20-65

# COLOR GRADING – PROCESSED ORANGE PRODUCTS

# 20-65.001	Standard Light Source

# 20-65.002	Equipment

# 20-65.003	Preparing Sample

# 20-65.004	Procedure

# 20-65.005	Calculation

## 20-65.001 Standard Light Source.

Where color grading of processed orange products is based on standard USDA color tubes, the grading shall be done under a standard light source of 7400º Kelvin, equivalent to that provided by the Macbeth Examolite color box, Model EBA-220.

Rulemaking Authority 601.10(7), 601.11 FS. Law Implemented 601.10(7), 601.11 FS. History–Formerly 105-1.19(1)(q), Amended 1-1-75, Formerly 20-65.01, Amended 1-30-86.

## 20-65.002 Equipment.

(1) Any of the following colorimeter or spectrophotometer devices shall be used for measuring color scores for frozen concentrated orange juice (FCOJ), canned concentrated orange juice (CCOJ), concentrated orange juice for manufacturing (COJFM), pasteurized orange juice (POJ), canned single strength orange juice (CSSOJ), and orange juice from concentrate (OJFC):

(a) HunterLab Model D45 Citrus Colorimeter,

(b) HunterLab Model D45D2 Citrus Colorimeter,

(c) Macbeth Color-Eye Spectrophotometer Model LS-1500,

(d) Macbeth Color-Eye Spectrophotometer Model 2020+ with Optiview Software Package,

(e) Macbeth Color-Eye Spectrophotometer Model 3000 with Optiview Software Package,

(f) Macbeth Color-Eye spectrophotometer Model 3100 with Optiview Software Package,

(g) Macbeth Color-Eye Spectrophotometer Model i5 with Color iQC Basic Software Package.

(h) Xrite Color-Eye Spectrophotometer Model i5 with Color IQC Professional SLITaper.

(i) Xrite Spectrophotometer Model Ci7600 with Color IQC Professional SLITaper.

(j) Xrite Spectrophotometer Model Ci7520 with Color IQC Professional SLITaper.

(k) Devices validated and approved by order of the Florida Department of Agriculture and Consumer Services.

(2) Each of the colorimeter or spectrophotometer devices shall be an equivalent method to the standard color tubes designated USDA OJ 2 through 6 for measuring color scores for canned orange juice (COJ).

(3) The colorimeter or spectrophotometer used shall be set, installed and maintained in a manner acceptable to and approved by the Processed Products Branch, Fruit and Vegetable Division, Agricultural Marketing Service, United States Department of Agriculture, which shall also be responsible for maintaining proper settings.

(4) A vacuum pump or aspirator for deaeration of the juice shall be provided for use with the colorimeter or spectrophotometer.

Rulemaking Authority 601.10(7) FS. Law Implemented 601.11, 601.24 FS. History–Formerly 105-1.19(1)(r), (l), Amended 1-1-75, Formerly 20-65.02, Amended 1-30-86, 1-12-92, 6-21-93, 1-17-95 10-26-97, 8-22-06, 6-19-12, 1-24-19, 11-26-20.

## 20-65.003 Preparing Sample.

Sampling for color determination of all products shall be at the appropriate Brix of the particular product being sampled, when reconstituted to single strength.

(1) In the case of product concentrated at 44.8 degrees Brix, such product shall be reconstituted and sampled at 12.8 degrees. Product concentrated at 43.2 degrees Brix shall be reconstituted and sampled at 12.3 degrees Brix. Product concentrated at 41.8 degrees Brix shall be reconstituted and sampled at 11.8 degrees Brix. If any frozen concentrated orange juice for manufacturing sampled for color determination prior to March 1, 1982, is required to be resampled subsequent thereto, the sample shall be prepared in accordance with this section, however the product shall be reconstituted at 12.3 degrees Brix.

(2) Concentrated orange juice for manufacturing shall be reconstituted and sampled at 11.8 degrees Brix.

(3) Brix shall be corrected for anhydrous citric acid and temperature of the refractometer.

(4) Reconstituted samples shall be deaerated for not less than 3 minutes and the temperature adjusted to 80 degrees F., plus or minus 2 degrees F.

Rulemaking Authority 601.10(7), 601.11, 601.24 FS. Law Implemented 601.10(7), 601.11, 601.24 FS. History–Formerly 105-1.19(1)(r), (2)(a), Amended 1-1-75, 3-20-82, Formerly 20-65.03, Amended 1-30-86.

## 20-65.004 Procedure.

(1) The colorimeter or spectrophotometer is to be kept with power connected and maintained on standby or in accordance with manufacturer’s instructions. Turn the instrument to “on” position for ten minutes before making measurements, or follow directions of manufacturer. All light sources shall conform to CIE (1931) illuminant C. All OJ4 color tube calibration values are to be obtained from the Winter Haven, FL laboratory of the Processed Products Branch, Fruit and Vegetable Division, Agricultural Marketing Service, United States Department of Agriculture.

(2) For the HunterLab Model D45 Citrus Colorimeter, standardize Citrus Red (CR) and Citrus Yellow (CY) with OJ4 color tube indexed for the instrument. Insert sample of juice at indexed position for the sample tube and read CR and CY and record.

(3) For the HunterLab Model D45D2 Citrus Colorimeter, standardize Y, CR and CY with OJ4 color tube indexed for the instrument. Insert sample of juice at indexed position for the sample tube and read CR, CY and Color Number (CN) and record.

(4) For the Macbeth Color-Eye Spectrophotometer Model 1500 or Model 2020+ with Optiview Software Package, equipped with the 19 mm aperture tube adapter and operated with small area view (SAV) setting with UV and gloss included (CRIIS), standardize X, Y, and Z with white tile with SAV plate in place. Replace SAV plate with tube adapter and insert sample of juice at indexed position for the sample tube and read color number values as shown on instrument, then record.

(5) For the Macbeth Color-Eye Spectrophotometer Model 3000 or 3100 with Optiview Software Package, equipped with the 19 mm aperture tube adapter and operated with small area view (SAV) setting with UV and gloss excluded (DREOS), standardize X, Y, and Z with white tile with SAV plate in place. Replace SAV plate with tube adapter and insert sample of juice at indexed position for the sample tube and read color number values as shown on instrument, then record.

(6) For the GretagMacbeth Color i5 Spectrophotometer with Color IQ Basic Software and 19mm Orange Juice Test Tube Holder. UV and Gloss exclusion is processed by software. Calibration is prompted by software. When completed replace Orange Juice Holder. Push F6 and Color Number is reported. For confirmation, lights illuminated on front of Color i5 should be: LAV, Reflection, D65, Excluded and Calibrated. Calibration required every 24 hours.

(7) To determine color number or equivalent color score, use the figures in subsections (2)-(5) to calculate the color number or equivalent color score as set forth in Rule 20-65.005, F.A.C.

Rulemaking Authority 601.10(7), 601.11, 601.24 FS. Law Implemented 601.10(7), 601.11, 601.24 FS. History–Formerly 105-1.19(1)(r), (2)(b), Amended 1-1-75, Formerly 20-65.04, Amended 1-30-86, 1-12-92, 6-21-93, 1-17-95, 10-26-97, 8-22-06.

## 20-65.005 Calculation.

(1) Step 1 – The color number is calculated as follows:

(a) For the HunterLab Model D45 Citrus Colorimeter the color number is equal to 22.510 plus 0.165CR plus 0.111CY.

(b) For the HunterLab Model D45D2 Citrus Colorimeter the color number is read directly from the instrument and recorded.

(c) For the Macbeth Color-Eye Spectrophotometer Model LS-1500, the color number is equal to 17 multiplied by the sum of 3.4 times X divided by Y minus Z divided by Y plus 5 divided by Y, and subtract a constant factor 12.6 from the resulting total.

CN = 17(3.4X/Y – Z/Y + 5/Y) – 12.6.

(d) For the Macbeth Color-Eye Spectrophotometer Model 2020+ with Optiview Software Package, the color number is read directly from the instrument and recorded.

(e) For the Macbeth Color-Eye Spectrophotometer Model 3000 and 3100 with Optiview Software Package, the color number is read directly from the instrument and recorded; or it may be calculated from the measured tristimulus X, Y and Z values by applying the equation CN = 14.5(3.15X/Y – Z/Y + 5.2/Y) – 2.6.

(f) For the GretagMacbeth Color-Eye Spectrophotometer Model i5 with Color iQC Basic Software, the color number is read directly from the instrument and recorded; or it may be calculated from the measured tristimulus X, Y and Z values by applying the equation CN = 14.5(3.15X/Y - Z/Y + 4.1/Y) -2.6.

(2) Step 2 – Apply the color number to the following table to determine the color score or equivalent color score for each appropriate product:

|  | Color | Color Score | Adjustment | Color |
| --- | --- | --- | --- | --- |
|  | Score for | for FCOJ & | for COJ & | Score for |
| Color Number | POJ & OJFC | CCOJ | COJFM | COJ & COJFM |
| 32.5 – 33.4 | 33 | 33 | +2 | 35 |
| 33.5 – 34.4 | 34 | 34 | +2 | 36 |
| 34.5 – 35.4 | — | 35 | +1 | 36 |
| 34.5 – 34.9 | 35 | — | — | — |
| 35.0 – 36.4 | 36 | — | — | — |
| 35.5 – 36.4 | — | 36 | +1 | 37 |
| 36.5 – 37.4 | 37 | 37 | +1 | 38 |
| 37.5 – 38.4 | 38 | 38 | 0 | 38 |
| 38.5 – 39.4 | 39 | 39 | 0 | 39 |
| 39.5 – 40.4 | 40 | 40 | 0 | 40 |

Rulemaking Authority 601.10(7), 601.11, 601.24 FS. Law Implemented 601.10(7), 601.11, 601.24 FS. History–Formerly 105-1.19(1)(r), (2)(c), Amended 1-1-75, 4-15-76, Formerly 20-65.05, Amended 1-30-86, 1-12-92, 6-21-93, 1-17-95, 10-26-97, 8-22-06.