- CHAPTER 64E-5

- CONTROL OF RADIATION HAZARDS

- 64E-5.1001 	Standards

- 64E-5.1002 	Monitoring Requirements

- 64E-5.1003 	Monitoring Fees

- 64E-5.101 	Definitions

- 64E-5.102 	Exemptions

- 64E-5.103 	Records

- 64E-5.104 	Tests

- 64E-5.105 	Prohibited Uses

- 64E-5.106 	Units of Exposure and Dose

- 64E-5.1101 	Prohibitions

- 64E-5.1102 	Storage and Transportation Precautions

- 64E-5.1103 	Radiation Survey Instruments

- 64E-5.1104 	Leak Testing of Sealed Sources

- 64E-5.1105 	Quarterly Inventory

- 64E-5.1106 	Utilization Records

- 64E-5.1107 	Design, Performance and Certification Criteria for Sealed Sources Used in Downhole Operations

- 64E-5.11071 	Uranium Sinker Bars

- 64E-5.11072 	Energy Compensation Source

- 64E-5.11073 	Tritium Neutron Generator Target Source

- 64E-5.1108 	Labeling

- 64E-5.1109 	Inspection and Maintenance

- 64E-5.1110 	Training Requirements

- 64E-5.1111 	Operating and Emergency Procedures

- 64E-5.1112 	Personnel Monitoring

- 64E-5.1113 	Security

- 64E-5.1114 	Handling Tools

- 64E-5.1115 	Subsurface Tracer Studies

- 64E-5.1116 	Radiation Surveys

- 64E-5.1117 	Documents and Records Required at Field Stations

- 64E-5.1118 	Temporary Jobsites

- 64E-5.1119 	Notification of Incidents, Abandonment and Lost Sources

- 64E-5.1120 	Subjects to Be Included in Training Courses for Logging Supervisors

- 64E-5.1201 	Radon Requirements and Enforcement

- 64E-5.1202 	Definitions

- 64E-5.1203 	General Provisions

- 64E-5.1204 	Certification Requirements for Radon Measurement Specialists and Technicians

- 64E-5.1205 	Certification Requirements for Radon Mitigation Specialists and Technicians

- 64E-5.1206 	Certification Requirements for Radon Measurement Businesses

- 64E-5.1207 	Certification Requirements for Radon Mitigation Businesses

- 64E-5.1208 	Measurement Requirements and Procedures

- 64E-5.1301 	Sealed or Unsealed Sources of Radioactive Material

- 64E-5.1302 	Operating and Emergency Procedures

- 64E-5.1303 	Leak Test Requirements for Possession of Sealed Sources

- 64E-5.1304 	Inventory Requirements

- 64E-5.1305 	Training Requirements, Authority, Duties and Responsibilities of the Radiation Safety Officer

- 64E-5.1306 	Opening Sealed Sources

- 64E-5.1307 	Training Requirements for Authorized Users

- 64E-5.1308 	Additional Requirements for General Licenses

- 64E-5.1309 	Training for Current Authorized Users

- 64E-5.1310 	Personnel Monitoring

- 64E-5.1311 	Storage, Security and Transportation Precautions

- 64E-5.1312 	Training and User Requirements for the Possession and Use of Sealed Sources in Portable Devices

- 64E-5.1313 	Training and User Requirements for the Possession and Use of Sealed Sources in Fixed Devices

- 64E-5.1314 	Possession of Survey Instruments

- 64E-5.1315 	Additional Requirements

- 64E-5.1316 	General Rules for the Safe Use of Unsealed Sources of Radioactive Materials

- 64E-5.1317 	Storage and Control of Volatiles and Gases

- 64E-5.1318 	Instrumentation

- 64E-5.1319 	Contamination Control Program

- 64E-5.1320 	Bioassy Program

- 64E-5.1401 	Irradiators

- 64E-5.1402 	Definitions

- 64E-5.1403 	Specific License for Large Irradiators

- 64E-5.1404 	Start of Construction

- 64E-5.1405 	Performance Criteria for Sealed Sources

- 64E-5.1406 	Access Control

- 64E-5.1407 	Shielding

- 64E-5.1408 	Fire Protection

- 64E-5.1409 	Radiation Monitors

- 64E-5.1410 	Control of Source Movement

- 64E-5.1411 	Irradiator Pools

- 64E-5.1412 	Source Rack Protection

- 64E-5.1413 	Power Failures

- 64E-5.1414 	Design Requirements

- 64E-5.1415 	Construction Control

- 64E-5.1416 	Training

- 64E-5.1417 	Operating and Emergency Procedures

- 64E-5.1418 	Personnel Monitoring

- 64E-5.1419 	Radiation Surveys

- 64E-5.1420 	Detection of Leaking or Contaminated Sources

- 64E-5.1421 	Inspection and Maintenance

- 64E-5.1422 	Pool Water Purity

- 64E-5.1423 	Attendance During Operation

- 64E-5.1424 	Entering and Leaving the Radiation Room

- 64E-5.1425 	Irradiation of Explosive or Highly Flammable Materials

- 64E-5.1426 	Records and Retention Periods

- 64E-5.1427 	Reports and Notifications

- 64E-5.1501 	Purpose and Scope

- 64E-5.1502 	Transportation of Radioactive Material

- 64E-5.1503 	Exemptions

- 64E-5.1504 	General Licenses for Carriers

- 64E-5.1505 	Routine Determinations

- 64E-5.1506 	Advance Notification of Shipment of Certain Quantities of Radioactive Waste

- 64E-5.1507 	Designation of Routes for Shipment of Radioactive Waste Requiring Advance Notification

- 64E-5.1508 	Inspection of Low-Level Radioactive Waste Shipments

- 64E-5.1509 	Permit Requirements

- 64E-5.1510 	Air Transport of Plutonium

- 64E-5.1511 	Notification in the Event of Suspected or Real Breach of Containment

- 64E-5.1512 	Inspections

- 64E-5.1513 	Communications

- 64E-5.1601 	Definitions

- 64E-5.1602 	Administrative Requirements

- 64E-5.1603 	Training and Education

- 64E-5.1604 	General Technical Requirements for Electronic Brachytherapy Facilities

- 64E-5.201 	Licensing of Radioactive Material

- 64E-5.202 	Source Material ‒ Exemptions

- 64E-5.203 	Radioactive Material Other Than Source Material ‒ Exemptions

- 64E-5.204 	Types of Licenses

- 64E-5.205 	General Licenses ‒ Source Material

- 64E-5.206 	General Licenses ‒ Radioactive Material Other Than Source Material

- 64E-5.207 	Filing Application for Specific Licenses

- 64E-5.208 	General Requirements for the Issuance of Specific Licenses

- 64E-5.209 	Special Requirements for Specific Licenses of Broad Scope

- 64E-5.210 	Special Requirements for a Specific License to Manufacture, Assemble, Repair or Distribute Commodities,

- Products or Devices Which Contain Radioactive Material

- 64E-5.211 	Special Requirements for Issuance of Specific Licenses for Source Material Milling

- 64E-5.212 	Issuance of Specific Licenses

- 64E-5.213 	Specific Terms and Conditions of License

- 64E-5.214 	Expiration and Termination of Licenses and Decommissioning of Sites and Separate Buildings or Outdoor Areas

- 64E-5.215 	Transfer of Material

- 64E-5.216 	Reciprocal Recognition of Licenses for Byproduct, Source, Naturally Occurring and Accelerator Produced

- Radioactive Material, and Special Nuclear Material in Quantities Not Sufficient to Form a Critical Mass

- 64E-5.217 	Bonding of Persons Licensed Pursuant to Subpart II C

- 64E-5.218 	Performance of Inspections

- 64E-5.219 	Emergency Planning

- 64E-5.220 	Radioactive Quantities

- 64E-5.221 	Radiological Criteria for License Termination

- 64E-5.222 	Radiological Criteria for Unrestricted Use

- 64E-5.223 	Criteria for License Termination Under Restricted Conditions

- 64E-5.224 	Alternate Criteria for License Termination

- 64E-5.225 	Public Notification and Public Participation

- 64E-5.226 	Minimizing Contamination

- 64E-5.301 	Standards for Protection Against Radiation

- 64E-5.302 	Implementation

- 64E-5.303 	Radiation Protection Programs

- 64E-5.304 	Occupational Dose Limits for Adults

- 64E-5.305 	Compliance with Requirements for Summation of External and Internal Doses

- 64E-5.306 	Determination of External Dose from Airborne Radioactive Material

- 64E-5.307 	Determination of Internal Exposure

- 64E-5.308 	Determination of Prior Occupational Dose

- 64E-5.309 	Planned Special Exposures

- 64E-5.310 	Occupation Dose Limits for Minors

- 64E-5.311 	Dose to an Embryo or Fetus

- 64E-5.312 	Dose Limits for Individual Members of the Public

- 64E-5.313 	Compliance with Dose Limits for Individual Members of the Public

- 64E-5.314 	General

- 64E-5.315 	Conditions Requiring Individual Monitoring of External and Internal Occupational Dose

- 64E-5.316 	Control of Access to High Radiation Areas

- 64E-5.317 	Control of Access to Very High Radiation Areas

- 64E-5.318 	Use of Process or Other Engineering Controls

- 64E-5.319 	Use of Individual Respiratory Protection Equipment

- 64E-5.320 	Security of Stored Sources of Radiation

- 64E-5.321 	Control of Sources of Radiation Not in Storage

- 64E-5.322 	Caution Signs

- 64E-5.323 	Posting Requirements

- 64E-5.324 	Exceptions to Posting Requirements

- 64E-5.325 	Labeling Containers and Radiation Machines

- 64E-5.326 	Exemptions to Labeling Requirements

- 64E-5.327 	Procedures for Receiving and Opening Packages

- 64E-5.328 	General Requirements

- 64E-5.329 	Method of Obtaining Approval of Proposed Disposal Procedures

- 64E-5.330 	Discharge by Release into Sanitary Sewerage

- 64E-5.331 	Disposal of Specific Wastes

- 64E-5.332 	Transfer for Disposal and Manifests

- 64E-5.333 	Classification and Characteristics of Low-Level Radioactive Waste for Near-Surface Land Disposal, Labeling and

- Manifest Requirements

- 64E-5.334 	General Provisions

- 64E-5.335 	Records of Radiation Protection Programs

- 64E-5.336 	Records of Surveys

- 64E-5.337 	Records of Tests for Leakage or Contamination of Sealed Sources

- 64E-5.338 	Records of Planned Special Exposures

- 64E-5.339 	Records of Individual Monitoring Results

- 64E-5.340 	Records of Waste Disposal or Transfers

- 64E-5.341 	Records of Testing Entry Control Devices for Very High Radiation Areas

- 64E-5.342 	Form of Records

- 64E-5.343 	Reports of Stolen, Lost, or Missing Licensed or Registered Sources of Radiation

- 64E-5.344 	Notification of Incidents

- 64E-5.345 	Reports of Exposures, Radiation Levels, Concentrations of Radioactive Material Exceeding the Constraints or

- Limits, Medical Events and Dose to an Embryo/Fetus or a Nursing Child

- 64E-5.346 	Reports of Planned Special Exposures

- 64E-5.347 	Notifications and Reports to Individuals

- 64E-5.348 	Reports of Leaking or Contaminated Sealed Sources

- 64E-5.349 	Vacating Premises

- 64E-5.350 	Reports of Transactions Involving Nationally Tracked Sources

- 64E-5.351 	Nationally Tracked Source Thresholds

- 64E-5.423 	Definitions

- 64E-5.424 	Requirements for Industrial Radiography Equipment Using Sealed Sources

- 64E-5.425 	Locking of Sources of Radiation, Storage Precautions, and Surveillance

- 64E-5.426 	Radiation Survey Instruments

- 64E-5.427 	Leak Testing, Repairing, Tagging, Opening, Modifying, and Replacing Sealed Sources and Devices

- 64E-5.428 	Quarterly Inventory

- 64E-5.429 	Source Movement Logs, Daily Survey Reports, and Individual Dosimeter Logs

- 64E-5.430 	Inspection and Maintenance

- 64E-5.431 	Permanent Radiographic Installations

- 64E-5.432 	Radiation Protection Program

- 64E-5.433 	Radiation Safety Officer

- 64E-5.434 	Training, Testing, Certification, and Audits

- 64E-5.435 	Conducting Industrial Radiographic Operations

- 64E-5.436 	Operating and Emergency Procedures

- 64E-5.437 	Personnel Monitoring

- 64E-5.438 	Radiation Surveys

- 64E-5.439 	Posting

- 64E-5.440 	Records

- 64E-5.441 	Reporting Requirements

- 64E-5.501 	Definitions

- 64E-5.502 	General Requirements

- 64E-5.503 	General Requirements for All Diagnostic X-Ray Systems

- 64E-5.504 	Fluoroscopic X-Ray Systems

- 64E-5.505 	Diagnostic Radiography Systems, Other than Fluoroscopic, Mammographic, Dental Intraoral or Veterinary

- Systems

- 64E-5.506 	Intraoral Dental Radiographic Systems

- 64E-5.507 	Therapeutic X-Ray Systems of Less Than 1 MeV

- 64E-5.508 	X-Ray and Electron Therapy Systems with Energies of 1 MeV and Above

- 64E-5.509 	Veterinary Medicine X-Ray Operations

- 64E-5.510 	Mammographic Systems

- 64E-5.511 	Registration of Radiation Machines

- 64E-5.601 	License Required

- 64E-5.6011 	Definitions

- 64E-5.602 	License Amendments

- 64E-5.603 	Notification

- 64E-5.604 	ALARA Program

- 64E-5.605 	Radiation Safety Officer

- 64E-5.606 	Radiation Safety Committee

- 64E-5.607 	Authority and Responsibilities

- 64E-5.608 	Supervision

- 64E-5.609 	Visiting Authorized User, Visiting Authorized Medical Physicist, or Visiting RSO

- 64E-5.610 	Mobile Medical Service Requirements

- 64E-5.611 	Quality Management Program and Notifications, Records and Reports of Medical Events

- 64E-5.612 	Suppliers

- 64E-5.613 	Quality Control of Diagnostic Instrumentation

- 64E-5.614 	Possession, Use, Calibration, and Check of Dose Calibrators in the Use of Unsealed Radiopharmaceuticals

- 64E-5.615 	Use, Calibration and Check of Survey Instruments

- 64E-5.616 	Determination of Dosages of Unsealed Radioactive Material for Medical Use

- 64E-5.617 	Authorization for Calibration, Transmission and Reference Sources

- 64E-5.618 	Requirements for Possession of Sealed Sources and Brachytherapy Sources

- 64E-5.619 	Syringe Shields and Labels

- 64E-5.620 	Vial Shields and Labels

- 64E-5.621 	Surveys for Contamination and Ambient Radiation Dose Rate

- 64E-5.622 	Release of Patients or Human Research Subjects Treated with Radiopharmaceuticals, Implants or Remote

- Afterloader Units

- 64E-5.623 	Storage of Volatiles and Gases

- 64E-5.624 	Decay in Storage

- 64E-5.625 	Safety Instructions and Precautions for Liquid Iodine, Radiopharmaceutical Therapy, Manual Brachytherapy,

- Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic Radiosurgery

- 64E-5.6251 	Manual Brachytherapy Therapy Related Computer Systems

- 64E-5.626 	Use of Radiopharmaceuticals for Uptake, Dilution, or Excretion Studies

- 64E-5.627 	Use of Unsealed Radiopharmaceuticals, Generators, and Reagent Kits for Imaging and Localization Studies

- 64E-5.628 	Generators

- 64E-5.629 	Control of Aerosols and Gases

- 64E-5.630 	Use of Radiopharmaceuticals for Therapy

- 64E-5.631 	Use of Sealed Sources for Diagnosis

- 64E-5.632 	Use of Sources for Manual Brachytherapy

- 64E-5.633 	Manual Brachytherapy Sources Inventory and Surveys

- 64E-5.6331 	Calibration Measurements of Manual Brachytherapy Sources

- 64E-5.6332 	Decay of Strontium-90 Sources for Ophthalmic Treatments

- 64E-5.634 	Use of a Sealed Source in a Remote Afterloader Unit, Teletherapy Unit, or Gamma Sterotactic Radiosurgery Unit

- 64E-5.635 	Installation, Adjustment, Maintenance and Repair Restrictions

- 64E-5.636 	Safety Procedures and Instructions for Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic

- Radiosurgery Units

- 64E-5.637 	Safety Precautions for Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic Radiosurgery Units

- 64E-5.638 	Radiation Monitoring Devices

- 64E-5.639 	Viewing Systems

- 64E-5.640 	Dosimetry Equipment Used With Remote Afterloading Units, Teletherapy Units, or Gamma Stereotactic

- Radiosurgery Units

- 64E-5.641 	Full Calibration Measurements on Teletherapy Units

- 64E-5.6411 	Full Calibration Measurements on Remote Afterloader Units

- 64E-5.6412 	Full Calibration Measurements on Gamma Stereotactic Radiosurgery Units

- 64E-5.642 	Periodic Spot-Checks or Teletherapy Units

- 64E-5.6421 	Periodic Spot-Checks for Remote Afterloader Units

- 64E-5.6422 	Periodic Spot-Checks for Gamma Stereotactic Radiosurgery Units

- 64E-5.6423 	Additional Technical Requirements for Mobile Remote Afterloader Units

- 64E-5.643 	Radiation Surveys for Teletherapy Facilities

- 64E-5.644 	Radiation Surveys for Remote Afterloader and Gamma Stereotactic Radiosurgery Facilities

- 64E-5.645 	Remote Afterloader, Gamma Stereotactic, and Teletherapy Therapy-Related Computer Systems

- 64E-5.646 	Reports of Teletherapy Surveys, Checks, Tests, and Measurements

- 64E-5.647 	Five Year Inspection for Teletherapy and Gamma Stereotactic Radiosurgery Units

- 64E-5.648 	Training for a Radiation Safety Officer

- 64E-5.649 	Training for Uptake, Dilution, or Excretion Studies

- 64E-5.650 	Training for Imaging and Localization Studies for Which a Written Directive Is Not Required

- 64E-5.651 	Training for Therapeutic Use of Radiopharmaceuticals (Repealed)

- 64E-5.652 	Training for Use of Manual Brachytherapy Sources

- 64E-5.653 	Training for Ophthalmic Use of Strontium 90

- 64E-5.654 	Training for Use of Sealed Sources for Diagnosis

- 64E-5.655 	Training for Use of Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic Radiosurgery Units

- 64E-5.656 	Training for and Authorized Medical Physicist

- 64E-5.657 	Training for Experienced RSO, Teletherapy or Medical Physicist, Authorized Medical Physicist, Authorized User,

- Nuclear Pharmacist, and Authorized Nuclear Pharmacist

- 64E-5.658 	Recentness of Training

- 64E-5.659 	Training for an Authorized Nuclear Pharmacist

- 64E-5.660 	Training for Use of Unsealed Radioactive Material for Which A Written Directive is Required in Rule 63E-5.626,

- 64E-5.627 or 64E-5.630, F.A.C.

- 64E-5.661 	Training for the Oral Administration of Sodium Iodide I-131 Requiring A Written Directive in Quantities Less

- Than or Equal to 1.22 Gigabecquerels (33 Millicuries)

- 64E-5.662 	Training for the Oral Administration of Sodium Iodide I-131 Requiring A Written Directive in Quantities Greater

- Than 1.22 Gigabecquerels (33 Millicuries)

- 64E-5.663 	Training for the Parenteral Administration of Unsealed Radioactive Material Requiring a Written Directive

- 64E-5.664 	Other Medical Uses of Radioactive Material or Radiation From Radioactive Material

- 64E-5.701 	Equipment Requirements

- 64E-5.702 	Area Requirements

- 64E-5.703 	Operating Requirements

- 64E-5.704 	Personnel Requirements

- 64E-5.801 	Registration Requirements

- 64E-5.802 	General Requirements for the Issuance of a Registration Certificate for Particle Accelerators

- 64E-5.803 	Particle Accelerators for Therapeutic Use on Humans

- 64E-5.804 	Limitations

- 64E-5.805 	Shielding and Safety Design Requirements

- 64E-5.806 	Particle Accelerator Controls and Interlock Systems

- 64E-5.807 	Warning Devices

- 64E-5.808 	Operating Procedures

- 64E-5.809 	Radiation Monitoring Requirements

- 64E-5.810 	Ventilation Systems

- 64E-5.901 	Posting of Notices to Workers

- 64E-5.902 	Instructions to Workers

- 64E-5.903 	Notification and Reports to Individuals

- 64E-5.904 	Presence of Representatives of Licensees or Registrants and Workers During Inspection

- 64E-5.905 	Consultation with Workers During Inspections

- 64E-5.906 	Request by Workers for Inspections

- 64E-5.907 	Inspections Not Warranted; Informal Review

- 64E-5.1001 Standards.

(1) Radiation exposure to the public from naturally occurring radioactive materials shall be maintained as low as reasonably achievable. For the purposes of this part, the normal background level of gamma radiation in buildings is 6 microroentgens (1.55 nC/kg) per hour and the normal background radon decay product concentration in buildings is 0.004 Working Level (WL).

(2) The mean gamma exposure rate in a building shall not exceed 20 microroentgens (5.16 nC/kg) per hour, including background, and the annual average radon decay product concentration shall not exceed 0.02 WL, including background.

Rulemaking Authority 404.051, 404.056 FS. Law Implemented 404.022(2), 404.051(4), 404.056 FS. History–New 2-16-86, Amended 1-3-89, Formerly 10D-91.1104.

- 64E-5.1002 Monitoring Requirements.

The department will perform the following tests both prior to phosphate mining and subsequent to reclamation:

(1) Gamma radiation exposure measurements.

(2) Soil characterization measurements consisting of

(a) Radon emanation determinations.

(b) Soil radium determinations.

(3) Air monitoring determinations.

(4) Surface and ground water monitoring of such water that is potentially affected by mining activities unless existing test data is available.

Rulemaking Authority 404.051, 404.056 FS. Law Implemented 404.022(2), 404.051(4), 404.056 FS. History–New 11-13-85, Amended 9-26-91, 5-15-96, Formerly 10D-91.1111.

- 64E-5.1003 Monitoring Fees.

(1) The costs incurred by the department and the subsequent fees associated with the monitoring program shall be annually prorated on the basis of the total number of measurements to be made on land proposed to be mined and land proposed to be reclaimed, excluding bodies of water.

(2) The annual fees paid by the mining companies shall be calculated by the following method:

(a) Gamma radiation exposure measurements will be made at the rate of one per acre. An annual fee of $11.63 per measurement shall be assessed the mining company by the department.

(b) Soil characterization measurements will be made at the rate of one per 20 acres. An annual fee of $496.00 per each 20 acres measured shall be assessed the mining company by the department.

(c) Air monitoring measurements will be assessed at the rate of $255.75 per measurement. The department shall prorate the cost of air monitoring measurements among the mining companies based on their share of the total acreage to be mined and reclaimed each year. The department will conduct no more than 272 air monitoring measurements per year for all mining companies for which a fee will be assessed.

(d) Surface and ground water monitoring measurements will be assessed at the rate of $300.00 per measurement. The department shall prorate the cost of water monitoring among the mining companies based on their share of the total acreage to be mined and reclaimed each year. The department will analyze no more than 160 water samples per year for all mining companies for which a fee will be assessed.

(3) The department shall bill each affected company by March for activities which are planned for the following July through December time period. The department shall bill each affected company by September for activities which are planned for the following January through June time period. Each mining company shall remit the fee to the department within 60 days of the billing date.

(4) Overpayments or underpayments from an affected company will be reconciled by the department on an annual basis.

Rulemaking Authority 404.022, 404.051, 404.056, 404.131 FS. Law Implemented 404.022(2), 404.051(4), 404.131(5) FS. History–New 11-13-85, Amended 9-26-91, Formerly 10D-91.1112, Amended 2-28-08.

- 64E-5.101 Definitions.

As used in these rules, these terms have the definitions set forth below. Additional definitions used only in a certain part are defined in that respective part.

(1) “A1” means the maximum activity of special form radioactive material permitted in a Type A package.

(2) “A2” means the maximum activity of radioactive material, other than special form or low specific activity radioactive material, permitted in a Type A package.

(3) “Absorbed dose” means the energy imparted by ionizing radiation per unit mass of irradiated material. The units of absorbed dose are the gray (Gy) and the rad.

(4) “Accelerator-produced material” means any material made radioactive by a particle accelerator.

(5) “Act” means the Florida Radiation Protection Act, Chapter 404, F.S.

(6) “Activity” means the rate of disintegration or transformation or decay of radioactive material. The units of activity are the becquerel (Bq) and the curie (Ci).

(7) “Address of use” means the building or buildings that are identified on the license and where radioactive material may be received, used, or stored.

(8) “Adult” means an individual 18 or more years of age.

(9) “Airborne radioactive material” means any radioactive material dispersed in the air in the form of dusts, fumes, particulates, mists, vapors or gases.

(10) “Airborne radioactivity area” means a room, enclosure or operating area in which airborne radioactive materials exist in concentrations:

(a) In excess of the derived air concentrations (DACs) specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, which is herein incorporated by reference and which can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03449 or at http://www.doh.state.fl.us/environment/radiation/regs/64e-5tab.htm, or

(b) To such a degree that an individual present in the area without respiratory protective equipment could exceed, during the hours an individual is present in a week, an intake of 0.6 percent of the annual limit on intake (ALI) or 12 DAC-hours.

(11) “ALARA” means as low as reasonably achievable making every reasonable effort to maintain exposures to radiation as far below the dose limits in these rules as practical, consistent with the purpose for which the licensed or registered activity is undertaken, taking into account the state of technology, the economics of improvements in relation to the state of technology, the economics of improvements in relation to benefits to the public health and safety, and other societal and socioeconomic considerations, and in relation to use of nuclear energy and licensed or registered sources of radiation in the public interest.

(12) “Analytical x-ray equipment” means equipment used for x-ray diffraction or fluorescence analysis.

(13) “Analytical x-ray system” means a group of components utilizing x-rays to determine the elemental composition or to examine the microstructure of materials.

(14) “Annual limit on intake” (ALI) means the derived limit for the amount of radioactive material taken into the body of an adult worker by inhalation or ingestion in a year. ALI is the smaller value of intake of a given radionuclide in a year by Reference Man that would result in a committed effective dose equivalent of 5 rem (0.05 sievert) or a committed dose equivalent of 50 rem (0.5 sievert) to any individual organ or tissue. ALI values for intake by ingestion and by inhalation of selected radionuclides are given in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, Table I, Columns 1 and 2.

(15) “Area of use” means a portion of a physical structure that has been set aside to receive, use, or store radioactive material.

(16) “Authorized user” means an individual who is identified on a department, NRC, agreement state, or licensing state specific license that authorizes the use of radioactive material.

(17) “Background radiation” means radiation from cosmic sources; naturally occurring radioactive materials, including radon, except as a decay product of source or special nuclear material, and including global fallout as it exists in the environment from the testing of nuclear explosive devices or from past nuclear accidents such as Chernobyl that contribute to background radiation. Background radiation does not include sources of radiation from radioactive materials regulated by the department.

(18) “Baggage x-ray system” means a cabinet x-ray system with a maximum energy less than 120 kVp that produces only fluoroscopic images and that is used for packages or carry-on baggage.

(19) “Becquerel” (Bq) means the SI unit of activity. One becquerel is equal to 1 disintegration or transformation per second (s-1).

(20) “Bioassay” means the determination of kinds, quantities or concentrations, and in some cases, the locations of radioactive material in the human body, whether by direct measurement, in vivo counting, or by analysis and evaluation of materials excreted or removed from the human body. For purposes of these rules, “radiobioassay” is an equivalent term.

(21) “Byproduct material” means:

(a) Any radioactive material, except special nuclear material, yielded in or made radioactive by exposure to the radiation incident to the process of producing or utilizing special nuclear material; and

(b) The tailings or wastes produced by the extraction or concentration of uranium or thorium from ore processed primarily for its source material content, including discrete surface waste resulting from uranium or thorium solution extraction processes. Underground ore bodies depleted by these solution extraction operations do not constitute byproduct material within this definition.

(c)1. Any discrete source of radium-226 that is produced, extracted, or converted after extraction for use for a commercial, medical, or research activity; or

2. Any material that meets the following:

a. Has been made radioactive by use of a particle accelerator; and

b. Is produced, extracted, or converted after extraction for use for a commercial, medical, or research activity; and

(d) Any discrete source of naturally occurring radioactive material, other than source material, that meets the following:

1. The NRC, in consultation with the Administrator of the Environmental Protection Agency, the Secretary of Energy, the Secretary of Homeland Security, and the head of any other appropriate Federal agency, determines would pose a threat similar to the threat posed by a discrete source of radium-226 to the public health and safety or the common defense and security; and

2. Is extracted or converted after extraction for use in a commercial, medical, or research activity.

(22) “Cabinet x-ray system or Cabinet x-ray” means an x-ray system with the x-ray tube installed in an enclosure independent of existing architectural structures. A cabinet x-ray system is intended to contain the material being irradiated, and exclude personnel from its interior during generation of radiation. To be certified as a cabinet x-ray, the cabinet must be shielded so that every location on the exterior meets the conditions of 0.5 mRem (0.005 millisievert) in any one hour, at a distance of 5 cm. An x-ray tube used within a shielded part of a building or x-ray equipment that may temporarily or occasionally incorporate portable shielding is not considered a cabinet x-ray system.

(23) “Calendar quarter” means not less than 12 consecutive weeks nor more than 14 consecutive weeks. The first calendar quarter of each year shall begin on January 1 and subsequent calendar quarters shall be arranged so that no day is included in more than 1 calendar quarter, no calendar quarter, or part thereof, is included in more than 1 calendar year, and no day in any 1 year is omitted from inclusion within a calendar quarter. No licensee or registrant shall change the method observed by him to determine calendar quarters for purposes of these rules except at the beginning of a calendar year.

(24) “Calibration” means:

(a) The determination of the response or reading of an instrument relative to a series of known radiation values over the range of the instrument; or

(b) The determination of the strength of a source of radiation relative to a standard.

(25) “Carrier” means a person engaged in the transportation of passengers or property by land or water as a common, contract, or private carrier or by civil aircraft.

(26) “Class” means a classification scheme for inhaled material according to its rate of clearance from the pulmonary region of the lung. Materials are classified as D, W, or Y, which applies to a range of clearance half-times: for Class D, Days, of less than 10 days, for Class W, Weeks, from 10 to 100 days, and for Class Y, Years, of greater than 100 days. For purposes of these rules, “lung class” and “inhalation class” are equivalent terms.

(27) “Collective dose” means the sum of the individual doses received in a given period of time by a specified population from exposure to a specified source of radiation.

(28) “Committed dose equivalent” (HT, 50) means the dose equivalent to organs or tissues of reference (T) that will be received from an intake of radioactive material by an individual during the 50-year period following the intake.

(29) “Committed effective dose equivalent” (HE, 50) is the sum of the products of the weighting factors applicable to each of the body organs or tissues that are irradiated and the committed dose equivalent to each of these organs or tissues (HE, 50 =  WT HT, 50).

(30) “Constraint” or “dose constraint” means a value above which specified licensee actions are required.

(31) “Curie” means a unit of quantity of radioactivity. One curie (Ci) is that quantity of radioactive material which decays at the rate of 3.7 × 1010 transformations per second (tps).

(32) “Declared pregnant woman” means a woman who has voluntarily informed her employer in writing of her pregnancy and the estimated date of conception. The declaration remains in effect until the declared pregnant woman withdraws the declaration in writing or is no longer pregnant.

(33) “Dedicated check source” means a radioactive source that is used to assure the consistent operation of a radiation detection or measurement device over several months or years. This source may also be used for other purposes.

(34) “Deep dose equivalent” (Hd), which applies to external whole body exposure, means the dose equivalent at a tissue depth of 1 centimeter (1,000 mg/cm2).

(35) “Decommission” means to remove a facility safely from service and reduce residual radioactivity to a level that permits release of the property for unrestricted use and termination of license or release of the property under restricted conditions and the termination of the license.

(36) “Depleted uranium” means the source material uranium in which the isotope uranium 235 is less than 0.711 weight percent of the total uranium present. Depleted uranium does not include special nuclear material.

(37) “Derived air concentration” (DAC) means the concentration of a given radionuclide in air which, if breathed by Reference Man for a working year of 2,000 hours under conditions of light work, results in an intake of one ALI. For purposes of these rules, the condition of light work is an inhalation rate of 1.2 cubic meters of air per hour for 2,000 hours in a year. DAC values are given in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, Table I, Column 3.

(38) “Derived air concentration-hour” (DAC-hour) means the product of the concentration of radioactive material in air, expressed as a fraction or multiple of the derived air concentration for each radionuclide, and the time of exposure to that radionuclide, in hours. A licensee can take 2,000 DAC-hours to represent one ALI, equivalent to a committed effective dose equivalent of 5 rem (0.05 sievert).

(39) “Dose” is a generic term that means absorbed dose, dose equivalent, effective dose equivalent, committed dose equivalent, committed effective dose equivalent, or total effective dose equivalent. For the purposes of these rules, “radiation dose” is an equivalent term.

(40) “Dose equivalent” (HT) means the product of the absorbed dose in tissue, quality factor, and all other necessary modifying factors at the location of interest. The units of dose equivalent are the sievert (Sv) and rem.

(41) “Dose limits” means the permissible upper bounds of radiation doses established as specified in these rules. For purposes of these rules, “limits” is an equivalent term.

(42) “Dosimetry processor” means an individual or an organization that processes and evaluates individual monitoring devices to determine the radiation dose delivered to the monitoring devices.

(43) “Effective dose equivalent” (HE) means the sum of the products of the dose equivalent to each organ or tissue (HT) and the weighting factor (wT) applicable to each of the body organs or tissues that are irradiated (HE =  WT HT).

(44) “Embryo” or “fetus” means the developing human organism from conception until birth.

(45) “Entrance or access point” means any opening through which an individual or extremity of an individual could gain access to radiation areas or to licensed or registered radioactive materials. This includes entry or exit portals of sufficient size to permit human entry, irrespective of their intended use.

(46) “Exposure”, when used as a noun, means the quotient of dQ by dm, where “dQ” is the absolute value of the total charge of the ions of 1 sign produced in air when all the electrons, negatrons and positrons liberated by photons in a volume element of air having mass “dm” are completely stopped in air. “Exposure”, when used as a verb, means being exposed to ionizing radiation or to radioactive material. The special unit of exposure is the roentgen (R). See Rule 64E-5.106, F.A.C., for the SI equivalent.

(47) “Exposure rate” means the exposure per unit of time, such as roentgen per minute and milliroentgen per hour.

(48) “External dose” means that portion of the dose equivalent received from any source of radiation outside the body.

(49) “Extremity” means hand, elbow, arm below the elbow, foot, knee, and leg below the knee.

(50) “Eye dose equivalent” means the external dose equivalent to the lens of the eye at a tissue depth of 0.3 centimeter (300 mg/cm2).

(51) “Fail-safe characteristics” means a design feature which causes beam port shutters to close, or otherwise prevents emergence of the primary beam, upon the failure of a safety or warning device.

(52) “Field station” means a temporary or portable facility where radioactive sources may be stored or used and from which equipment is dispatched to temporary jobsites.

(53) “Former U.S. Atomic Energy Commission (AEC) or U.S. Nuclear Regulatory Commission (NRC) licensed facilities” means nuclear reactors, nuclear fuel reprocessing plants, uranium enrichment plants, or critical mass experimental facilities where AEC or NRC licenses have been terminated.

(54) “Generally applicable environmental radiation standards” means standards issued by the U.S. Environmental Protection Agency (EPA) under the authority of the Atomic Energy Act of 1954, as amended, that impose limits on radiation exposures or levels, or concentrations or quantities of radioactive material, in the general environment outside the boundaries of locations under the control of persons possessing or using radioactive material.

(55) “Gray” (Gy) means the SI unit of absorbed dose. One gray is equal to an absorbed dose of 1 joule per kilogram (100 rad).

(56) “Healing arts” means professions concerned with diagnosis or treatment of human and animal maladies, including the practice of medicine, dentistry, veterinary medicine, osteopathy, chiropractic, podiatry and naturopathy.

(57) “High radiation area” means an area, accessible to individuals, in which radiation levels from radiation sources external to the body could result in an individual receiving a dose equivalent in excess of 1 mSv (0.1 rem) in 1 hour at 30 centimeters from any source of radiation or from any surface that the radiation penetrates. For purposes of these rules, rooms or areas in which diagnostic x-ray systems are used for healing arts purposes are not considered high radiation areas.

(58) “Human use” means the internal or external administration of radiation or radioactive material to human beings.

(59) “Individual” means any human being.

(60) “Individual monitoring” means the assessment of:

(a) Dose equivalent by the use of individual monitoring devices or by the use of survey data; or

(b) Committed effective dose equivalent by bioassay or by determination of the time-weighted air concentrations to which an individual has been exposed.

(61) “Individual monitoring devices” means devices designed to be worn by a single individual for the assessment of dose equivalent such as film badges, thermoluminescence dosimeters, pocket ionization chambers, and personal or lapel air sampling devices. For purposes of these rules, individual monitoring equipment and personnel monitoring equipment are equivalent terms. Examples of individual monitoring devices are film badges, thermoluminescent dosimeters (TLDs), optically stimulated luminescent devices (OSLDs), pocket ionization chambers, and personal air sampling devices.

(62) “Industrial radiography” means nondestructive testing using ionizing radiation to make radiographic images or radiographs to detect flaws in objects.

(63) “Inhalation class” (see “Class”).

(64) “Injection tool” means a device used for controlled subsurface injection of radioactive tracer material.

(65) “Interlock” means a device arranged or connected such that the occurrence of an event or condition is required before a second event or condition can occur or continue to occur.

(66) “Internal dose” means that portion of the dose equivalent received from radioactive material taken into the body.

(67) “Large irradiator” means an irradiator where radiation dose rates exceeding 500 rems (5 sieverts) per hour exist at 1 meter from the sealed radioactive sources in air or in water. This does not include irradiators in which both sealed source and the area subject to irradiation are contained within a device and are not accessible to personnel, or to radioactive materials used for medical radiology, teletherapy, industrial radiography, gauging, calibration of radiation detection instruments, or open-field agricultural irradiations.

(68) “Lens dose equivalent (LDE)” applies to the external exposure of the lens of the eye and is taken as the dose equivalent at the tissue depth of 0.3 centimeter (300 mg/cm2).

(69) “License” means a license issued by the Department in accordance with the rules adopted by the Department.

(70) “Licensed material” means radioactive material received, possessed, used, transferred or disposed of under a general or specific license issued by the department.

(71) “Licensee” means any person who is licensed by the Department in accordance with these rules and the Act.

(72) “Licensing State” means any state with rules equivalent to the Suggested State Regulations for Control of Radiation for the regulatory control of NARM and which has been granted final designation by the Conference of Radiation Control Program Directors, Inc.

(73) “Local components” means parts of an analytical x-ray system and includes areas that are struck by x-rays, such as radiation source housings, port and shutter assemblies, collimators, sample holders, cameras, goniometers, detectors and shielding, but does not include power supplies, transformers, amplifiers, readout devices and control panels.

(74) “Logging supervisor” means the individual who provides personal supervision of the utilization of sources of radiation at the well site.

(75) “Logging tool” means a device used subsurface to perform well-logging.

(76) “Lost or missing licensed material” means licensed material whose location is unknown. This definition includes licensed material that has been shipped but has not reached its planned destination and whose location cannot be readily traced in the transportation system.

(77) “Low specific activity material (LSA)” means that as defined in 49 C.F.R. section 173.403, 10-1-12 edition, which is herein incorporated by reference and may be obtained at https://www.flrules.org/Gateway/reference.asp?No=Ref-03472 or at http://www.myfloridaeh.com/radiation/radmat1.htm.

(78) “Lung class” (see “Class”).

(79) “Major processor” means a user processing, handling or manufacturing radioactive material exceeding A2 quantities as unsealed sources or material, or exceeding 4 times A1 quantities as sealed sources, but does not include nuclear medicine programs, universities, industrial radiographers or small industrial programs. A1 and A2 quantities can be found in Part XV.

(80) “Management” means the chief executive officer or other individual, or a delegate or the delegates of the chief executive officer or other individual, having the authority to manage, direct, or administer the licensee’s activities.

(81) “Medical institution” means any establishment that:

(a) Offers services more intensive than those required for room, board, personal services, and general nursing care, and offers facilities and beds for use beyond 24 hours by individuals requiring diagnosis, treatment, or care for illness, injury, deformity, infirmity, abnormality, disease, or pregnancy; and

(b) Regularly makes available at least clinical laboratory services, diagnostic X-ray services, and treatment facilities for surgery or obstetrical care, or other definitive medical treatment of similar extent.

(82) “Member of the public” means any individual except when that individual is receiving an occupational dose.

(83) “Mineral logging” means any logging performed for the purpose of mineral exploration other than oil or gas.

(84) “Minor” means an individual less than 18 years of age.

(85) “Medical event” means the administration of:

(a) Radioactive materials or radiation from radioactive materials requiring a written directive that results in the following:

1. A dose that differs from the prescribed dose or dose that would have resulted from the prescribed dosage by more than 0.05 Sv (5 rem) effective dose equivalent, 0.5 Sv (50 rem) to an organ or tissue, or 0.5 Sv (50 rem) shallow dose equivalent to the skin;

2. When the total dose delivered differs from the prescribed dose by 20 percent or more;

3. The total dosage delivered differs from the prescribed dosage by 20 percent or more or falls outside the prescribed dosage range;

4. The fractionated dose delivered differs from the prescribed dose, for a single fraction, by 50 percent or more;

5. A dose that exceeds 0.05 Sv (5 rem) effective dose equivalent, 0.5 Sv (50 rem) to an organ or tissue, or 0.5 Sv (50 rem) shallow dose equivalent to the skin;

6. An administration of a wrong radioactive drug containing radioactive material;

7. An administration of a radioactive drug containing radioactive material by the wrong route of administration;

8. An administration of a dose or dosage to the wrong individual or human research subject;

9. An administration of a dose or dosage delivered by the wrong mode of treatment;

10. A leaking sealed source where the patient or human research subject is contaminated;

11. A dose to the skin or an organ or tissue other than the treatment site that exceeds by 0.5 Sv (50 rem) to an organ or tissue and 50 percent or more of the dose expected from the administration defined in the written directive (excluding, for permanent implants, seeds that were implanted in the correct site but migrated outside the treatment site); or

12. Any medical use that results or will result in unintended permanent functional damage to an individual’s organ or a physiological system, as determined by a physician.

(b) Radioactive materials or radiation from radioactive materials not requiring a written directive that result in either of the following:

1. A dose that differs from the prescribed dose or dose that would have resulted from the prescribed dosage by more than 0.05 Sv (5 rem) effective dose equivalent, 0.5 Sv (50 rem) to an organ or tissue, or 0.5 Sv (50 rem) shallow dose equivalent to the skin; and

a. When the total dose delivered differs from the prescribed dose by 20 percent or more;

b. The total dosage delivered differs from the prescribed dosage by 20 percent or more or falls outside the prescribed dosage range;

c. The fractionated dose delivered differs from the prescribed dose, for a single fraction, by 50 percent or more; or

2. A dose that exceeds 0.05 Sv (5 rem) effective dose equivalent, 0.5 Sv (50 rem) to an organ or tissue, or 0.5 Sv (50 rem) shallow dose equivalent to the skin; and

a. An administration of a wrong radioactive drug containing radioactive material;

b. An administration of a radioactive drug containing radioactive material by the wrong route of administration;

c. An administration of a dose or dosage to the wrong individual or human research subject;

d. An administration of a dose or dosage delivered by the wrong mode of treatment; or

e. A leaking sealed source where the patient or human research subject is contaminated.

3. Any medical use that results or will result in unintended permanent functional damage to an individual’s organ or a physiological system, as determined by a physician.

(c) Radiation from a therapeutic x-ray machine or particle accelerator that result in any of the following:

1. Any medical use that results or will result in unintended permanent functional damage to an individual’s organ or a physiological system, as determined by a physician;

2. An administration of a dose to the wrong individual or human research subject;

3. An administration of a dose delivered by the wrong mode of treatment, wrong treatment, or wrong treatment site;

4. When treatment consists of three or fewer fractions and the calculated total administered dose differs from the total prescribed dose by more than 10 percent of the total prescribed dose;

5. When the calculated weekly administered dose is 30 percent greater than the weekly prescribed dose; or

6. When the calculated total administered dose differs from the total prescribed dose by more than 20 percent of the total prescribed dose.

(86) “Monitoring” means the measurement of radiation, radioactive material concentrations, surface area activities or quantities of radioactive material and the use of the results of these measurements to evaluate potential exposures and doses. For purposes of these rules, radiation monitoring and radiation protection monitoring are equivalent terms.

(87) “NARM” means any naturally occurring or accelerator-produced radioactive material. To meet the definition of licensing state, NARM only refers to discrete sources of NARM. Diffuse sources of NARM, which are large in volume and low in activity, are excluded from consideration by the Conference of Radiation Control Program Directors, Inc., for licensing state designation purposes.

(88) “Natural radioactivity” means radioactivity of naturally occurring nuclides.

(89) “Nonstochastic effect” means a health effect the severity of which varies with the dose and for which a threshold is believed to exist. Radiation-induced cataract formation is an example of a nonstochastic effect. For purposes of these rules, “deterministic effect” is an equivalent term.

(90) “Normal form” means radioactive material which has not been demonstrated to qualify as “special form”; also referred to as “nonspecial form”.

(91) “Normal operating procedures” means operating procedures for conditions suitable for analytical purposes with shielding and barriers in place. These do not include maintenance but do include routine alignment procedures. Routine and emergency radiation safety considerations are part of these procedures.

(92) “Nuclear Regulatory Commission” (NRC) means the U.S. Nuclear Regulatory Commission or its duly authorized representatives.

(93) “Occupational dose” means the dose received by an individual in the course of employment in which the individual’s assigned duties involve exposure to sources of radiation, whether in the possession of the licensee, registrant, or other person. Occupational dose does not include dose received from background radiation, from any medical administration the individual has received, from exposure to individuals administered radioactive material and released as specified in Rule 64E-5.622, F.A.C., from voluntary participation in medical research programs, or as a member of the public.

(94) “Offshore” means within the territorial waters of the State of Florida as specified in Article II, Section 1 of the Constitution of the State of Florida.

(95) “Open-beam configuration” means an analytical x-ray system in which an individual could accidentally place some part of his body in the primary beam path during normal operation.

(96) “Package” means that as defined in 49 C.F.R. section 173.403, 10-1-12 edition.

(97) “Packaging” means, for radioactive materials, the assembly of components necessary to ensure compliance with the packaging requirements of the U.S. Nuclear Regulatory Commission and the U.S. Department of Transportation. It may consist of one or more receptacles, absorbent materials, spacing structures, thermal insulation, radiation shielding, and devices for cooling or absorbing mechanical shocks. The conveyance, tie-down system, and auxiliary equipment may sometimes be designated as part of the packaging.

(98) “Particle accelerator” means any machine capable of accelerating electrons, protons, deuterons, or other charged particles in a vacuum and of discharging the resultant particulate or other radiation into a medium at energies usually in excess of 1 MeV.

(99) “Permanent radiographic installation” means an enclosed shielded room, cell, or vault, as specified in Rule 64E-5.431, F.A.C., in which industrial radiography is performed.

(100) “Permit” means the written authorization issued by the Department for the transportation of radioactive waste as described in Rule 64E-5.1509, F.A.C.

(101) “Personal supervision” means supervision in which the radiographer or logging supervisor is physically present at the site where sources of radiation and associated equipment are being used, watching the performance of the radiographer’s assistant or supervised individual and in such proximity that immediate assistance can be given if required.

(102) “Planned special exposure” means an infrequent exposure to radiation, separate from and in addition to the annual occupational dose limits.

(103) “Prescribed Dosage” means the quantity of radiopharmaceutical activity as documented:

(a) In a written directive; or

(b) Either in the diagnostic clinical procedures manual or in any appropriate record as specified in the directions of the authorized user for diagnostic procedures in which a written directive is not required.

(104) “Prescribed Dose” means:

(a) For gamma stereotactic radiosurgery, the total dose as documented in the written directive;

(b) For manual brachytherapy, either the total source strength and exposure time or the total dose as documented in the written directive;

(c) For teletherapy, particle accelerator or therapeutic x-ray machine, the total dose and dose per fraction as documented in the written directive; or

(d) For remote brachytherapy afterloaders, the total dose and dose per fraction as documented in the written directive.

(105) “Primary beam” means the radiation which passes through an aperture of the source housing in a direct path from the x-ray tube located in the radiation source housing.

(106) “Public dose” means the dose received by a member of the public from exposure to radiation or radioactive materials released by a licensee or registrant, or to any other sources of radiation under the control of the licensee or registrant. Public dose does not include occupational dose or doses received from background radiation, from any medical administration the individual has received, from exposure to individuals administered radioactive materials and released as specified in Rule 64E-5.622, F.A.C., or from voluntary participation in medical research programs.

(107) “Quality factor” (Q) means the modifying factor listed in the tables in subsections 64E-5.106(3) and (4), F.A.C., used to derive dose equivalent from absorbed dose.

(108) “Quarter” means a period of time equal to one-fourth of the year observed by the licensee or registrant of approximately 13 consecutive weeks. The beginning of the first quarter in a year shall coincide with the starting date of the year and no day shall be omitted or duplicated in consecutive quarters.

(109) “Rad” means the special unit of absorbed dose. One rad is equal to an absorbed dose of 100 ergs per gram or 0.01 joule per kilogram (0.01 gray).

(110) “Radiation” means alpha particles, beta particles, gamma rays, x-rays, neutrons, high-speed electrons, high-speed protons, and other particles capable of producing ions. For purposes of these rules, “ionizing radiation” is an equivalent term. Radiation, as used in these rules, does not include nonionizing radiation, such as radio waves or microwaves, visible, infrared, or ultraviolet light.

(111) “Radiation area” means any area, accessible to individuals, in which radiation levels could result in an individual’s receiving a dose equivalent in excess of 0.05 mSv (0.005 rem) in 1 hour at 30 centimeters from the source of radiation or from any surface that the radiation penetrates.

(112) “Radiation machine” means any device capable of producing radiation except those devices with radioactive material as the only source of radiation.

(113) “Radiation Safety Officer or RSO” means a person who has the knowledge and responsibility to apply appropriate radiation protection rules.

(114) “Radioactive marker” means radioactive material placed subsurface or on a structure intended for subsurface use for the purpose of depth determination or direction orientation.

(115) “Radioactivity” means the transformation of unstable atomic nuclei by the emission of radiation.

(116) “Radiographer” means any individual who has completed successfully the training and testing requirements specified in subsection 64E-5.434(2), F.A.C., performs or personally supervises radiographic operations, and is responsible to the licensee or registrant for assuring compliance with the requirements of these rules and all license or certificate of registration conditions.

(117) “Radiographer’s assistant or assistant radiographer” means any individual who has completed successfully the training and testing requirements specified in subsection 64E-5.434(1), F.A.C., and who, under the personal supervision of a radiographer, conducts radiographic operations.

(118) “Radiographic exposure device” means any instrument containing a sealed source, fastened or contained therein, in which the sealed source or shielding thereof may be moved, or otherwise changed from a shielded position to an unshielded position for the purpose of making a radiographic exposure. It also is known as a camera or a projector.

(119) “Recordable event” means the administration of:

(a) A radiopharmaceutical or radiation without a written directive where a written directive is required;

(b) A radiopharmaceutical or radiation where a written directive is required without daily recording of each administered radiopharmaceutical dosage or radiation dose in the appropriate record;

(c) Iodine 131 as sodium iodide in quantities greater than 30 microcuries (1.11 megabecquerels) when;

1. The administered dosage differs from the prescribed dosage by more than 10 percent of the prescribed dosage; and

2. The difference between the administered dosage and the prescribed dosage exceeds 15 microcuries.

(d) A therapeutic administration of a radiopharmaceutical other than iodine 131 as sodium iodide, when the administered dosage differs from the prescribed dosage by more than 10 percent from the prescribed dosage;

(e) A brachytherapy radiation dose when the calculated administered dose differs from the prescribed dose by more than 10 percent of the prescribed dose; or

(f) A teletherapy, particle accelerator, gamma stereotactic radiosurgery or therapeutic x-ray machine radiation dose when the calculated weekly administered dose is 15 percent greater than the weekly prescribed dose.

(120) “Reference Man” means a hypothetical aggregation of human physical and physiological characteristics determined by international consensus. These characteristics can be used by researchers and public health workers to standardize results of experiments and to relate biological insult to a common base. A description of Reference Man is contained in the International Commission on Radiological Protection report, ICRP Publication 23, “Report of the Task Group on Reference Man.”

(121) “Registrant” means any person who is registered with the Department and is legally obliged to register with the Department pursuant to these rules and the Act.

(122) “Regulations of the U.S. Department of Transportation” means the regulations in 49 C.F.R. Parts 100-189.

(123) “Rem” means the special unit of any of the quantities expressed as dose equivalent. The dose equivalent in rem is equal to the absorbed dose in rad multiplied by the quality factor (1 rem = 0.01 sievert).

(124) “Research and development” means:

(a) Theoretical analysis, exploration or experimentation; or

(b) The extension of investigative findings and theories of a scientific or technical nature into practical application for experimental and demonstration purposes, including the experimental production and testing of models, devices, equipment, materials and processes. Research and development does not include the internal or external administration of radiation or radioactive material to human beings.

(125) “Respiratory protective equipment” means an apparatus, such as a respirator, used to reduce an individual’s intake of airborne radioactive materials.

(126) “Restricted area” means an area, access to which is limited by the licensee or registrant to protect individuals against undue risks from exposure to sources of radiation. A restricted area does not include areas used as residential quarters, but separate rooms in a residential building can be set apart as a restricted area.

(127) “Roentgen” means the special unit of exposure. One roentgen (R) equals 2.58 × 10-4 coulombs per kilogram of air.

(128) “Sanitary sewerage” means a system of public sewers for carrying off waste water and refuse, but excluding sewage treatment facilities, septic tanks, and leach fields owned or operated by the licensee or registrant.

(129) “Sealed source” means radioactive material that is encased in a capsule designed to prevent release or escape of the radioactive material.

(130) “Shallow dose equivalent” (HS), which applies to the external exposure of the skin of the whole body or the skin of an extremity, means the dose equivalent at a tissue depth of 0.007 centimeter (7 mg/cm2).

(131) “Shielded position” means the location within the radiographic exposure device or source changer where the sealed source is secured and restricted from movement.

(132) “Shipping paper” means a shipping order, bill of lading, manifest or other shipping document serving a similar purpose and containing the information required by 49 C.F.R., Parts 172.202, 172.203 and 172.204.

(133) “SI” means an abbreviation of the International System of Units.

(134) “Sievert” means the SI unit of any of the quantities expressed as dose equivalent. The dose equivalent in sievert is equal to the absorbed dose in gray multiplied by the quality factor (1 Sv = 100 rem).

(135) “Source changer” means a device designed and used for replacement of sealed sources in radiographic exposure devices, including those source changers also used for transporting and storage of sealed sources.

(136) “Source holder” means a housing or assembly into which a radioactive source is placed for the purpose of facilitating the handling and use of the source in well-logging operations.

(137) “Source material” means:

(a) Uranium or thorium, or any combination thereof, in any physical or chemical form; or

(b) Ores which contain by weight one-twentieth of 1 percent (0.05 percent) or more of uranium, thorium or any combination of uranium and thorium. Source material does not include special nuclear material.

(138) “Source material milling” means any activity that results in the production of byproduct material as defined by Rule 64E-5.101, F.A.C.

(139) “Source of radiation” means any radioactive material or any device or equipment emitting, or capable of producing, radiation.

(140) “Special form” means radioactive material which satisfies all of the following conditions:

(a) It is either a single solid piece or is contained in a sealed capsule that can be opened only by destroying the capsule;

(b) The piece or capsule has at least one dimension not less than 5 millimeters; and

(c) It satisfies the test requirements of 49 C.F.R., Part 173.469. Special form encapsulations designed in accordance with the requirements of 49 C.F.R., Part 173.389 in effect on June 30, 1983, and constructed prior to July 1, 1985, may continue to be used. Special form encapsulations either designed or constructed after June 30, 1985, must meet the requirements of this part.

(141) “Special nuclear material in quantities not sufficient to form a critical mass” means uranium enriched in the isotope U-235 in quantities not exceeding 350 grams of contained U-235; uranium 233 in quantities not exceeding 200 grams; plutonium in quantities not exceeding 200 grams; or any combination of them in accordance with the following formula: For each kind of special nuclear material, determine the ratio between the quantity of that special nuclear material and the quantity specified above for the same kind of special nuclear material. The sum of such ratios for all of the kinds of special nuclear material in combination shall not exceed 1. For example, the following quantities in combination would not exceed the limitation and are within the formula:

| 175 (grams contained U-235) | + | 50 (grams U-233) | + | 50 (grams Pu) | = | 1 |
| --- | --- | --- | --- | --- | --- | --- |
| 350 |  | 200 |  | 200 |  |  |

(142) “Specific activity” means the activity of a radionuclide per unit mass of that nuclide. The specific activity of a material in which the radionuclide is essentially uniformly distributed is the activity per unit mass of the material.

(143) “Stochastic effect” means a health effect that occurs randomly and for which the probability of the effect occurring, rather than its severity, is assumed to be a linear function of dose without threshold. Hereditary effects and cancer incidence are examples of stochastic effects. For the purposes of these rules, “probabilistic effect” is an equivalent term.

(144) “Storage area” means any location, facility, or vehicle which is used to store, transport, or secure a radiographic exposure device, a storage container, or a sealed source when it is not in use and which is locked or has a physical barrier to prevent accidental exposure, tampering with, or unauthorized removal of the device, container, or source.

(145) “Storage container” means a container in which sealed sources are secured and stored.

(146) “Subsurface tracer study” means the release of a substance tagged with radioactive material for the purpose of tracing the movement or position of the tagged substance in the well-bore or adjacent formation.

(147) “Survey” means an evaluation of the radiological conditions and potential hazards incident to the production, use, transfer, release, disposal or presence of sources of radiation. When appropriate, such evaluation includes tests, physical examinations, and measurements of levels of radiation or concentrations of radioactive material present.

(148) “Teletherapy” means therapeutic irradiation in which the source of radiation is at a distance from the body.

(149) “Temporary job site” means a site, base or facility that is created and maintained to support a single job.

(150) “Test” means the process of verifying compliance with an applicable rule.

(151) “Total effective dose equivalent” (TEDE) means the sum of the effective dose equivalent for external exposures and the committed effective dose equivalent for internal exposures.

(152) “Type B packaging” means a packaging designed to retain the integrity of containment and shielding required by U.S. Nuclear Regulatory Commission and U.S. Department of Transportation regulations when subjected to the normal conditions of transport and hypothetical accident test conditions set forth in 10 C.F.R., Part 71.

(153) “Unrefined and unprocessed ore” means ore in its natural form prior to any processing, such as grinding, roasting, beneficiating or refining.

(154) “Unrestricted area” means an area, access to which is neither limited nor controlled by the licensee or registrant. For purposes of these rules, “uncontrolled area” is an equivalent term.

(155) “U.S. Department of Energy” means the Department of Energy established by Public Law 95-91, August 4, 1977, 91 Stat. 565, 42 U.S.C. 7101 et seq., to the extent that the Department exercises functions formerly vested in the U.S. Atomic Energy Commission, its Chairman, members, officers and components and transferred to the U.S. Energy Research and Development Administration and to the Administrator thereof as specified in sections 104(b), (c) and (d) of the Energy Reorganization Act of 1974 (Public Law 93-438, October 11, 1974, 88 Stat. 1233 at 1237, 42 U.S.C. 5814, effective January 19, 1975) and retransferred to the Secretary of Energy as specified in section 301(a) of the Department of Energy Organization Act (Public Law 95-91, August 4, 1977, 91 Stat. 565 at 577-578, 42 U.S.C. 7151, effective October 1, 1977.)

(156) “Very high radiation area” means an area, accessible to individuals, in which radiation levels from radiation sources external to the body could result in an individual receiving an absorbed dose in excess to 500 rad (5 gray) in 1 hour at 1 meter from a source of radiation or from any surface that the radiation penetrates. At very high doses received at high dose rates, units of absorbed dose, gray and rad, are appropriate, rather than units of dose equivalent, sievert and rem.

(157) “Visiting authorized user” means an authorized user who is not identified on the license.

(158) “Waste handling licensees” means persons licensed to receive and store radioactive wastes prior to disposal or persons licensed to dispose of radioactive waste.

(159) “Weighting factor” (WT) for an organ or tissue (T) means the proportion of the risk of stochastic effects resulting from irradiation of that organ or tissue to the total risk of stochastic effects when the whole body is irradiated uniformly. For calculating the effective dose equivalent, the values of WT are:

ORGAN DOSE WEIGHTING FACTORS

| Organ or Tissue | WT |
| --- | --- |
| Gonads | 0.25 |
| Breast | 0.15 |
| Red Bone Marrow | 0.12 |
| Lung | 0.12 |
| Thyroid | 0.03 |
| Bone Surfaces | 0.03 |
| Remainder | 0.30* |
| Whole Body | 1.00** |

*The 0.30 weighting factor for remainder results from 0.06 for each of 5 “remainder” organs, excluding the skin and the lens of the eye, that receive the highest doses.

**To weight the external whole body dose to add it to the internal dose, a single weighting factor, WT = 1.0, has been specified. The department will consider the use of other weighting factors for external exposure.

(160) “Well-bore” means a drilled hole in which wireline service operations and subsurface tracer studies are performed.

(161) “Well-logging” means the lowering and raising of measuring devices or tools which may contain sources of radiation into well-bores or cavities for the purpose of obtaining information about the well or adjacent formations.

(162) “Whole body” means, for purposes of external exposure, head, trunk including male gonads, arms above the elbow, or legs above the knee.

(163) “Wireline” means a cable containing one or more electrical conductors which is used to lower and raise logging tools in the well-bore.

(164) “Wireline service operation” means any evaluation or mechanical service which is performed in the well-bore using devices on a wireline.

(165) “Worker” means an individual engaged in work in a restricted area under the authority of a license or registration issued by the Department.

(166) “Working level” (WL) means any combination of short-lived radon daughters in 1 liter of air that will result in the ultimate emission of 1.3 × 105 MeV of potential alpha particle energy. The short-lived radon daughters are:

(a) For radon 222: polonium 218, lead 214, bismuth 214, and polonium 214;

(b) For radon 220: polonium 216, lead 212, bismuth 212, and polonium 212.

(167) “Working level month” (WLM) means an exposure to 1 working level for 170 hours. Two thousand working hours per year divided by 12 months per year is approximately equal to 170 hours per month.

(168) “Written directive” means a written order for a specific patient or human research subject, dated and signed by an authorized user prior to the administration of a radiopharmaceutical or radiation, which shall contain the following information:

(a) For a therapeutic administration of a radiopharmaceutical, the radiopharmaceutical, dosage, and route of administration;

(b) For any administration of iodine 131 as sodium iodide in quantities greater than 30 microcuries (1.11 megabecquerels), the dosage;

(c) For gamma stereotactic radiosurgery, target coordinate settings per treatment for each anatomically distinct treatment site, collimator size, plug pattern, and total dose;

(d) For teletherapy, particle accelerator or therapeutic x-ray machine, the total dose, dose per fraction, treatment site, number of fractions and overall treatment period;

(e) For high dose rate remote afterloading brachytherapy, the radioisotope, treatment site, dose per fraction, number of fractions, and total dose; and

(f) For all other brachytherapy, including low, medium, and pulsed dose rate remote afterloaders,

1. Prior to implantation, the radioisotope, treatment site, dose, number of sources, and source strengths; and

2. After implantation but prior to completion of the procedure, the radioisotope, treatment site, total source strength and exposure time or total dose.

(169) “Year” means the period of time beginning in January used to determine compliance with the provisions of these rules. The licensee or registrant can change the starting date of the year used to determine compliance by the licensee or registrant if the change is made at the beginning of the year and if no day is omitted or duplicated in consecutive years.

(170) “Principal activities” means activities authorized by the license that are essential to achieve the purpose for which the department issued or amended the license. Storage during which no licensed material is accessed for use or disposal and activities incidental to decontamination or decommissioning are not principal activities.

(171) “Authorized nuclear pharmacist” means a pharmacist who satisfies the following:

(a) Meets the requirements in subsection 64E-5.659(1) and Rule 64E-5.658, F.A.C.; or

(b) Authorized on a radioactive materials license by the department or identified as an authorized nuclear pharmacist on one of the following:

1. A specific license issued by the NRC or agreement state that authorizes medical use or the practice of nuclear pharmacy;

2. A permit issued by a NRC master material licensee that authorizes medical use or the practice of nuclear pharmacy;

3. A permit issued by a NRC or agreement state broad scope medical use licensee that authorizes medical use or the practice of nuclear pharmacy; or

4. A permit issued by a NRC master material broad scope licensee that authorizes medical use or the practice of nuclear pharmacy; or

(c) Is identified as an authorized nuclear pharmacist by a commercial nuclear pharmacy that has been authorized to identify authorized nuclear pharmacists; or

(d) Is designated as an authorized nuclear pharmacist in accordance with subparagraph 64E-5.210(10)(b)3., F.A.C.

(172) “Critical Group” means the group of individuals reasonably expected to receive the greatest exposure to residual radioactivity for any applicable set of circumstances.

(173) “Distinguishable from background” means that the detectable concentration of a radionuclide is statistically different from the background concentrations of that radionuclide in the vicinity of the site or, in the case of structures, in similar materials using adequate measurement technology, survey, and statistical techniques.

(174) “Residual radioactivity” means radioactivity in structures, materials, soils, groundwater, and other media at a site resulting from activities under the licensee’s control. This includes radioactivity from all licensed and unlicensed sources used by the licensee but excludes background radiation. It also includes radioactive material as a result of routine or accidental releases of radioactive material at the site and previous burials at the site even if those burial sites were made as specified in Part III of this Chapter.

(175) “Assigned protection factor” or “APF” means the expected workplace level of respiratory protection provided by a properly functioning respirator or a class of respirators to properly fitted and trained users. Operationally, the inhaled concentration can be estimated by dividing the ambient airborne concentration by the APF.

(176) “Atmosphere-supplying respirator” means a respirator that supplies the respirator user with breathing air from a source independent of the ambient atmosphere and includes supplied-air respirators and self-contained breathing apparatus units.

(177) “Energy compensation source” or “ECS” means a small sealed source with an activity not exceeding 100 microcuries (3.7 MBq) used within a logging tool or other tool components to provide a reference standard to maintain the tool’s calibration when in use.

(178) “Fit factor” means a quantitative estimate of the fit of a particular respirator to a specific individual and typically estimates the ratio of the concentration of a substance in ambient air to its concentration inside the respirator when worn.

(179) “Fit test” means the use of a protocol to evaluate qualitatively or quantitatively the fit of a respirator on an individual.

(180) “Self-contained breathing apparatus” or “SCBA” means an atmosphere-supplying respirator for which the breathing air source is designed to be carried by the user.

(181) “Supplied-air respirator” or “air-line respirator” means an atmosphere-supplying respirator for which the source of breathing air is not designed to be carried by the user.

(182) “Tritium neutron generator target source” means a tritium source used within a neutron generator tube to produce neutrons for use in well logging applications.

(183) “User seal check” or “fit check” means an action conducted by the respirator user to determine if the respirator is seated to the face properly. Examples include negative pressure check, positive pressure check, irritant smoke check, and isoamyl acetate check.

(184) “Annual or Annually” means an interval not to exceed 12 months.

(185) “Semiannual or Semiannually” means an interval not to exceed six months.

(186) “Daily” means an interval not to exceed a consecutive 24 hour period or once every calendar day worked.

(187) “Mobile C-arm” means a mobile c-arm fluoroscope designed for use without a specific patient support device. This includes machines moved from room to room to assist in surgical procedures.

(188) “C-arm system” means a mobile C-arm used in the same room with the same patient support device.

(189) “Nationally tracked source” means a sealed source containing a quantity equal to or greater than Category 1 or Category 2 levels of any radioactive material listed in Rule 64E-5.351, F.A.C. In this context a sealed source is defined as radioactive material that is sealed in a capsule or closely bonded, in a solid form, and which is not exempt from regulatory control. It does not mean material encapsulated solely for disposal, or nuclear material contained in any fuel assembly, subassembly, fuel rod, or fuel pellet. Category 1 nationally tracked sources are those containing radioactive material at a quantity equal to or greater than the Category 1 threshold. Category 2 nationally tracked sources are those containing radioactive material at a quantity equal to or greater than the Category 2 threshold but less than the Category 1 threshold.

(190) “Sealed Source and Device Registry” means the national registry that contains all the registration certificates, generated by both NRC and the agreement states, that summarize the radiation safety information for the sealed sources and devices and describe the licensing and use conditions approved for the product.

(191) “Stereotactic radiosurgery” means the use of external radiation in conjunction with a stereotactic guidance device to very precisely deliver a therapeutic dose to a tissue volume.

(192) “Structured educational program” means an educational program designed to impart particular knowledge and practical education through interrelated studies and supervised training.

(193) “C-arm fluoroscope” means a fluoroscopic machine where the image receptor and the x-ray tube housing assembly are ganged allowing a change in the direction of the beam axis with respect to the patient without moving the patient.

(194) “Extremity-use-only fluoroscope” means a fluoroscope manufactured after June 10, 2006, having a maximum source-image receptor distance of less than 45 centimeters and labeled “Extremity-use-only”.

(195) “Discrete source” means a radionuclide that has been processed so that its concentration within a material has been purposely increased for use for commercial, medical, or research activities.

(196) “Positron Emission Tomography (PET) radionuclide production facility” means a facility operating a cyclotron or accelerator for the purpose of producing PET radionuclides.

(197) “Consortium” means an association of medical use licensees and a PET radionuclide production facility in the same geographical area that jointly own or share in the operation and maintenance cost of the PET radionuclide production facility that produces PET radionuclides for use in producing radioactive drugs within the consortium for noncommercial distributions among its associated members for medical use. The PET radionuclide production facility within the consortium must be located at an educational institution or a Federal facility or a medical facility.

(198) “Waste” or “Radioactive Waste” means those low-level radioactive wastes containing source, special nuclear or other radioactive material that are acceptable for disposal in a land disposal facility. For the purposes of this definition, low-level radioactive waste means radioactive waste not classified as high-level radioactive waste, transuranic waste, spent nuclear fuel, or radioactive material as defined in paragraphs 64E-5.101(21)(b), (c) and (d).

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.031, 404.051, 404.061, 404.20, 404.22 FS. History–New 7-17-85, Amended 4-4-89, 5-12-93, 1-1-94, 5-15-96, Formerly 10D-91.102, Amended 5-18-98, 10-8-00, 8-6-01, 9-11-01, 12-18-01, 9-28-06, 8-16-07, 2-28-08, 2-11-10, 5-8-13, 12-26-13.

- 64E-5.102 Exemptions.

(1) The Department may, upon application therefor or upon its own initiative, grant such exemptions or exceptions from the requirements of these regulations as it determines are authorized by law and will not result in undue hazard to public health and safety or property or the environment.

(2) Any U.S. Department of Energy contractor or subcontractor and any U.S. Nuclear Regulatory Commission contractor or subcontractor of the following categories operating within this State is exempt from these regulations to the extent that such contractor or subcontractor under his contract receives, possesses, uses, transfers, transports or acquires sources of radiation:

(a) Prime contractors performing work for the U.S. Department of Energy at U.S. Government-owned or-controlled sites, including the transportation of sources of radiation to or from such sites and the performance of contract services during temporary interruptions of such transportation;

(b) Prime contractors of the U.S. Department of Energy performing research in, or development, manufacture, storage, testing, or transportation of, atomic weapons or components thereof;

(c) Prime contractors of the U.S. Department of Energy using or operating nuclear reactors or other nuclear devices in a United States Government-owned vehicle or vessel; and

(d) Any other prime contractor or subcontractor of the U.S. Department of Energy or of the U.S. Nuclear Regulatory Commission when the State and the U.S. Nuclear Regulatory Commission jointly determine:

1. That the exemption of the prime contractor or subcontractor is authorized by law; and

2. That under the terms of the contract or subcontract, there is adequate assurance that the work thereunder can be accomplished without undue risk to the public health, safety and environment.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.051(10), 404.061(4), 404.111(3), 404.121(1), 404.20 FS. History–New 7-17-85, Amended 5-12-93, Formerly 10D-91.103.

- 64E-5.103 Records.

Each licensee and registrant shall maintain records showing the receipt, transfer and disposal of all sources of radiation. Additional record requirements are specified elsewhere in these regulations.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.061(2), 404.081, 404.20(2), 404.22(2) FS. History–New 7-17-85, Formerly 10D-91.104.

- 64E-5.104 Tests.

Each licensee and registrant shall perform upon instructions from the Department, and shall permit the Department to perform, such reasonable tests as the Department deems appropriate and necessary, including tests of:

(1) Sources of radiation;

(2) Facilities wherein sources of radiation are used or stored;

(3) Radiation detection and monitoring instruments; and

(4) Other equipment and devices used in connection with utilization or storage of licensed or registered sources of radiation.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (7), 404.061(2), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.106.

- 64E-5.105 Prohibited Uses.

(1) A hand-held fluoroscopic screen shall not be used unless it has been listed in the Registry of Sealed Source and Devices or accepted for certification by the U.S. Food and Drug Administration, Center for Devices and Radiological Health.

(2) A shoe-fitting fluoroscopic device shall not be used.

Rulemaking Authority 404.051, 404.061, 404.141 FS. Law Implemented 404.022, 404.051, 404.061(2), 404.141, 404.22(3) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.110.

- 64E-5.106 Units of Exposure and Dose.

(1) As used in these regulations, the unit of exposure is the coulomb per kilogram (C/kg). One roentgen is equal to 2.58 × 10-4 coulomb per kilogram of air.

(2) As used in these regulations, the units of dose are:

(a) Gray (Gy) is the SI unit of absorbed dose. One gray is equal to an absorbed dose of 1 joule per kilogram (100 rad).

(b) Rad is the special unit of absorbed dose. One rad is equal to an absorbed dose of 100 erg per gram or 0.01 joule per kilogram (0.01 Gy).

(c) Rem is the special unit of any of the quantities expressed as dose equivalent. The dose equivalent in rem is equal to the absorbed dose in rad multiplied by the quality factor (1 rem = 0.01 Sv).

(d) Sievert is the SI unit of any of the quantities expressed as dose equivalent. The dose equivalent in sievert is equal to the absorbed dose in gray multiplied by the quality factor (1 Sv = 100 rem).

(3) As used in these regulations, the quality factors for converting absorbed dose to dose equivalent are shown below:

QUALITY FACTORS AND ABSORBED DOSE EQUIVALENCIES

|  |  | Absorbed Dose Equal |
| --- | --- | --- |
|  | Quality Factor | to a Unit Dose |
| TYPE OF RADIATION | (Q) | Equivalenta |
| X, gamma, or beta radiation and | 1 | 1 |
| high-speed electrons |  |  |
| Alpha particles, multiple-charged | 20 | 0.05 |
| particles, fission fragments and heavy |  |  |
| particles of unknown charge |  |  |
| Neutrons of unknown energy | 10 | 0.1 |
| High-energy protons | 10 | 0.1 |

a Absorbed dose in rad equal to 1 rem or the absorbed dose in gray equal to 1 Sv.

(4) If it is more convenient to measure the neutron fluence rate than to determine the neutron dose equivalent rate in sievert per hour or rem per hour, as provided in Rule 64E-5.106, F.A.C., above, 0.01 Sv (1 rem) of neutron radiation of unknown energies can, for purposes of these regulations, be assumed to result from a total fluence of 25 million neutrons per square centimeter incident upon the body. If sufficient information exists to estimate the approximate energy distribution of the neutrons, the licensee or registrant can use the fluence rate per unit dose equivalent or the appropriate Q value from the table below to convert a measured tissue dose in gray or rad to dose equivalent in sievert or rem.

MEAN QUALITY FACTORS, Q, AND FLUENCE PER UNIT DOSE 
EQUIVALENT FOR MONOENERGETIC NEUTRONS

|  | Fluence per Unit |  |  | Fluence per Unit |
| --- | --- | --- | --- | --- |
|  | Neutron | Quality | Dose Equivalentb | Dose Equivalentb |
|  | Energy | Factora | (neutrons | (neutrons |
|  | (MeV) | (Q) | cm-2 rem-1) | cm-2 Sv-1) |
| (thermal) | 2.5 × 10-8 | 2 | 980 × 106 | 980 × 108 |
|  | 1 × 10-7 | 2 | 980 × 106 | 980 × 108 |
|  | 1 × 10-6 | 2 | 810 × 106 | 810 × 108 |
|  | 1 × 10-5 | 2 | 810 × 106 | 810 × 108 |
|  | 1 × 10-4 | 2 | 840 × 106 | 840 × 108 |
|  | 1 × 10-3 | 2 | 980 × 106 | 980 × 108 |
|  | 1 × 10-2 | 2.5 | 1010 × 106 | 1010 × 108 |
|  | 1 × 10-1 | 7.5 | 170 × 106 | 170 × 108 |
|  | 5 × 10-1 | 11 | 39 × 106 | 39 × 108 |
|  | 1 | 11 | 27 × 106 | 27 × 108 |
|  | 2.5 | 9 | 29 × 106 | 29 × 108 |
|  | 5 | 8 | 23 × 106 | 23 × 108 |
|  | 7 | 7 | 24 × 106 | 24 × 108 |
|  | 10 | 6.5 | 24 × 106 | 24 × 108 |
|  | 14 | 7.5 | 17 × 106 | 17 × 108 |
|  | 20 | 8 | 16 × 106 | 16 × 108 |
|  | 40 | 7 | 14 × 106 | 14 × 108 |
|  | 60 | 5.5 | 16 × 106 | 16 × 108 |
|  | 1 × 102 | 4 | 20 × 106 | 20 × 108 |
|  | 2 × 102 | 3.5 | 19 × 106 | 19 × 108 |
|  | 3 × 102 | 3.5 | 16 × 106 | 16 × 108 |
|  | 4 × 102 | 3.5 | 14 × 106 | 14 × 108 |

aValue of quality factor at the point where the dose equivalent is maximum in a 30-cm diameter cylinder tissue-equivalent phantom.

bMonoenergetic neutrons incident normally on a 30-cm diameter cylinder tissue-equivalent phantom.

Rulemaking Authority 404.042, 404.051, 404.061 FS. Law Implemented 404.022(2) FS. History–New 1-1-94, Formerly 10D-91.113.

- 64E-5.1101 Prohibitions.

(1) No licensee shall perform wire-line service operations with a sealed source unless, prior to commencement of the operation, the licensee has a written agreement with the well operator, well owner, drilling contractor or land owner that:

(a) In the event a sealed source is lodged downhole, a reasonable effort at recovery will be made; and

(b) In the event a decision is made to abandon the sealed source downhole, the requirements of Rule 64E-5.1119, F.A.C., shall be met.

(2) No registrant shall permit above-ground testing of particle accelerators, designed for use in well-logging, which results in the production of radiation, except in areas or facilities controlled or shielded so that the requirements of Part III, as applicable, are met.

Rulemaking Authority 404.051, 404.061, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.1203.

- 64E-5.1102 Storage and Transportation Precautions.

(1) No licensee shall perform wire-line service operations with a sealed source unless, prior to commencement of the operation, the licensee has a written agreement with the well operator, well owner, drilling contractor or land owner that:

(a) In the event a sealed source is lodged downhole, a reasonable effort at recovery will be made; and

(b) In the event a decision is made to abandon the sealed source downhole, the requirements of Rule 64E-5.1119, F.A.C., shall be met.

(2) No registrant shall permit above-ground testing of particle accelerators, designed for use in well-logging, which results in the production of radiation, except in areas or facilities controlled or shielded so that the requirements of Part III, as applicable, are met.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2) FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.1204.

- 64E-5.1103 Radiation Survey Instruments.

(1) The licensee or registrant shall maintain sufficient calibrated and operable radiation survey instruments at each field station and temporary jobsite to make physical radiation surveys as required by this part and by Part III. Instrumentation shall be capable of measuring 0.1 milliroentgen (0.001 mSv) per hour through at least 50 milliroentgens (0.5 mSv) per hour.

(2) Radiation survey instruments used to establish dose rates shall be calibrated:

(a) At energies and geometries appropriate for use;

(b) At intervals not to exceed 6 months, and after each instrument servicing;

(c) Such that accuracy within plus or minus 20 percent can be demonstrated; and

(d) For linear scale instruments, at two points located approximately 1/3 and 2/3 of full-scale on each scale; for logarithmic scale instruments, at midrange of each decade, and at two points of at least one decade; and for digital instruments, at appropriate points.

(3) Records of survey instrument calibrations shall be maintained for 3 years after the calibration date for inspection by the Department.

Rulemaking Authority 404.051, 404.061, 404.081, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081(1), 404.22 FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.1205, Amended 10-8-00.

- 64E-5.1104 Leak Testing of Sealed Sources.

(1) Requirements. Each licensee using sealed sources containing radioactive material shall have the sources tested for leakage. Records of leak test results shall be kept in units of microcuries and maintained for inspection by the Department for 3 years after the leak test is performed or until transfer or disposal of the sealed source.

(2) Method of testing. Tests for leakage shall be performed only by persons specifically authorized to perform such tests by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State, or a licensing state using a leak test kit or method approved by the department, the U.S. Nuclear Regulatory Commission, an Agreement State, or a Licensing State. The test sample shall be taken from the nearest accessible point to the sealed source where contamination might accumulate. The test sample shall be analyzed for radioactive contamination, and the analysis shall be capable of detecting the presence of 0.005 microcurie (185 Bq) of radioactive material on the test sample.

(3) Test frequency.

(a) Each sealed source except an energy compensation source or ECS containing radioactive material shall be tested at intervals not to exceed 6 months. In the absence of a certificate from a transferor indicating that a test has been made within the 6 months before the transfer, the sealed source shall not be used until tested. If, for any reason, it is suspected that a sealed source may be leaking, it shall be removed from service immediately and tested for leakage as soon as practical.

(b) Each ECS that is not exempt from testing as specified in subsection 64E-5.1104(5), F.A.C., below, shall be tested at intervals not to exceed 3 years. In the absence of a certificate from a transferor that a test has been made within the 3 years before transfer, the ECS shall not be used until tested.

(4) Removal of Leaking or Contaminated Sources from service. If the test specified in subsection (3), above, reveals the presence of 0.005 microcurie (185 Bq) or more of removable radioactive material, the licensee shall remove the sealed source from service immediately and shall cause it to be decontaminated, repaired, or disposed of by a person licensed by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State, or a Licensing State to perform these functions. The licensee shall check the equipment associated with the leaking source for radioactive contamination, and if contaminated, have it decontaminated or disposed of by a person licensed by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State to perform these functions. A report describing the equipment involved, the test results, any contamination which resulted from the leaking source, and corrective action taken shall be filed with the Department within 5 days of receiving the test results.

(5) Exemptions. The following sources are exempted from the periodic leak test requirements of subsections 64E-5.1104(1) through (4), F.A.C.:

(a) Hydrogen 3 sources;

(b) Sources containing radioactive material with a half-life of 30 days or less;

(c) Sealed sources containing radioactive material in gaseous form;

(d) Sources of beta-emitting or gamma-emitting radioactive material with an activity of 100 microcuries (3.7 MBq) or less; and

(e) Sources of alpha-emitting radioactive material with an activity of 10 microcuries (0.370 MBq) or less.

Rulemaking Authority 404.022, 404.051(1), (4), 404.061, 404.081(1) FS. Law Implemented 404.022, 404.051(1), (5), 404.061(2), 404.081 FS. History–New 4-4-89, Formerly 10D-91.12051, Amended 9-28-06.

- 64E-5.1105 Quarterly Inventory.

Each licensee or registrant shall conduct a quarterly physical inventory to account for all sources of radiation. Records of inventories shall be maintained for 2 years from the date of the inventory for inspection by the Department and shall include the quantities and kinds of sources of radiation, the location where sources of radiation are assigned, the date of the inventory and the name of the individual conducting the inventory.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1), 404.081(1), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.1206.

- 64E-5.1106 Utilization Records.

Each licensee using radioactive materials shall maintain utilization records, which shall be kept available for inspection by the Department for 2 years from the date of the recorded event, showing the following information for each source of radiation:

(1) Make, model number and a serial number or a description of each source of radiation used;

(2) The identity of the well logging supervisor or field unit to whom assigned;

(3) Locations where used and dates of use; and

(4) In the case of tracer materials and radioactive markers, the utilization record shall indicate the radionuclide and activity used in a particular well.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081(1) FS. History–New 7-17-85, Formerly 10D-91.1207.

- 64E-5.1107 Design, Performance and Certification Criteria for Sealed Sources Used in Downhole Operations.

(1) A licensee can use a sealed source in well logging applications if:

(a) The sealed source is doubly encapsulated;

(b) The sealed source contains radioactive material whose chemical and physical forms are as insoluble and nondispersible as practical; and

(c) The sealed source meets requirements specified in subsection (2), (3), or (4), below.

(2) A licensee can use a sealed source manufactured on or before July 14, 1989, in well logging applications if it meets the requirements of USASI N5.10 - 1968, “Classification of Sealed Radioactive Sources”, which is herein incorporated by reference and available from the Department, or the requirements specified in subsections (3) and (4), below.

(3) A licensee can use a sealed source manufactured after July 14, 1989, in well logging applications if it meets the oil-well logging requirements specified in ANSI/HPS N43.6 1997, “Sealed Radioactive Sources Classification,” which is herein incorporated by reference and available from the Department.

(4) A licensee can use a sealed source manufactured after July 14, 1989, in well logging applications if:

The sealed source’s prototype has been tested and found to maintain its integrity after each of the following tests:

(a) Temperature. The test source is held at -40° C for 20 minutes, 600° C for 1 hour, and then subjected to a thermal shock test with a temperature drop from 600° C to 20° C within 15 seconds.

(b) Impact test. A 5 kg steel hammer 2.5 cm in diameter is dropped from a height of 1 m onto the test source.

(c) Vibration test. The test source is subjected to a vibration from 25 Hz to 500 Hz at 5 g amplitude for 30 minutes.

(d) Puncture test. A 1 gram hammer and 0.3 cm diameter pin is dropped from a height of 1 m onto the test source.

(e) Pressure test. The test source is subjected to an external pressure of 24,600 pounds per square inch absolute (1.695 x 107 pascals).

(5) The requirements of subsections (1) through (4), above, do not apply to sealed sources that contain licensed material in gaseous form.

(6) The requirements of subsections (1) through (4), above, do not apply to ECSs. ECSs shall be registered with the Department as specified in subsection 64E-5.210(14), F.A.C., the U.S. Nuclear Regulatory Commission, an Agreement State, or a Licensing State.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.071(1), 404.081(1) FS. History–New 7-17-85, Formerly 10D-91.1208, Amended 9-28-06.

- 64E-5.11071 Uranium Sinker Bars.

The licensee can use a uranium sinker bar in well logging applications only if it is legibly impressed with the words “CAUTION – RADIOACTIVE – DEPLETED URANIUM” and “NOTIFY CIVIL AUTHORITIES (OR COMPANY NAME) IF FOUND.”

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.071(1), 404.081(1) FS. History–New 9-28-06.

- 64E-5.11072 Energy Compensation Source.

The licensee can use an ECS that is contained within a logging tool or other tool components only if the ECS contains 100 microcuries (3.7 MBq) or less of licensed material.

(1) For well logging applications with a surface casing for protecting fresh water aquifers, use of the ECS is subject only to the requirements specified in Rules 64E-5.1104, 64E-5.1105, and 64E-5.1106, F.A.C., above.

(2) For well logging applications without a surface casing for protecting fresh water aquifers, use of the ECS is subject only to the requirements specified in Rules 64E-5.1101, 64E-5.1104, 64E-5.1105, 64E-5.1106, 64E-5.1119(5), and 64E-5.343 through 64E-5.349, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.071(1), 404.081(1) FS. History–New 9-28-06, Amended 2-28-08.

- 64E-5.11073 Tritium Neutron Generator Target Source.

(1) Use of a tritium neutron generator target source containing quantities not exceeding 30 curies (1,110 MBq) and in a well with a surface casing to protect fresh water aquifers is not subject to the requirements specified in Rules 64E-5.1101, 64E-5.1107, 64E-5.1119(5), and 64E-5.343 through 64E-5.349, F.A.C.

(2) Use of a tritium neutron generator target source containing more than 30 curies (1,110 MBq) or in a well without a surface casing to protect fresh water aquifers is not subject to the requirements specified in Rule 64E-5.1107, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.071(1), 404.081(1) FS. History–New 9-28-06.

- 64E-5.1108 Labeling.

(1) Each source, source holder or logging tool containing radioactive material shall bear a durable, legible and clearly visible marking or label, which has, as a minimum, the standard radiation caution symbol as described and illustrated in Rule 64E-5.322, F.A.C., without the conventional color requirement, and the following wording:

DANGER (OR “CAUTION”)

RADIOACTIVE

This label shall be on the smallest component transported as a separate piece of equipment.

(2) Each transport container shall have permanently attached to it a durable, legible and clearly visible label which has, as a minimum, the standard radiation caution symbol as described and illustrated in Rule 64E-5.322, F.A.C., and the following wording:

DANGER (OR “CAUTION”)

RADIOACTIVE.

NOTIFY CIVIL AUTHORITIES IF FOUND

Rulemaking Authority 404.051, 404.061, 404.081, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081(1), 404.20(1) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.1209.

- 64E-5.1109 Inspection and Maintenance.

(1) Each licensee possessing radioactive material shall conduct, at intervals not to exceed 6 months, a program of inspection and maintenance of source holders, logging tools, source handling tools, storage containers, transport containers and injection tools to assure proper labeling and physical condition. Records of inspection and maintenance shall be maintained for a period of 2 years for inspection by the Department.

(2) If any inspection conducted pursuant to this section reveals damage to labeling or components critical to radiation safety, the device shall be removed from service until repairs have been made.

(3) The repair, opening or modification of any sealed source device shall be performed only by persons specifically authorized to do so by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.071(1), 404.081(1) FS. History–New 7-17-85, Formerly 10D-91.1210.

- 64E-5.1110 Training Requirements.

(1) No licensee or registrant shall permit any individual to act as a logging supervisor as defined in this part until such individual has:

(a) Received, in a course taught by an individual who has been licensed by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State, instruction in the subjects outlined in this part and demonstrated an understanding thereof;

(b) Read and received instruction in the regulations contained in this part and the applicable sections of Parts I, III and IX, or their equivalent, conditions of the appropriate license or certificate of registration, and the licensee’s or registrant’s operating and emergency procedures, and demonstrated an understanding thereof; and

(c) Demonstrated competence to use sources of radiation, related handling tools and radiation survey instruments which will be used on the job.

(2) No licensee or registrant shall permit any individual to assist in the handling of sources of radiation until such individual has:

(a) Read or received instruction in the licensee’s or registrant’s operating and emergency procedures and demonstrated an understanding thereof; and

(b) Demonstrated competence to use, under the personal supervision of the logging supervisor, the sources of radiation, related handling tools and radiation survey instruments which will be used on the job.

(3) The licensee or registrant shall maintain employee training records for inspection by the Department for 2 years following termination of employment.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1), 404.081(1), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.1211.

- 64E-5.1111 Operating and Emergency Procedures.

The licensee’s or registrant’s operating and emergency procedures shall include appropriate instructions in at least the following:

(1) Handling and use of sources of radiation to be employed so that no individual is likely to be exposed to radiation doses in excess of the standards established in Part III;

(2) Methods and occasions for conducting radiation surveys;

(3) Methods and occasions for locking and securing sources of radiation;

(4) Personnel monitoring and the use of personnel monitoring equipment;

(5) As applicable, the transportation of radioactive sources to temporary job sites and field stations, including the packaging and placing of such sources in vehicles, placarding of vehicles and securing the sources during transportation;

(6) Minimizing exposure of individuals in the event of an accident;

(7) Procedure for notifying proper personnel in the event of an accident;

(8) Maintenance of records;

(9) As applicable, inspection and maintenance of source holders, logging tools, source handling tools, storage containers, transport containers and injection tools;

(10) As applicable, procedures to be followed in the event a sealed source is lodged downhole; and

(11) An applicable, procedures to be used for picking up, receiving and opening packages containing radioactive material.

Rulemaking Authority 404.051, 404.061, 404.081, 404.20, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081(1), 404.20(1), 404.22 FS. History–New 7-17-85, Formerly 10D-91.1212.

- 64E-5.1112 Personnel Monitoring.

No licensee or registrant shall permit any individual to act as a logging supervisor or to assist in the use of sources of radiation unless such individual wears a film badge, optically stimulated luminescent device (OSLD), or a thermoluminescent dosimeter (TLD) that is processed and evaluated by an accredited NVLAP processor. Each film badge, OSLD, or TLD shall be assigned to and worn by only one individual. Film badges shall be replaced at least monthly and OSLDs and TLDs shall be replaced at least quarterly. Each film badge, OSLD, and TLD shall be processed promptly after replacement. The licensee shall retain records of personnel dosimeters and bioassay results until the Department terminates each pertinent license or registration requiring the records.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081(1), (2) FS. History–New 7-17-85, Amended 5-15-96, Formerly 10D-91.1213, Amended 10-8-00, 9-28-06.

- 64E-5.1113 Security.

During each logging or tracer application, the logging supervisor or other designated employee shall maintain direct surveillance of the operation to protect against unauthorized or unnecessary entry into a restricted area, as defined in Rule 64E-5.101, F.A.C.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.031, 404.051(1), (4), 404.061(2) FS. History–New 7-17-85, Amended 8-29-94, Formerly 10D-91.1214.

- 64E-5.1114 Handling Tools.

The licensee shall provide and require the use of tools that will assure remote handling of sealed sources except for low-activity calibration sources that result in a gamma exposure rate at contact of less than 100 milliroentgens (2.58 × 10-5 uC per kg) per hour.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2) FS. History–New 7-17-85, Formerly 10D-91.1215.

- 64E-5.1115 Subsurface Tracer Studies.

(1) Protective gloves and other appropriate protective clothing and equipment shall be used by all personnel handling radioactive material. Precautions shall be taken to avoid ingestion or inhalation of radioactive material.

(2) No licensee shall intentionally inject radioactive material into any fresh water aquifers unless the Department of Health and the Department of Environmental Protection determine that such injection will not endanger the public health, safety and welfare.

(3) No licensee shall inject radioactive material into any well unless it can be demonstrated to the Department that the procedure will not result in any liquids or gases distributed to the public exceeding the following criteria:

(a) For gases, the air concentration in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table II, Column 1, shall apply.

(b) For liquids, the water concentration values in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, Table II, Column 2, shall apply.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.031, 404.051(1), (4), 404.061(2) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.1216, Amended 12-26-13.

- 64E-5.1116 Radiation Surveys.

(1) Radiation surveys and personnel exposure calculations shall be made and recorded for each area where radioactive materials are stored.

(2) Radiation surveys and personnel exposure calculations shall be made and recorded for the radiation levels in occupied positions and on the exterior of each vehicle used to transport radioactive material. Such surveys and calculations shall include each source of radiation or combination of sources to be transported in the vehicle.

(3) After removal of the sealed source from the logging tool and before departing the jobsite, a survey meter shall be used to assure that the logging tool is free of contamination.

(4) Radiation surveys shall be made and recorded at the jobsite or well-head for each tracer operation, except those using tritium, carbon 14 and sulfur 35. These surveys shall include measurements of radiation levels before and after the operation. If radiation levels, post operation, exceed twice background, the area shall be decontaminated or restricted until radiation levels reach twice background.

(5) Records required pursuant to this section shall include the dates, the identification of individuals making the survey, the identification of survey instruments used and an exact description of the location of the survey. Records of these surveys shall be maintained for inspection by the Department for 2 years after completion of the survey.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.071(1), 404.081(1), 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.1217.

- 64E-5.1117 Documents and Records Required at Field Stations.

Each licensee or registrant shall maintain, for inspection by the Department, the following documents and records for the specific devices and sources used at the field station:

(1) Appropriate license or certificate of registration;

(2) Operating and emergency procedures;

(3) A copy of these regulations;

(4) Records of the latest survey instrument calibrations pursuant to Rule 64E-5.1103 and Part III, F.A.C.;

(5) Records of the latest leak test results pursuant to license conditions;

(6) Quarterly inventories required pursuant to Rule 64E-5.1105, F.A.C.;

(7) Utilization records required pursuant to Rule 64E-5.1106, F.A.C.;

(8) Records of inspection and maintenance required pursuant to Rule 64E-5.1109, F.A.C.; and

(9) Survey records required pursuant to Rule 64E-5.1116, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1), 404.081(1) FS. History–New 7-17-85, Formerly 10D-91.1218.

- 64E-5.1118 Temporary Jobsites.

Each licensee or registrant conducting operations at a temporary jobsite, which is a location to which radioactive materials have been dispatched to perform wireline service operations or subsurface tracer studies, shall have the following documents and records available at that site for inspection by the Department:

(1) Operating and emergency procedures;

(2) Survey records required pursuant to Rule 64E-5.1116, F.A.C., for the period of operation at the site;

(3) Evidence of current calibration for the radiation survey instruments in use at the site; and

(4) When operating in the State under reciprocity, a copy of the appropriate license, certificate of registration or equivalent documents.

Rulemaking Authority 404.051, 404.061, 404.081, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (11), 404.061(2), 404.081(1), 404.22 FS. History–New 7-17-85, Amended 5-15-96, Formerly 10D-91.1219.

- 64E-5.1119 Notification of Incidents, Abandonment and Lost Sources.

(1) Notification shall be made of radiation incidents and radioactive sources lost in other than downhole logging operations in accordance with appropriate provisions of Part III.

(2) Whenever a sealed source or device containing radioactive material is lodged downhole, the licensee shall:

(a) Monitor at the surface for the presence of radioactive contamination with a radiation survey instrument or logging tool during logging tool recovery operations; and

(b) Notify the Department immediately by telephone or telegraph if radioactive contamination is detected at the surface or if the source appears to be damaged.

(3) When it becomes apparent that efforts to recover the radioactive source will not be successful, the licensee shall:

(a) Advise the well-operator and the Department of an appropriate method of abandonment, which shall include:

1. The immobilization and sealing in place of the radioactive source with a cement plug,

2. The setting of a whipstock or other deflection device, and

3. The mounting of a permanent identification plaque, at the surface of the well, containing the appropriate information required by this section;

(b) Notify the Department by telephone of the circumstances  that resulted in the inability to retrieve the source and obtain the Department’s approval to implement abandonment procedures or notify the Department that the licensee implemented abandonment before receiving Department approval because the licensee believed there was an immediate threat to public health and safety; and

(c) File a written report with the Department within 30 days of the abandonment, setting forth the following information:

1. Date of occurrence and a brief description of attempts to recover the source;

2. A description of the radioactive source involved, including radionuclide, quantity and chemical and physical form;

3. Surface location and identification of well;

4. Results of efforts to immobilize and set the source in place;

5. Depth of the lodged radioactive source;

6. Depth of the top of the cement plug;

7. Depth of the well;

8. Information contained on the permanent identification plaque;

9. The immediate threat to public health and safety that justified abandonment before Department approval as specified in paragraph (3)(b), above; and

(d) Develop and implement a means to prevent inadvertent intrusion on the source unless the source is not accessible to any subsequent drilling operations.

(4) Whenever a sealed source containing radioactive material is abandoned downhole, the licensee shall provide a permanent plaque, as described below, for posting the well or well-bore at the surface of the well unless the mounting of the plaque is not practical. The size of the plaque shall be at least 7 inches (17 cm) square and 1/8 inch (3 mm) thick. This plaque shall:

(a) Be constructed of long-lasting material, such as stainless steel, brass, bronze, or monel, and

(b) Contain the following information engraved on its face:

1. The word “CAUTION”;

2. The radiation symbol without the conventional color requirement;

3. The date of abandonment;

4. The name of the well operator or well owner;

5. The well name and well identification numbers or other designation;

6. The sealed sources by radionuclide and quantity of activity;

7. The source depth and the depth to the top of the plug; and

8. An appropriate warning, depending on the specific circumstances of each abandonment which may include:

a. “Do not drill below plug-back depth”;

b. “Do not enlarge casing”; or

c. “Do not reenter the hole”, followed by the words, “before contacting the Department of Health”.

(5) The licensee shall immediately notify the Department by telephone or telegraph, and subsequently by confirming letter, if the licensee knows or has reason to believe that radioactive material has been lost in or to an underground potable water source. Such notice shall designate the well location and shall describe the magnitude and extent of loss of radioactive material, assess the consequences of such loss and explain efforts planned or being taken to mitigate these consequences.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.081(1) FS. History–New 7-17-85, Formerly 10D-91.1220, Amended 9-28-06.

- 64E-5.1120 Subjects to Be Included in Training Courses for Logging Supervisors.

The following subjects: must be included in training courses for logging supervisors.

(1) Fundamentals of radiation safety, including:

(a) Characteristics of radiation;

(b) Units of radiation dose and, if appropriate, quantity of radioactivity;

(c) Significance of radiation dose, including:

1. Radiation protection standards; and

2. Biological effects of radiation dose;

(d) Levels of radiation from sources of radiation; and

(e) Methods of minimizing radiation dose, including:

1. Working time;

2. Working distances; and

3. Shielding.

(2) Radiation detection instrumentation to be used, including:

(a) Use of radiation survey instruments, including operation, calibration and limitations;

(b) Survey techniques; and

(c) Use of personnel monitoring equipment.

(3) Equipment to be used, including:

(a) Handling equipment, if appropriate;

(b) Sources of radiation;

(c) Storage precautions, if appropriate, and control of equipment; and

(d) Operation and control of equipment.

(4) The requirements of these regulations.

(5) The licensee’s or registrant’s written operating and emergency procedures.

(6) The licensee’s or registrant’s record keeping procedures.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.081(1), 404.22 FS. History–New 7-17-85, Formerly 10D-91.1221.

- 64E-5.1201 Radon Requirements and Enforcement.

(1) This part establishes rules, requirements and procedures for the certification of radon measurement and mitigation specialists, technicians and businesses, and the protocols for mandatory measurements.

(2) The General Statement of Policy and Procedures for Radon Enforcement Actions January 1993, which is available from the department and which is incorporated herein by reference, will be used to determine enforcement actions.

Rulemaking Authority 404.042, 404.051, 404.056 FS. Law Implemented 404.051(1), (3), (7), 404.056 FS. History–New 1-3-89, Amended 12-9-96, Formerly 10D-91.1301.

- 64E-5.1202 Definitions.

The following definitions shall be applicable to this part:

(1) “Business day” means any day of the year, exclusive of Saturdays, Sundays, and State of Florida holidays.

(2) “Diagnostic tests” means tests performed or procedures used to determine appropriate mitigation methods for a building.

(3) “Certified person” or “certified individual” means a certified radon measurement specialist, certified radon measurement business, certified radon measurement technician, certified radon mitigation business, certified radon mitigation specialist or certified radon mitigation technician as defined by this part.

(4) “Certified radon measurement business” means a commercial business which provides on-site measurement services for radon or radon progeny and which meets the requirements for certification provided by this part. For the purpose of this section, an analytical service provider which does not perform on-site measurement services directly to the public is not considered a radon measurement business.

(5) “Certified radon measurement specialist” means an individual who performs on-site radon or radon progeny measurements; provides professional advice on such radon or radon progeny measurements, health risks, radon-related exposure, radon entry routes, and other radon-related activities; and meets the requirements for certification provided by this part.

(6) “Certified radon measurement technician” means an individual who performs on-site radon or radon progeny measurement activities, provides information on measurement results, and meets the requirements for certification provided by this part. Such an individual can perform these activities only under the supervision of a certified radon measurement specialist.

(7) “Certified radon mitigation business” means a commercial business which designs or installs systems in existing buildings to mitigate radon or radon progeny contamination and which meets the requirements for certification provided by this part.

(8) “Certified radon mitigation specialist” means an individual who performs or evaluates diagnostic tests to determine appropriate radon or radon progeny mitigation strategies for a building, designs and installs mitigation systems, and meets the requirements for certification provided by this part.

(9) “Certified radon mitigation technician” means an individual who installs or supervises the installation of radon or radon progeny mitigation techniques on existing buildings and meets the requirements for certification provided by this part. Such an individual can perform these activities only under the supervision of a certified radon mitigation specialist.

(10) “Mitigate” means to apply materials or install systems and materials to reduce radon or radon progeny concentrations in the indoor atmosphere or to prevent entry of radon or radon progeny into the indoor atmosphere.

(11) “Picocurie per liter (pCi/l)” means 2.22 disintegrations per minute of radioactive material per liter. One picocurie is equivalent to 10-12 Curies.

(12) “Radon” means the radioactive noble gas radon 222.

(13) “Radon progeny” means the short-lived radionuclides formed as a result of the decay of radon 222, including polonium 218, bismuth 214, lead 214, and polonium 214.

(14) “Working level (WL)” means the concentration of short-lived radon progeny that will result in 130,000 million electron volts of potential alpha particle energy per liter of air. Working level is a measure of radon progeny concentration in air.

Rulemaking Authority 404.042, 404.051, 404.056 FS. Law Implemented 404.051(1), (3), (4), 404.056 FS. History–New 1-3-89, Amended 5-21-90, 10-24-90, 1-18-94, 12-9-96, Formerly 10D-91.1303.

- 64E-5.1203 General Provisions.

(1) Beginning January 1, 1989, no person may test for or mitigate the presence of radon in Florida for a fee or other remuneration unless such person has been certified as provided by this part. These regulations in no way exempt any person from other state and local occupational licensure requirements.

(2) The certification requirements of this part shall not apply to a commercial business that subcontracts radon measurements to a certified radon business when the complete radon report from the certified radon business is provided to the client and interpretations of the results are provided by the certified business.

(3) No certification shall be approved unless the applicant demonstrates to the department that the following conditions are met:

(a) The applicant has not been found to be in violation of Chapter 404, Florida Statutes, or this part, and has not been decertified;

(b) The applicant has filed an accurate and complete application with the application fee describing compliance with the relevant certification requirements.

(4) Requirements for continued certification shall include the following conditions:

(a) The certified person shall conduct his activities as described in the approved certification and shall remain in compliance with Chapter 404, F.S., and this part.

(b) The certified person shall allow authorized representatives of the department to have access during normal business hours to his facilities, offices and files for inspection and examination of radon-related records, measurement procedures and mitigation system design and installations. The certified person shall also allow authorized representatives of the department to inspect the on site radon measurement or mitigation activities.

(c) The following changes shall be reported in writing to the department within 10 business days of the change: changes to the name or address of a certified individual or business; changes in the information provided with the original business application, including changes in certified personnel; changes in analytical service providers used; changes in USEPA proficiency status; and changes to the measurement or mitigation methods used. New measurement methods shall comply with the Quality Assurance Plan requirements specified in subsection 64E-5.1206(5), F.A.C.

(5) Prior to performing radon measurement or radon mitigation services, including sample collection, analysis, or interpretation of measurement results a certified individual shall own, be employed by, or be retained as a consultant to a certified radon measurement or certified radon mitigation business.

(6) A certification will be valid for 1 year following the date of issuance. If renewal is requested more than 90 days after an expired certification, a new application must be submitted for recertification. After 1 year, previously certified applicants must attend a department-approved training course. A duplicate certificate can be obtained by written request to the department and payment of a five dollar fee.

(7)(a) A request for annual certification renewal shall be submitted with a nonrefundable renewal fee. For renewal requests received prior to expiration of the certification, the fee schedule set forth in this part will apply. For renewal requests received after certification has expired, the fee schedule for application and certification will apply. A certification renewal shall be issued or denied according to the criteria set forth in this part.

(b) Annual certification renewal fee schedule:

1. Radon Measurement or Mitigation Business

Renewal request received on or before June 30, 2009:						$200

Renewal request received between July 1, 2009 and June 30, 2010:				$275

Renewal request received between July 1, 2010 and June 30, 2011:				$350

Renewal request received between July 1, 2011 and June 30, 2012:				$425

Renewal request received on or after July 1, 2012:						$500

2. Radon Measurement or Mitigation Specialist

Renewal request received on or before June 30, 2009:						$200

Renewal request received between July 1, 2009 and June 30, 2010:				$250

Renewal request received between July 1, 2010 and June 30, 2011:				$300

Renewal request received between July 1, 2011 and June 30, 2012:				$350

Renewal request received on or after July 1, 2012:						$400

3. Radon Measurement or Mitigation Technician

Renewal request received on or before June 30, 2009:						$200

Renewal request received between July 1, 2009 and June 30, 2010:				$225

Renewal request received between July 1, 2010 and June 30, 2011:				$250

Renewal request received between July 1, 2011 and June 30, 2012:				$275

Renewal request received on or after July 1, 2012:						$300

(8) A certified individual in good standing who becomes an active duty member of the Armed Forces of the United States prior to the expiration of the renewal period shall remain in good standing while on active duty and for a period of 6 months after discharge from active duty. No radon measurement or mitigation activity for which certification is required shall be conducted after the expiration of the term of certification.

(9) A certified individual in good standing who is absent from the state because of their spouse’s active duty with the Armed Forces of the United States prior to the expiration of the renewal period shall remain in good standing while their spouse is on active duty and for a period of 6 months after their spouse’s discharge from active duty. No radon measurement or mitigation activity for which certification is required shall be conducted after the expiration of the term of certification.

(10) All applications should clearly label any information considered trade secrets and must segregate such information from other information as required by Section 119.165, F.S.

(11) The department will approve training courses for certification based upon evaluation of course materials and content. These courses shall meet the following requirements.

(a) The course subjects shall be applicable to Florida, shall enable the students to meet training objectives provided by the department and shall include at least the following:

1. Radon Measurement Specialist Training.

a. Two hours of instruction on radioactivity, radon, and radon’s health effects.

b. One and one-half hours of instruction on Florida’s radon legislation and administrative rules and federal radon legislation and proficiency programs.

c. Two and one-half hours of instruction on radon entry and behavior.

d. One hour of instruction on worker health and safety.

e. Seven hours of instruction on radon measurement devices, techniques and protocols.

f. One and one-quarter hours of instruction on building investigation and diagnostics and mitigation strategies and techniques applicable to Florida.

g. One-half hour of instruction on radon-resistant building techniques for new construction applicable to Florida.

h. One hour of instruction on measurement and mitigation techniques for radon in domestic water.

2. Radon Measurement Technician Training.

a. One and one-quarter hours of instruction on radioactivity, radon, and radon’s health effects.

b. One and one-half hours of instruction on Florida’s radon legislation and administrative rules and federal radon legislation and proficiency programs.

c. One and three-quarters hours of instruction on radon entry and behavior.

d. One hour of instruction on worker health and safety.

e. Five hours of instruction on radon measurement devices, techniques and protocols, including 1 hour of experience deploying, retrieving and reporting results of various measurement devices.

f. One-half hour of instruction on mitigation strategies and techniques applicable to Florida.

g. One-half hour of instruction on radon-resistant building techniques for new construction applicable to Florida.

h. One-half hour of instruction on measurement and mitigation techniques for radon in domestic water.

3. Radon Mitigation Specialist Training.

a. Two hours of instruction on radioactivity, radon, and radon’s health effects.

b. One and one-half hours of instruction on Florida’s radon legislation and administrative rules and federal radon legislation and proficiency programs.

c. Two and one-half hours of instruction on radon entry and behavior.

d. One hour of instruction on worker health and safety.

e. Three-quarters hour of instruction on radon measurement devices, techniques and protocols.

f. One-half hour of instruction on radon-resistant building techniques for new construction applicable to Florida.

g. One hour of instruction on measurement and mitigation techniques for radon in domestic water.

h. Three and one-half hours of instruction on building investigation and diagnostics and mitigation strategies and techniques applicable to Florida.

i. Six and one-half hours of instruction on mitigation system theory, design and installation applicable to Florida.

4. Radon Mitigation Technician Training.

a. One and one-quarter hours of instruction on radioactivity, radon, and radon’s health effects.

b. One and one-half hours of instruction on Florida’s radon legislation and administrative rules and federal radon legislation and proficiency programs.

c. One and three-quarters hours of instruction on radon entry and behavior.

d. One hour of instruction on worker health and safety.

e. Three-quarters hour of instruction on radon measurement devices, techniques and protocols.

f. One-half hour of instruction on radon-resistant building techniques for new construction applicable to Florida.

g. Three-quarters hour of instruction on measurement and mitigation techniques for radon in domestic water.

h. Three hours of instruction on building investigation and diagnostics and mitigation strategies and techniques applicable to Florida.

i. Five hours of instruction on mitigation system theory and installation applicable to Florida.

(b) Participants shall be given an opportunity to evaluate in writing the learning experiences, instructors, instructional methods, facilities and resources used for the course. Results of these evaluations shall be available for inspection by the department for 3 years after completion of the course.

(c) The subject matter shall be kept current and represent accepted protocols and techniques.

(d) The course provider shall furnish each participant who has attended the entire course a written verification of completion which includes:

1. Date of course;

2. Signature and name of provider;

3. Course title; and

4. Name of participant.

(e) Course providers shall send to the department a roster of participants who have attended the entire course within 30 days after each course on DH Form 1812 Radon Certification Training Roster, Apr 90 available from the department and herein incorporated by reference. Providers shall maintain security of attendance records. Course providers shall maintain records of individual courses for 3 years after completion of the course for inspection by the department.

(f) A course provider seeking approval to conduct a radon certification training course shall make application on DH Form 1813 Application for Radon Certification Course Provider, Apr 90, available from the department and herein incorporated by reference.

(g) The course provider also shall submit to the department a copy of the proposed curriculum, all training materials, and information concerning the proposed instructors.

(h) The course provider shall allow authorized representatives of the department to observe and evaluate training sessions, instructors and course material without charge.

(i) Any proposed changes in course materials or instructors shall be reported in writing to the department at least 15 business days before implementation of the change.

(j) Approval of a course shall be valid for a period of 24 months. Approval shall be withdrawn if the department determines that the course approval requirements are not met.

Rulemaking Authority 404.051, 404.056 FS. Law Implemented 404.051, 404.056, 404.071 FS. History–New 1-3-89, Amended 5-21-90, 10-24-90, 1-18-94, 12-9-96, Formerly 10D-91.1304, Amended 3-16-09.

- 64E-5.1204 Certification Requirements for Radon Measurement Specialists and Technicians.

(1) The following is required for certification as a radon measurement specialist and measurement technician:

(a) The individual shall submit an application on DH Form 1751 Application for Certification as Radon Specialist or Technician, Jan. 93, which is herein incorporated by reference and which is available from the department. The application shall include the nonrefundable application and certification fee according to the fee schedule set forth in this part and documentation of the applicant’s relevant education and experience, as specified on DH Form 1751 Application for Radon Certification as Radon Specialist or Technician, Jan. 93.

(b) Individuals who have never been certified must have attended within the last 2 years of the training examination date a course on radon measurement that has been approved by the department. The technician applicant also shall have received a score of 65 percent or greater on the measurement technician training examination to be administered by the department at least three times per year. The specialist applicant also shall have received a score of 75 percent or greater on the measurement specialist training examination to be administered by the department at least three times each year. If a specialist applicant receives a score of less than 75 percent but greater than or equal to 65 percent, he can be certified as a measurement technician by making a written request to the department for a change from specialist to technician. If an applicant wishes to be reexamined for specialist or technician certification, he must reapply by submitting a completed application with experience history and a $200 nonrefundable application and certification fee. On approval of the new application, the applicant can take the next scheduled department exam. Applicants can review their exam results during normal working hours at the department’s office in Tallahassee by requesting an appointment at least 5 business days prior to the review.

(c) Application and certification fee schedule:

1. Radon Measurement Specialist

Application received on or before June 30, 2009:						$200

Application received between July 1, 2009 and June 30, 2010:					$275

Application received between July 1, 2010 and June 30, 2011:					$325

Application received between July 1, 2011 and June 30, 2012:					$375

Application received on or after July 1, 2012:						$425

2. Radon Measurement Technician

Application received on or before June 30, 2009:						$200

Application received between July 1, 2009 and June 30, 2010:					$250

Application received between July 1, 2010 and June 30, 2011:					$275

Application received between July 1, 2011 and June 30, 2012:					$300

Application received on or after July 1, 2012:						$325

(2) In addition to the above, a measurement specialist applicant shall possess 4 years of relevant radiation or radiological safety professional or work experience. Relevant post-secondary education may be substituted on a year for year basis for 3 of the 4 years of relevant professional or work experience. Relevant post-secondary education includes curriculum in a natural science, engineering, radiation, radiological safety or health physics, mathematics or statistics. An academic year shall be 30 semester hours at a C- grade level or better, of which at least 10 hours shall be in the above subject areas, or 45 quarter hours at a C- grade level or better, of which at least 15 hours shall be in the above subject areas. For other relevant radiation education such as military or technical school training, 15 classroom hours of instruction shall be equivalent to 1 credit hour of college level course work. Full time training and advanced military courses will be evaluated by the department to determine their academic equivalent. Relevant professional or work experience includes radiation instrumentation, radiation physics, radiation protection, radiation biology, and radiation risk communication. Professional or work experience also shall include employment in such occupations as nuclear medicine technologist, radiation therapy technologist, radiographer, health physicist, military health physics specialist, nuclear navy engineering lab technician, health physics technician, hospital radiation safety officer, or occupations requiring equivalent experience.

Rulemaking Authority 404.051, 404.056, 404.071 FS. Law Implemented 404.051, 404.056, 404.071  FS. History–New 1-3-89, Amended 5-21-90, 10-24-90, 1-18-94, 12-9-96, Formerly 10D-91.1305, Amended 3-16-09.

- 64E-5.1205 Certification Requirements for Radon Mitigation Specialists and Technicians.

(1) The following is required for certification as a radon mitigation specialist and technician:

(a) The individual shall submit an application on DH Form 1751 Application for Certification as Radon Specialist or Technician, Jan. 93, a nonrefundable application and certification fee according to the fee schedule set forth in this part, and documentation of the applicant’s relevant education and experience, as specified on DH Form 1751 Application for Radon Certification as Radon Specialist or Technician, Jan. 93.

(b) Individuals who have never been certified must have attended within 2 years of the training examination date a course on radon diagnostics and mitigation that has been approved by the department. The technician applicant also shall have received a score of 60 percent or greater on the mitigation technician training examination to be administered by the department at least three times per year. The specialist applicant also shall have received a score of 70 percent or greater on the mitigation specialist training examination to be administered by the department at least three times each year. If a specialist applicant receives a score of less than 70 percent but greater than or equal to 60 percent, he can be certified as a mitigation technician by making a written request to the department for a change from specialist to technician. If an applicant wishes to be reexamined for specialist or technician certification, he must reapply by submitting a completed application with experience history and a nonrefundable $200 application and certification fee. After approval of the new application, the applicant can take the next scheduled department exam. Applicants can review their exam results during normal working hours at the department’s office in Tallahassee by requesting an appointment at least 5 business days prior to the review.

(c) Application and certification fee schedule:

1. Radon Mitigation Specialist

Application received on or before June 30, 2009:						$200

Application received between July 1, 2009 and June 30, 2010: 					$275

Application received between July 1, 2010 and June 30, 2011:					$325

Application received between July 1, 2011 and June 30, 2012:					$375

Application received on or after July 1, 2012: 						$425

2. Radon Mitigation Technician

Application received on or before June 30, 2009: 						$200

Application received between July 1, 2009 and June 30, 2010: 					$250

Application received between July 1, 2010 and June 30, 2011: 					$275

Application received between July 1, 2011 and June 30, 2012: 					$300

Application received on or after July 1, 2012: 						$325

(2) In addition to the above, a mitigation specialist applicant shall possess a valid residential, building or general contractor’s license issued by the State of Florida, or possess 4 years of relevant professional or work experience. Relevant post-secondary education may be substituted on a year for year basis for 3 of the 4 years of relevant professional or work experience. Relevant post-secondary education includes curriculum in architecture, engineering, building construction, or the physical sciences. Relevant professional or work experience includes the design or construction of buildings, or associated heating, ventilation, and air conditioning systems, or closely related activities.

Rulemaking Authority 404.051, 404.056, 404.071 FS. Law Implemented 404.051, 404.056, 404.071 FS. History–New 1-3-89, Amended 5-21-90, 10-24-90, 1-18-94, 12-9-96, Formerly 10D-91.1306, Amended 3-16-09.

- 64E-5.1206 Certification Requirements for Radon Measurement Businesses.

The following qualifications are required for certification as a radon measurement business:

(1) Submission of an application for certification on DH Form 1749 Application for Certification as a Radon Business, Jan. 93, which is herein incorporated by reference and which is available from the department, and a nonrefundable application and certification fee from the following schedule:

Application received on or before June 30, 2009:						$200

Application received between July 1, 2009 and June 30, 2010:					$300

Application received between July 1, 2010 and June 30, 2011:					$375

Application received between July 1, 2011 and June 30, 2012:					$450

Application received on or after July 1, 2012:						$525

(2) Identification of all radon and radon progeny measurement methods or related services to be offered, including the purpose of each measurement service, the type of measurement equipment and a summary of the procedure to be used in performing the service. All measurements reported to clients shall be performed with devices that meet measurement proficiency requirements of the USEPA. U.S. Environmental Protection Agency Indoor Radon and Radon Decay Product Measurement Device Protocols, Publication No. EPA 402-R-92-004 and Protocols for Radon and Radon Decay Product Measurements in Homes, Publication No. EPA 402-R-92-003, which are herein incorporated by reference and which are available from the department, shall be followed for all measurements reported to clients with the exception of nonresidential measurements performed as specified in subpart B of this rule.

(3) Identification of certified radon measurement specialists and certified radon measurement technicians who will perform measurements for the applicant. An applicant shall maintain on staff or retain as a consultant a certified radon measurement specialist. All radon or radon progeny measurements will be performed only by certified radon measurement specialists or certified radon measurement technicians. This shall include the initial placement and final retrieval of all measurement devices. The certified radon measurement specialist shall monitor the applicant’s measurement activities and shall review, approve, and sign monthly reports to the department, assess quality assurance and quality control measures, evaluate operating procedures, and ensure compliance with state and federal regulations. The certified radon measurement specialist shall be present when requested by the department during scheduled visits by the department and shall physically observe and evaluate each certified radon measurement technician in the performance of his measurement duties at least once each calendar quarter to insure adequate supervision. A record of this evaluation shall be signed by the specialist and maintained on file by the business. If the measurement reports submitted to meet the requirements of paragraph 64E-5.1206(12)(a) and subsection 64E-5.1208(3), F.A.C., show that no radon or radon progeny measurements were performed during an entire calendar quarter by any of the certified radon measurement technicians working for a certified radon measurement business, the visit, observation and evaluation by the specialist are not required for that quarter. The quarterly visit, observations and evaluations by the specialist must be resumed within the same calendar quarter in which measurement activities are resumed. The interval between visits, observations and evaluations by the specialist shall not exceed 1 year.

(4) Identification of the proficient analytical service providers which will process devices for all measurement methods used by the business. A radon measurement business must maintain proficiency as an analytical service provider with the USEPA with every measurement method for which it analyzes devices or obtains results directly from measurement instruments.

(5) Disclosure and adherence to a plan of quality assurance and quality control for each measurement method used by the applicant to assure the reliability and validity of radon and radon progeny measurements. This plan shall include provisions to monitor the accuracy and precision of measurements and can include the use of periodic calibrations of instruments, and blank, duplicate and spiked measurements.

(6) Disclosure of all sample reporting forms mailed to clients, including any guidance provided concerning the need for further measurement or mitigation.

(7) Disclosure of copies of current publications and advertisements of radon-related services made by the applicant.

(8) Development, disclosure and adherence to a health and safety program to limit employees’ exposure to radon progeny during the course of employment. Such a program shall include measures to keep each employee’s exposure, which shall not exceed 4 working level months per year, as low as reasonably achievable. The certified radon measurement business shall maintain records showing the radiation exposures of all individuals for whom personnel monitoring is required. Such records shall be kept on DH Form 1834 Working Level Month Explosure Form, Aug. 90, which is herein incorporated by reference and which is available from the department, or on clear and legible records containing all the information required by that form. Records shall be maintained in working level months showing the results of exposure during radon measurement activities and shall be preserved for 20 years. The discontinuance or curtailment of activities does not relieve the certified business of responsibility for retaining all radiation exposure records.

(9) Maintenance of the following records for 5 years:

(a) Records of all radon tests performed;

(b) Records of instrument calibrations and quality control;

(c) Records of participation in a proficiency program;

(d) Copies of certification for radon measurement specialists and radon measurement technicians employed or used as consultants.

(e) Records of specialist site visits and observations and evaluations of technicians.

(10) If a certified radon measurement business loses its certified radon measurement specialist, the certified radon measurement business shall notify the department in writing within 5 business days. The certified radon measurement business shall obtain another certified radon measurement specialist within 30 days of the loss of the certified radon measurement specialist.

(11) Identification of all branch office locations and the certified specialists and technicians at branch locations who will perform measurements for the certified business. Branch offices shall identify the name and the certification number of the certified business on correspondence and reports to clients and shall accept payments in the name of the certified business. The branch office shall be considered a part of the certified business and make reports and follow procedures approved for the certified business.

(12) A certified radon measurement business shall provide the following reports.

(a) A certified radon measurement business shall submit measurement reports to the department by the 15th of each month. Reports shall include building information, the name of the person performing the measurements and the results of all available radon and radon progeny measurements taken during the previous month which are not otherwise reported as specified in subpart B of this rule. If no measurements were taken during a month, the report shall indicate none were taken. Data shall be submitted on DH Form 1750 Monthly Report for Certified Radon Businesses, Residential Measurement Report, Mar 94, or DH Form 1833 Monthly Report for Certified Radon Businesses, Non Residental Measurement Report, Mar 94, which are herein incorporated by reference and which are available from the department.

(b) The certified radon measurement business shall report measurement results for radon and radon progeny to the client. Radon results shall be reported in picocuries per liter. Radon progeny results shall be reported in working levels. The report provided to the client shall include the name of the certified business, the business certification number, the name and certification number of the individual who performed the measurements, the type of device used, the date and time of the start and end of the measurement, and the following statements: This notice is provided to you by an organization or individual certified by the Florida Department of Health to perform radon or radon progeny measurements or radon mitigation services. Any questions, comments, or complaints regarding the persons performing these measurement or mitigation services may be directed to the Florida Department of Health, Bureau of Facility Programs, Radon and Indoor Air Quality, 4052 Bald Cypress Way, Bin #A08, Tallahassee, Florida 32399-1710 (1-800-543-8279).

(c) In addition, the certified radon measurement specialist or business shall notify the client by telephone and certified mail within 2 business days of any measurement with results equal to or greater than 100 pCi/l or 0.5 WL and advise the client to contact the radon information line at 1-800-543-8279 or at other telephone numbers provided by the department. The results of this measurement shall also be provided to the department by phone and certified mail within the same 2 business day period.

Rulemaking Authority 404.051, 404.056, 404.071 FS. Law Implemented 404.051, 404.056, 404.071 FS. History–New 1-3-89, Amended 10-24-90, 1-18-94, 12-9-96, Formerly 10D-91.1309, Amended 3-16-09.

- 64E-5.1207 Certification Requirements for Radon Mitigation Businesses.

The following qualifications are required for certification as a radon mitigation business:

(1) Submission of an application on DH Form 1749 Working Level Month Exposure Form, Jan. 93, and a nonrefundable application and certification fee from the following schedule:

Application received on or before June 30, 2009:						$200

Application received between July 1, 2009 and June 30, 2010:					$300

Application received between July 1, 2010 and June 30, 2011:					$375

Application received between July 1, 2011 and June 30, 2012:					$450

Application received on or after July 1, 2012:						$525

(2) Identification of all mitigation materials and systems offered, diagnostic tests performed, and other related services offered.

(3) Identification of the certified radon mitigation specialists and certified radon mitigation technicians to be used by the business.

(4) If a certified radon mitigation business loses its certified radon mitigation specialist, the certified radon mitigation business shall notify the department in writing within 5 business days. The certified radon mitigation business shall obtain another certified radon mitigation specialist within 30 days of the loss.

(5) Identification of procedures and instrumentation used to perform diagnostic tests.

(6) Disclosure of all reporting forms and system operating instructions provided to clients.

(7) Disclosure of copies of current publications and advertisements of radon-related services made by the applicant.

(8) Development, disclosure and adherence to a health and safety program to limit employees’ exposure to radon progeny during the course of their employment. Such a program shall include measures to keep each employee’s exposure, which shall not exceed 4 working level months per year, as low as reasonably achievable. The certified radon mitigation business shall maintain records showing the radiation exposures of all individuals for whom personnel monitoring is required. Such records shall be kept on DH Form 1834 Working Level Month Exposure Form, Aug. 90, or on clear and legible records containing all the information required by that form. Records shall be maintained in working level months showing the results of exposure during radon mitigation activities and shall be preserved for 20 years. The discontinuance or curtailment of activities does not relieve the certified business of responsibility for retaining all radiation exposure records.

(9) The certified radon mitigation business shall maintain on staff or retain as a consultant a certified radon mitigation specialist. The certified radon mitigation specialist shall monitor the applicant’s mitigation activities, ensure compliance with state and federal regulations and be responsible for evaluating diagnostic tests in a building and designing mitigation systems. The certified radon mitigation specialist shall be present when requested by the department during scheduled visits by the department and shall physically observe and evaluate each certified radon mitigation technician in the performance of his mitigation duties at least once each calendar quarter to insure adequate supervision. A record of this evaluation shall be signed by the specialist and maintained on file by the business. If the mitigation reports submitted to meet the requirements of subsection 64E-5.1207(18), F.A.C., show that no diagnostics or mitigation installations were performed during an entire calendar quarter by any of the certified radon mitigation technicians working for a certified radon mitigation business, the visit, observations and evaluations by the specialist are not required for that quarter. The quarterly visit, observations and evaluations by the specialist must be resumed within the same quarter in which diagnostic or mitigation activities are resumed. The interval between visits, observations and evaluations by the specialist shall not exceed 1 year.

(10) The certified radon mitigation business shall assure that radon mitigation system installations are performed under the supervision of a certified radon mitigation specialist or certified radon mitigation technician. A certified radon mitigation business may contract with a non-certified business to perform mitigation installation provided that work performed by non-certified persons is under the supervision of a certified radon mitigation specialist or certified radon mitigation technician.

(11) The certified mitigation business shall install radon mitigation systems as specified in Chapter 9B-53, F.A.C., Florida Standard for Mitigation of Radon in Existing Buildings, June 1, 1994.

(12) The certified radon mitigation business shall provide all warranty information on the reduction of the radon level, or the proper functioning of mitigation equipment in writing to clients.

(13) The certified radon mitigation business shall ensure that each building is measured for radon levels before and after mitigation work is performed. Such measurements shall be of sufficient type, duration and consistency to allow for comparison of before and after mitigation radon levels, shall be performed with devices that meet proficiency requirements of the USEPA measurement proficiency program, and shall be performed by a certified measurement business. The post-mitigation measurements shall be started no sooner than 24 hours after the mitigation system is fully operational. The mitigation system shall be operated normally and continuously during the entire measurement period. The results of both the pre-mitigation and the post-mitigation measurements shall be sent to the department as specified in subsection 64E-5.1207(18), F.A.C.

(14) The certified radon mitigation business shall maintain at a minimum the following records for 5 years:

(a) Records of all mitigation work performed, including client name, address, initial and follow-up measurement results, diagnostic test results, a description of each mitigation system and materials installed, cost of each system, post-mitigation measurement methods and results and all pertinent dates.

(b) Records of mitigation plans developed and signed by a certified radon mitigation specialist.

(c) Records of all instrument calibrations and warrantied equipment installed.

(d) Copies of the certification for radon mitigation specialists and radon mitigation technicians employed or used as consultants.

(e) Records of specialist site visits and observations and evaluations of technicians.

(15) The certified radon mitigation specialist or certified radon mitigation technician shall perform a visual inspection and diagnostic tests as appropriate before system installation to determine the appropriate mitigation system to be installed. Observations made during visual inspections shall be documented by the business.

(16) The certified radon mitigation business shall provide the client with written instructions on the operation and maintenance of each component of the mitigation system. Information on adverse effects produced by the operation of a mitigation system, including added energy costs, shall be included.

(17) Identification of all branch office locations and the certified specialists and technicians at branch locations who will perform mitigation services for the certified business. Branch offices shall identify the name of the certified business on correspondence and reports to clients and shall accept payments in the name of the certified business. The branch office shall be considered a part of the certified business and make reports and follow procedures approved for the certified business.

(18) A certified radon mitigation business shall provide the following reports; a certified radon mitigation business shall submit mitigation reports to the department by the 15th day of each month. Reports shall include building information on all radon mitigation work completed during the previous month, the name of the certified radon measurement business and individual performing the pre-mitigation and post-mitigation radon and radon progeny measurements, and the measurement results. If no radon mitigations were performed during a month, the report shall indicate none were performed. Reports shall be submitted on DH Form 1753 Monthly Report for Certified Radon Mitigation Businesses, Radon Mitigation Report, Mar 90.

Rulemaking Authority 404.051, 404.056, 404.071 FS. Law Implemented 404.051, 404.056, 404.071 FS. History–New 1-3-89, Amended 10-24-90, 1-18-94, 12-9-96, Formerly 10D-91.1310, Amended 3-16-09.

- 64E-5.1208 Measurement Requirements and Procedures.

(1) The following buildings are required to be measured for radon or radon progeny.

(a) All public and private school buildings or school sites housing students in kindergarten through grade 12.

(b) All state-owned, state-operated, state-regulated or state-licensed 24-hour care facilities. This shall include but not be limited to the following: hospitals, nursing homes, foster homes, assisted living facilities, mental institutions, nursery facilities providing 24-hour care, detention centers, correctional institutions, and pre-release centers.

(c) All state-licensed day care centers for children or minors.

(2) The radon measurement procedures for use in fulfillment of the requirements of Section 404.056, F.S., Mandatory Radon Measurement Protocols (Brochure DH 150-334), may be obtained from the Department of Health and is incorporated herein by reference. This measurement protocol will be used to identify structures in which the potential exists for elevated radon concentrations, to specify longer term measurements to establish occupant exposure levels when elevated concentrations are found, and to periodically remeasure facilities.

(3) Measurement reports for dwelling units which have been measured using section 2 of the USEPA Radon and Radon Decay Product Measurements in Homes, Publication No. EPA 402-R-92-003, shall be reported to the department on DH Form 1778 Residential Radon Measurement Report, Feb. 96, herein incorporated by reference and available from the department. Measurements reports for non-residential buildings which have been measured using sections 3 and 4 of the Mandatory Radon Measurement Protocols (Brochure DH 150-334), shall be reported to the department on DH Form 1777 Non Residential Radon Measurement Report, Feb. 96, herein incorporated by reference and available from the department.

Rulemaking Authority 404.051, 404.056 FS. Law Implemented 404.022, 404.051(1), (4), (7), 404.056(1), (5), (7), (9) FS. History–New 10-8-89, Amended 1-18-94, 12-9-96, Formerly 10D-91.1317.

PART XIII RADIATION SAFETY REQUIREMENTS FOR POSSESSION AND USE OF SEALED OR UNSEALED SOURCES OF RADIOACTIVE MATERIALS

- 64E-5.1301 Sealed or Unsealed Sources of Radioactive Material.

The rules in this part establish radiation safety requirements for licensees possessing or using sealed or unsealed sources of radioactive materials not otherwise specified in a license or addressed in these rules. The requirements of this part are in addition to and not in substitution for other applicable requirements of these rules. Licenses of broad scope are exempt from the requirements of Rule 64E-5.1313, subsections 64E-5.1318(2) and 64E-5.1319(1), (2), (3), (4), F.A.C. Except for Rule 64E-5.1320, F.A.C., the requirements of this part do not apply to persons licensed as specified in Parts IV, VI, and XI. General licensees as specified in subsections 64E-5.206(7) and (8), F.A.C., are exempt from the requirements of this part.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), 404.061(2), (3), 404.081(1) FS. History–New 5-12-93, Amended 5-15-96, Formerly 10D-91.1401, Amended 2-11-10.

- 64E-5.1302 Operating and Emergency Procedures.

The licensee’s operating and emergency procedures shall be posted in accordance with Rule 64E-5.901, F.A.C., and shall accompany portable devices at all times. The procedures shall include instructions in the following as applicable to the type of use:

(1) The uses of sources of radiation so that exposures are maintained as low as reasonably achievable and no individual is likely to be exposed to radiation doses in excess of the standards in Part III;

(2) Methods and occasions for conducting radiation surveys;

(3) Methods and occasions for locking and securing sources of radiation;

(4) Personnel monitoring and the use of personnel monitoring equipment;

(5) Minimizing exposure of individuals in the event of an accident;

(6) Notifying proper personnel in the event of damage, loss, theft, or accident involving sources of radiation;

(7) General guidelines for the safe handling and use of unsealed sources of radioactive materials;

(8) Maintenance of records;

(9) Procedures for picking up, receiving and opening packages containing radioactive materials; and

(10) The transportation of radioactive sources to temporary job sites, including the packaging, marking, labeling and placing of such sources in vehicles, placarding of vehicles, securing the sources during transportation and possessing proper shipping papers and emergency response information.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Amended 1-1-94, 5-15-96, Formerly 10D-91.1403.

- 64E-5.1303 Leak Test Requirements for Possession of Sealed Sources.

(1) A licensee in possession of a sealed source shall assure that:

(a) The sealed source is tested for leakage before its first use unless the licensee has a certificate from the supplier indicating that the sealed source was tested semiannually before transfer to the licensee;

(b) The sealed source is tested for leakage at least semiannually or at intervals approved by the department;

(c) Leak tests are capable of detecting 0.005 microcurie (185 Bg) of radioactive material on the test sample, or, in the case of radium, the escape of radon of 0.001 microcurie (37 Bg) each 24 hours;

(d) Test samples are taken from the sealed source or from the surfaces of the device in which the sealed source is mounted or stored on which radioactive contamination might be expected to accumulate; and

(e) Device test samples are taken when the sealed source is in the off or shielded position.

(f) Leak tests are analyzed by individuals who are licensed by the department, U.S. Nuclear Regulatory Commission, an agreement state or a licensing state to perform leak test services.

(2) A licensee shall retain leak test records for 3 years. The records shall contain the manufacturer’s name, the model and serial numbers of each sealed source tested, the identity of each sealed source radionuclide and its estimated activity, the measured activity of each test sample expressed in microcuries (becquerels), the date of the test, and the signature of the radiation safety officer or designee.

(3) If the leak test reveals the presence of 0.005 microcurie (185 Bq) or more of removable contamination, the licensee shall:

(a) Immediately withdraw the sealed source from use and cause it to be decontaminated and repaired or to be disposed of in accordance with these regulations; and

(b) File a report with the department within 5 days of receiving the leak test results describing the equipment involved, the test results, and the action taken.

(4) A leak test is not required on the following sealed sources:

(a) Sealed sources containing only radioactive material with a half-life of less than 30 days;

(b) Sealed sources containing only radioactive material as a gas;

(c) Sealed sources containing 100 microcuries (3.7 MBq) or less of beta or photo-emitting material or 10 microcuries (370 kBq) or less of alpha-emitting material; and

(d) Sealed sources that are listed on a department license for storage only. The licensee shall test each such sealed source for leakage before any use or transfer unless it has been tested for leakage within 6 months before the date of use or transfer.

(5) The department is authorized to approve leak test frequencies other than semiannually. Criteria used to determine these frequencies include:

(a) The isotope and activity;

(b) The requested usage;

(c) The environmental hazards to which the sealed sources may potentially be exposed;

(d) The manufacturer’s recommended frequency; and

(e) Specific information on the sealed source or device provided by the U.S. Nuclear Regulatory Commission, an agreement state or a licensing state, including:

1. Sealed Source and Device Registry sheets;

2. Naturally Occurring or Accelerator Produced Radioactive Materials Sealed Source and Device Registry sheets.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1404.

- 64E-5.1304 Inventory Requirements.

A licensee who possesses sealed sources shall conduct a physical inventory of all such sealed sources semiannually unless another interval is specified in the license. Inventory records shall be retained for 3 years. Inventory records shall contain the following:

(1) The model and serial number of each sealed source;

(2) The identity of each sealed source radionuclide and its estimated activity;

(3) The location of each sealed source;

(4) The date of the inventory; and

(5) The signature of the radiation safety officer or designee.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1405.

- 64E-5.1305 Training Requirements, Authority, Duties and Responsibilities of the Radiation Safety Officer.

(1) The licensee shall appoint a radiation safety officer with the authority to fulfill the duties and responsibilities listed in this part.

(2) The radiation safety officer shall have sufficient training and experience with radioactive materials to be a user of the requested licensed materials, unless otherwise specified in the license. This training shall include practical experience in the safe use of radioactive materials and knowledge of procedures, facilities and equipment.

(3) The duties and responsibilities of the radiation safety officer shall include the following:

(a) Ensure that all terms and conditions of the license and these regulations are complied with;

(b) Ensure that the sealed sources are leak tested timely and as prescribed by the manufacturer or by the license;

(c) Ensure that radioactive materials are used only by individuals who are authorized by the license and that all individuals wear required personnel monitoring equipment;

(d) Maintain all records required by the license and these regulations. These records shall include personnel monitoring records, leak test records, inventory records, training records for users and receipt, transfer and disposal records;

(e) Ensure that radioactive materials are properly secured against unauthorized access or removal;

(f) Serve as a contact with the department for events such as the loss, theft or damage of radioactive material; and

(g) Ensure that all users read and understand the licensee’s emergency, operating and radiation safety procedures.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1406.

- 64E-5.1306 Opening Sealed Sources.

Unless otherwise specifically licensed by the department, the U.S. Nuclear Regulatory Commission, an agreement state or a licensing state to perform such services, the licensee shall not open sealed sources.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1407.

- 64E-5.1307 Training Requirements for Authorized Users.

(1) Radioactive materials shall be used by individuals who are qualified by training and experience to protect public health, safety and the environment. A description of this training must be submitted and approved by the department and include the following:

(a) Principles and fundamentals of radiation protection and safety practices related to the use of radioactive materials, including ALARA principles;

(b) Radioactivity measurements;

(c) Use of radiation detection instruments and monitoring techniques;

(d) Biological effects of radiation;

(e) Transportation of radioactive materials;

(f) Practical experience with the use of radioactive materials; and

(g) Licensee’s operating and emergency procedures.

(2) For licensees who propose to train their own personnel to be authorized users, the following must be provided to and approved by the department:

(a) Instructor qualifications, including training and experience with radioactive materials specifically relating to the topics of instruction;

(b) A detailed training program, including duration of training for each of the topics listed in (1) of this section;

(c) The method of testing the knowledge of students, such as a written and practical examination, and whether the examination is open or closed book; and

(d) If an examination is used, the passing score, method of retesting students who do not pass and an example of the examination with the correct answers indicated.

(3) Records of training shall be maintained during the employment of the individual or 5 years, whichever is greater.

(4) Unless otherwise specified in the license, a licensee’s authorized user training program is not transferable to another licensee.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1409.

- 64E-5.1308 Additional Requirements for General Licenses.

Specific licensees authorized for a general license described in subsection 64E-5.205(4), 64E-5.206(1), (4) or (6), F.A.C., shall comply with the regulations that are applicable to that general license and Rules 64E-5.1304 and 64E-5.1305, F.A.C. Specific licensees authorized for the general license in subsection 64E-5.205(4), F.A.C., or possess generally licensed devices described in subsection 64E-5.206(1) or (4), F.A.C., are not required to remit the annual fees specified in subparagraph 64E-5.204(1)(c)1., 2., or 5., F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1410.

- 64E-5.1309 Training for Current Authorized Users.

Individuals who are authorized users on a department, U.S. Nuclear Regulatory Commission, agreement state or licensing state license on 5-12-93 who perform only those procedures for which they are authorized on that date need not comply with the training requirements in Rules 64E-5.1307, 64E-5.1312 and 64E-5.1313, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1411.

- 64E-5.1310 Personnel Monitoring.

(1) Unless otherwise specified in the license, no licensee shall permit any individual to use or to assist in the use of sealed sources of radiation in portable devices unless such individual wears a film badge, OSLD, or a TLD.

(2) Unless otherwise specified in the license, no license shall permit any individual to perform installations, maintenance or service, initial radiation surveys, relocations or removal from service of sealed sources in fixed devices unless such individual wears a film badge, OSLD, or a TLD.

(3) Licensees who use iodine 125, iodine 131, hydrogen 3, uranium 234, uranium 235 or uranium 238 and are required to have a bioassay program must submit a description of their bioassay program for approval by the department.

(4) A whole body film badge, OSLD, or TLD is required to be worn by any individual using or assisting in the use of unsealed sources of radioactive materials of any gamma-emitting isotope with a gamma ray energy greater than 50 kiloelectron volts or the use of any beta-emitting isotope with a maximum beta energy of 300 kiloelectron volts or more.

(5) An extremity film badge, OSLD, or TLD is required to be worn by any individual using or assisting in the use of unsealed sources of radioactive materials of 1,000 microcuries (37 MBq) or more of beta-emitting isotopes with a maximum beta energy of 1,000 kiloelectron volts or more in any month or by any individual who receives a dose of 40 millirem (400 µSv) or more on a whole body film badge, OSLD, or TLD for 2 consecutive months.

(6) Each film, OSLD, and TLD badge shall be assigned to and worn by only one individual. Film badges and extremity OSLDs and TLDs must be replaced monthly. Whole body OSLDs, and TLDs must be replaced quarterly. After replacement, each film badge, OSLD, and TLD must be promptly processed.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), 404.061(2), 404.081(1), (2) FS. History–New 5-15-96, Formerly 10D-91.14111, Amended 10-8-00.

- 64E-5.1311 Storage, Security and Transportation Precautions.

(1) Each sealed source of radioactive material shall be provided with a storage or transport container. The container shall be equipped with a lock or tamper seal to prevent unauthorized removal of or exposure to the source of radiation.

(2) All portable gauge licensees must comply with either paragraph (2)(a), or (2)(b) below. Effective January 1, 2007, portable gauge licensees must comply only with paragraph (2)(b).

(a) Sealed sources must have a minimum of two locks between the device and the public when being transported or stored.

(b) Each portable gauge licensee shall use a minimum of two independent physical controls that form tangible barriers to secure portable gauges from unauthorized removal, whenever portable gauges are not under the control and constant surveillance of the licensee.

(3) Transport containers shall be physically secured in the transporting vehicle to prevent accidental loss, tampering, or unauthorized removal. The sealed source shall be transported as far away from occupied areas of the vehicle as possible.

(4) Sealed sources not in storage or being transported must be under the constant surveillance and immediate control of the licensee.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1412, Amended 9-28-06.

- SUBPART B REQUIREMENTS FOR THE POSSESSION AND USE OF SEALED SOURCES IN PORTABLE DEVICES

- 64E-5.1312 Training and User Requirements for the Possession and Use of Sealed Sources in Portable Devices.

(1) Users of sealed sources in portable devices must have completed a minimum of 8 hours of training from individuals approved by the department. This training must include the areas described in Rule 64E-5.1307, F.A.C.

(2) Documentation of training for each user must be maintained for the duration of employment or 5 years, whichever is greater.

(3) Sealed sources in portable devices may be used by individuals who are under the direct supervision and in the physical presence of an authorized user.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1415.

- SUBPART C REQUIREMENTS FOR THE POSSESSION AND USE OF SEALED SOURCES IN FIXED DEVICES

- 64E-5.1313 Training and User Requirements for the Possession and Use of Sealed Sources in Fixed Devices.

Unless otherwise specifically licensed by the department, the U.S. Nuclear Regulatory Commission, an agreement state or a licensing state to perform such services, the licensee shall not remove sealed sources from source holders; remove source holders containing sealed sources from devices; perform maintenance or repair of devices or sources holders containing sealed sources, including repair or maintenance of the shutter; perform installations, replacement, removal from service, relocations, or disposal of sealed sources, source holders or devices containing sealed sources; or perform initial radiation surveys of devices or source holders.

(1) Users of sealed sources in fixed devices must have completed a minimum of 8 hours of training from individuals approved by the department. This training must include the areas described in Rule 64E-5.1307, F.A.C.

(2) Individuals who perform installations, maintenance or service, initial radiation surveys, relocations, or removal from service must have completed a minimum of 40 hours of training from individuals approved by the department. This training must include the following:

(a) The principles and fundamentals of radiation protection and safety practices related to the use of radioactive material;

(b) Radiation measurements, use of radiation detection instruments and monitoring techniques;

(c) Biological effects of radiation;

(d) Procedures for performing services; and

(e) Actual practice in performing the services.

(3) Documentation of training for each user must be maintained for the duration of employment or 5 years, whichever is greater.

(4) Sealed sources in fixed devices may be used by individuals under the supervision of an authorized user. An authorized user must be available at all times when sealed sources in fixed devices are being used.

(5) Installations, maintenance or service, initial radiation surveys, relocations or removal from service may be performed by individuals who are under the direct supervision and in the physical presence of an individual who is an authorized user for these operations.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1418.

- 64E-5.1314 Possession of Survey Instruments.

A licensee authorized to perform installations, maintenance or service, initial radiation surveys, relocations or removal from service of sealed sources in fixed devices shall possess portable radiation survey instruments with a range from 1 millirem (10 µSv) per hour to 200 millirem (2 µSv) per hour. The instruments shall be operable and calibrated as provided in Rule 64E-5.314, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Amended 1-1-94, Formerly 10D-91.1419.

- 64E-5.1315 Additional Requirements.

A licensee must post and provide to personnel lock-out procedures that prevent employees from entering the radiation beam during maintenance, repairs, or other work in, on, or around a bin, tank, hopper or pipe on which a device is mounted. The department is authorized to require a physical barrier around certain types of devices where the possibility exists that an individual could be exposed to the beam of radiation.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1420.

- 64E-5.1316 General Rules for the Safe Use of Unsealed Sources of Radioactive Materials.

The licensee shall assure that all individuals who handle unsealed sources of radioactive materials comply with the following, unless otherwise specified in the license:

(1) Laboratory coats or other protective clothing are worn at all times in areas where radioactive materials are used;

(2) Disposable gloves are worn at all times while handling radioactive materials;

(3) Eating, drinking, smoking, or applying cosmetics in any area where radioactive material is stored or used is prohibited;

(4) Storing food, drinks, or personal effects in areas where radioactive material is stored or used is prohibited;

(5) If applicable, personnel monitoring devices are worn at all times while in areas where radioactive materials are used or stored;

(6) Radioactive waste is disposed of only in designated, labeled, and properly shielded receptacles; and

(7) Radioactive materials are confined in clearly labeled appropriate containers.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1422.

- 64E-5.1317 Storage and Control of Volatiles and Gases.

(1) A licensee shall store volatile radioactive materials and radioactive gases in the shippers’ radiation shield and container or an equivalent shield and container.

(2) Unless otherwise specified in the license, a licensee shall store and use radioactive volatiles and gases in a properly functioning glove box or fume hood that will maintain airborne concentrations within the limits prescribed by State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table I, Column 2, and Table II, Column 1.

(3) Unless otherwise specified in the license, the glove box or fume hood shall either be directly vented to the atmosphere through an air exhaust or provide for collection and decay or disposal of the volatile or gas.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Amended 1-1-94, Formerly 10D-91.1423, Amended 12-26-13.

- 64E-5.1318 Instrumentation.

(1) The licensee shall have instruments available to detect radioactive materials listed on the license, unless otherwise authorized by the department. Instrumentation shall be sensitive enough to detect activities required for adequate contamination control described in this part.

(2) The licensee must submit a description of the equipment and procedures to be followed in measuring contamination for departmental approval. These procedures shall include the following:

(a) Type of instrument detection system used, such as a Geiger-Mueller or scintillation detector with a scaler, single or multichannel analyzer, and type of radiation detected;

(b) Background counting times and average background counts;

(c) Sample counting times;

(d) Instrument efficiency and calculation of efficiency determination;

(e) Instrument’s lower limit of detection using the sample counting time and the isotope and activity of calibration standards;

(f) Sample calculation converting counting results into activity or activity per unit time, if applicable; and

(g) Frequency of instrument calibration.

(3) If portable radiation survey instruments are used to determine compliance with the contamination control action levels described in this part, the instruments shall be operable and calibrated as provided in Rule 64E-5.314, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Amended 1-1-94, Formerly 10D-91.1425.

- 64E-5.1319 Contamination Control Program.

The licensee shall establish and submit for department approval a contamination control program to limit the spread of unsealed sources of radioactive materials. This program shall include contamination action levels, corresponding actions taken if these levels are exceeded, frequency of measurement and maintenance of records.

(1) Unless otherwise specified in the license, the licensee shall perform the following:

(a) A survey with a radiation survey instrument shall be completed at the end of each day of use or receipt of all areas where radioactive materials are used or received.

(b) A weekly survey with a radiation survey instrument shall be completed of all areas where radioactive materials or radioactive waste are stored.

(c) A wipe survey shall be completed for removable contamination weekly during weeks of use of all areas where radioactive materials or waste are routinely used or stored.

(d) Surveys required by paragraphs (1)(a), (b) and (f) of this section shall be completed with an instrument capable of measuring dose rates as low as 0.1 millirem (1 µSv) per hour, if applicable.

(e) A wipe survey shall be completed for removable contamination at the end of each day of use of all areas where radioactive materials are routinely used if the radioactive materials authorized by the license are not detectable with instruments described in (1)(d) above.

(f) Monitoring shall be performed of individuals for contamination prior to leaving the restricted area.

(g) The radiation safety officer shall be notified if the contamination detected during the surveys required by this section exceeds the action levels specified in subsection (2), below.

(2) Unless otherwise specified in the license, the licensee shall take the following actions if the corresponding action levels are exceeded:

(a) Follow decontamination procedures and resurvey to determine effectiveness of decontamination efforts if the action levels are exceeded when conducting surveys described in paragraphs (1)(a), (b) or (f) above;

(b) Follow emergency procedures if the radiation detected during the survey required in paragraph (1)(f) of this section is above action levels; and

(c) Follow decontamination procedures and resurvey to determine effectiveness of decontamination efforts if the action levels listed below are exceeded for the survey required in paragraph (1)(c) or (e) above;

1. 100 dpm per 100 square centimeters of any alpha-emitting radioactive materials not listed in this section;

2. 1,000 dpm per 100 square centimeters of any beta- or gamma-emitting radioactive materials not listed in this section;

3. 50 dpm per 100 square centimeters of any transuranic;

4. 2,000 dpm per 100 square centimeters of any uranium; or

5. 2,000 dpm per 100 square centimeters of any radioactive material with a half-life of less than 80 hours.

(3) The licensee shall retain a record of each survey required by this section for 3 years. These records shall include:

(a) The date of the survey;

(b) An annotated diagram of each area surveyed;

(c) Background levels;

(d) Measured dose rates, keyed to the diagram, expressed in millirem (microsieverts) per hour or the removable contamination, keyed to the diagram, expressed in dpm per 100 square centimeters, or counts per minute if performed with a radiation survey instrument, with the action level described in paragraph (2)(a) of this section;

(e) The serial number and model number of the instrument used to make the survey or analyze the samples; and

(f) The initials of the person who performed the survey.

(4) Records of surveys described in paragraph (1)(f) need not be recorded unless the action level described in paragraph (2)(a) of this section is exceeded.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.1426.

SUBPART D REQUIREMENTS FOR POSSESSION AND USE OF UNSEALED

SOURCES OF RADIOACTIVE MATERIALS

- 64E-5.1320 Bioassay Program.

The licensee shall establish and submit for department approval a bioassay program used to evaluate internal doses. At a minimum an acceptable program shall include the following action levels for organ uptakes, corresponding actions taken if these levels are exceeded, frequency of measurement and maintenance of records.

(1) Routine bioassay is required when an individual handles in open form unsealed quantities of radioactive iodine that exceed those shown in the table 1 below. The quantities shown apply to both the quantity handled at any one time or integrated as the total amount of activity introduced into a process by an employee over any 3-month period.

| TABLE 1 |  |  |
| --- | --- | --- |
| I-125 or I-131 Activity Handled in Unsealed Form Requiring Bioassay |  |  |
| Type of Operation | Volatile or Dispersible | Bound to Nonvolatile Agent |
| Processes in open room or bench, with possible escape of iodine from process vessels | 1.0 mCi (37 MBq) | 1.0 mCi (37 MBq) |
| Processes with possible escape of iodine carried out within a fume hood of adequate design, face velocity, and performance reliability | 1.0 mCi (37 MBq) | 10.0 mCi (370 MBq) |
| Processed carried out within glove boxes, ordinarily closed, but with possible release of iodine from process and occasional exposure to contaminated box and box leakage | 10.0 mCi (370 MBq) | 100.0 mCi (3700 MBq) |

(a) A bioassay shall be taken within 72 hours of initial use of radioiodine and every 2 weeks thereafter. When radioiodine use is on an infrequent basis (less than every 2 weeks), a bioassay shall be taken within 10 days of the last day of use.

(b) If the thyroid burden at the time of measurement exceeds 0.12 microcurie (4.44 KBq) of iodine 125 or 0.04 microcurie (1.48 KBq) of iodine 131, the following actions shall be taken:

1. An investigation of the operations involved, including air and other facility surveys, shall be carried out to determine the cause(s);

2. Corrective actions that will eliminate or lower the potential for further exposures shall be implemented;

3. A repeat bioassay shall be taken within 2 weeks of the previous measurement and shall be evaluated within 24 hours after the measurement in order to confirm the presence of internal radioiodines; and

4. Notification reports must be provided as required by Rules 64E-5.345 and 64E-5.347, F.A.C., or as required by conditions of the license; and

(c) A record of each bioassay shall be maintained for inspection by the department in an auditable form for 3 years and shall include the date of the bioassay, the name of the individual, and the thyroid burden at the time of the measurement.

(2) Routine bioassay is required when an individual handles in open form unsealed quantities of tritium that exceed those shown in table 2 below. The quantities shown apply to both the quantity handled at any one time or integrated as the total amount of activity introduced into a process by an employee over a 1-month period.

| TABLE 2 |  |  |
| --- | --- | --- |
| Tritium Activity Handled in Levels or Concentrations Requiring Bioassay |  |  |
| Type of Operation | HTO and Other Tritiated Compounds (Including Nucleotide Precursors) | Tritium (HT or T ) Gas in Sealed Process Vessels |
| Processes in open room or bench with possible escape of tritium from process vessels | 0.1 Ci ( 3.70 GBq) | 100 Ci (3.7 TBq) |
| Processes with possible escape of tritium carried out within a fume hood of adequate design, face velocity, and performance reliability | 1 Ci (37 GBq) | 1,000 Ci (37 TBq) |
| Processes carried out within glove boxes that are ordinarily closed but with possible release of tritium from process vessels and occasional exposure to contaminated box and leakage | 10 Ci (370 GBq) | 10,000 Ci (370 TBq) |

(a) A bioassay shall be taken within 72 hours of initial use of tritium and every 2 weeks thereafter. When work with tritium is on an infrequent basis (less frequent than every 2 weeks), a bioassay shall be taken within 10 days of the last day of use.

(b) If the urinary tritium concentration exceed 5 microcuries (185 KBq) per liter at the time of the measurement the following actions shall be taken:

1. An investigation of the operations involved, including air and other facility surveys, shall be carried out to determine the cause(s);

2. Corrective actions that will eliminate or lower the potential for further exposures shall be implemented;

3. A repeat bioassay shall be taken within 1 week of the previous measurement and shall be evaluated within 1 week after the measurement. Internal dose commitments shall be estimated using at least two bioassays and other survey data, including the probable times of intake of tritium; and

4. Notification reports must be provided as required by Rules 64E-5.345 and 64E-5.347, F.A.C., or as required by conditions of the license; and

(c) A record of each bioassay shall be maintained for inspection by the department in an auditable form for 3 years and shall include the date of the bioassay, the name of the patient, and the urinary tritium concentration at the time of the measurement.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 2-11-10.

- 64E-5.1401 Irradiators.

(1) This part contains requirements for licenses authorizing the use of sealed sources containing radioactive materials in irradiators to irradiate objects or material. This part also contains radiation safety requirements for operating irradiators.

(2) The rules in this part apply to panoramic irradiators which have either dry or wet storage of radioactive sealed sources and to underwater irradiators in which both the source and the product being irradiated are under water. Irradiators covered by the rules in this part are those whose radiation dose rates exceed 500 rads (5 grays) per hour at 1 meter from the radioactive sealed sources in air or in water.

(3) The rules in this part do not apply to self-contained dry-source-storage irradiators in which both the source and the area subject to irradiation are contained within a device and are not accessible by personnel. This part also does not apply to radioactive material used for medical radiology, teletherapy, industrial radiography, gauging, calibration of radiation instruments, or open-field agricultural irradiations.

(4) The requirements of this part are in addition to other applicable requirements of these rules.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1501.

- 64E-5.1402 Definitions.

(1) “Doubly encapsulated sealed source” means a sealed source in which the radioactive material is sealed within a capsule which is sealed within another capsule.

(2) “Irradiator” means a facility which uses radioactive sealed sources to irradiate objects or materials and in which radiation dose rates exceeding 500 rads (5 grays) per hour exist at 1 meter from the sealed radioactive sources in air or water, as applicable for the irradiator type, but does not include irradiators in which both the sealed source and the area subject to irradiation are contained within a device and are not accessible to personnel.

(3) “Irradiator operator” means an individual who successfully completed the training and testing specified in Rule 64E-5.1416, F.A.C., and who is authorized by the licensee and approved by the department to operate the irradiator without a supervisor present.

(4) “Panoramic dry-source-storage irradiator” means an irradiator in which the irradiation occurs in air in areas potentially accessible to personnel and in which the sources are stored in shields made of solid materials. The term includes beam-type dry-source-storage irradiators in which only a narrow beam of radiation is produced to perform irradiations.

(5) “Panoramic irradiator” means an irradiator in which the irradiation occurs in air in areas potentially accessible to personnel. The term includes beam-type irradiators.

(6) “Panoramic wet-source-storage irradiator” means an irradiator in which the irradiation occurs in air in areas potentially accessible to personnel and in which the sources are stored under water in a storage pool.

(7) “Pool irradiator” means an irradiator in which the sources are stored or used in a pool of water, including panoramic wet-source-storage irradiators and underwater irradiators.

(8) “Product conveyor system” means a system to move the product to be irradiated to, from, and within the area where irradiation occurs.

(9) “Radiation room” means a shielded room in which irradiation occurs. Underwater irradiators do not have radiation rooms.

(10) “Seismic area” means any area where the probability of a horizontal acceleration in rock of more than 0.3 times the acceleration of gravity in 250 years is greater than 10 percent as designated by the U.S. Geological Survey.

(11) “Underwater irradiator” means an irradiator in which the sources always remain shielded under water and personnel cannot access the sealed source or the space subject to irradiation without entering the pool.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1502.

- 64E-5.1403 Specific License for Large Irradiators.

An application for a specific license to use sealed sources in a large irradiator shall be made as specified in Rule 64E-5.207, F.A.C. A separate license is required for each large irradiator, radiation room or underwater irradiator. The department will approve an application for a specific license for the use of licensed material in an irradiator if the applicant meets the requirements contained in this section.

(1) The applicant must satisfy the general requirements specified in Rule 64E-5.208, F.A.C., and the requirements contained in this part.

(2) The applicant must describe training for irradiator operators which includes the following:

(a) At least 40 hours of classroom training;

(b) At least 160 hours of on-the-job or simulator training;

(c) Safety reviews;

(d) The means the applicant will use to evaluate the operator’s knowledge and understanding of and ability to comply with the department’s rules and licensing requirements and the applicant’s operating and emergency procedures; and

(e) The minimum qualifications of personnel who provide training.

(3) The applicant shall submit an outline or summary of the written operating and emergency procedures specified in Rule 64E-5.1417, F.A.C. The outline or summary must include important radiation safety aspects of the procedures.

(4) The applicant shall describe the radiation safety responsibilities and authority of the radiation safety officer and other management personnel and specify who has the authority to stop unsafe operations. The applicant also shall describe the qualifications required of the radiation safety officer.

(5) The applicant shall submit a description of the access control systems required by Rule 64E-5.1406, F.A.C., the radiation monitors required by Rule 64E-5.1409, F.A.C., the method to detect leaking sources required by Rule 64E-5.1420, F.A.C., including the sensitivity of the method, and a diagram of the facility which shows the position of all required interlocks and radiation monitors.

(6) The applicant shall assure that any radioactive source not used in the irradiation process shall be removed from the irradiator pool and disposed of or returned to the manufacturer unless otherwise approved by the department.

(7) If the applicant intends to perform leak testing of dry-source-storage sealed sources, the applicant shall establish procedures for leak testing and submit a description of these procedures to the department for approval. The procedures must include the following:

(a) Instruments to be used;

(b) Methods of performing the analysis; and

(c) Pertinent experience of individuals who analyze the samples.

(8) If licensee personnel load or unload sources, the applicant shall describe the qualifications of the personnel and the procedures to be used. If the applicant contracts for source loading or unloading, the loading or unloading must be done by an organization licensed by the U.S. Nuclear Regulatory Commission, an agreement state or a licensing state to load or unload irradiator sources.

(9) The applicant shall perform operational tests on the following to ensure proper functioning of all equipment and safety devices before the irradiator is loaded with sources:

(a) Interlock and radiation safety systems;

(b) Pool integrity and plumbing;

(c) Source rack mechanical positioning system;

(d) Source rack movement and position sensing systems;

(e) Source rack electrical control system;

(f) Uninterruptable electrical power supply for radiation monitoring warning systems;

(g) Fire protection system;

(h) Emergency systems for returning a stuck source rack into the pool;

(i) Systems used for transferring sources to and from transport vehicles; and

(j) Product conveyor system.

(10) The applicant shall describe the operational inspection and maintenance program, including the frequency of operational checks required by Rule 64E-5.1421, F.A.C.

(11) The roof plug opening or removable shielding providing access for the loading and removal of sources shall be large enough to accommodate the largest applicable transportation cask.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1503.

- 64E-5.1404 Start of Construction.

(1) The applicant shall not begin construction of a new irradiator before sending a license application to the department for the irradiator. As used in this paragraph, the term “construction” includes the construction of any portion of the permanent facility on the site but does not include:

(a) Engineering and design work;

(b) Purchase of the site;

(c) Site surveys or soil testing;

(d) Site preparation;

(e) Site excavation;

(f) Construction of warehouse structures; and

(g) Other similar tasks.

(2) Site requirements for an irradiator include geological, radiological and hydrological testing to ensure a stable environment before construction begins. Records of the results of the siting requirements shall be maintained for the life of the facility.

(3) Any activities undertaken before the issuance of a license will be entirely at the risk of the applicant and will have no bearing on the issuance of a license.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1504.

- 64E-5.1405 Performance Criteria for Sealed Sources.

(1) The licensee shall assure that sealed sources installed after August 14, 1996 meet the requirements of this section. A prototype of the sealed source must be leak tested and found leak-free after each of the following tests:

(a) Temperature. The test source must be held at minus 40 degrees Celsius for 20 minutes, 600 degrees Celsius for 1 hour, and then be subjected to a thermal shock test with a temperature drop from 600 degrees Celsius to 20 degrees Celsius within 15 seconds.

(b) Pressure. The test source must be subjected twice to an external pressure of 290 pounds per square inch absolute for at least 5 minutes.

(c) Impact. A 2 kilogram steel weight, 2.5 centimeters in diameter, must be dropped from a height of 1 meter onto the test source.

(d) Vibration. The test source must be subjected three times for 10 minutes each time to a range of vibration from 25 Hertz to 500 Hertz at 5 times the acceleration of gravity for 30 minutes. Each test source must be vibrated for 30 minutes at each resonant frequency found.

(e) Puncture. A 50 gram weight and pin, 0.3 centimeter pin diameter, must be dropped from a height of 1 meter onto the test source.

(f) Bend. If the length of the source is more than 15 times larger than the minimum cross-sectional dimension, the test source must be subjected to a force of 2,000 newtons at its center equidistant from two support cylinders. The distance between the support cylinders is 10 times the minimum cross-sectional dimension of the source.

(2) Sealed sources installed after 8-14-96 must be doubly encapsulated, the isotope or isotope and matrix contained within the sealed sources must be as nondispersible and insoluble as practical or rendered insoluble in water if the source is used in a wet-source-storage or wet-source-change irradiator and must satisfy the requirements specified in subsection 64E-5.210(14), F.A.C.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1505.

- 64E-5.1406 Access Control.

(1) Panoramic irradiators shall not be operated unless the following are met:

(a) Each entrance to a radiation room must have a door or other physical barrier to prevent inadvertent entry of personnel while the sources are exposed. Product conveyor systems can serve as barriers as long as they reliably and consistently function as a barrier. It must not be possible to move the sources out of their shielded position if any door or barrier to the radiation room is open. Opening the door or barrier while the sources are exposed must cause the sources to return promptly to their shielded position. The primary entry door must have a lock which is operated by the same key used to control source movement. The doors and barriers must not prevent any individual in the radiation room from leaving.

(b) Each entrance to a radiation room must have an independent backup access control to detect personnel entry while the sources are exposed if the primary access control fails. Entry while the sources are exposed must cause the sources to return to their fully shielded position and also must activate a visible and audible alarm to make the individual entering the room aware of the hazard. The alarm also must alert at least one other individual of the entry who is on site and who is trained to render or summon assistance promptly.

(c) A radiation monitor must be provided to detect the presence of high radiation levels in the radiation room before personnel entry. The monitor must be integrated with personnel access door locks to prevent room access when the monitor detects high radiation levels. The monitor must generate audible and visible alarms if high radiation levels are detected when personnel entry is attempted. The monitor can be located in the entrance or maze but not in the direct radiation beam.

(d) Before sources move from their shielded position, the source control automatically must activate conspicuous visible and audible alarms to alert people in the radiation room that the sources will be moved from their shielded position. The alarms must give individuals enough time to leave the room before the sources leave the shielded position.

(e) Each radiation room must have a clearly visible and readily accessible control which will allow an individual in the room to return the sources to their fully shielded position.

(f) Each radiation room must contain a control which allows the sources to move from the shielded position only if the control has been activated and the door or barrier to the radiation room subsequently has been closed within a preset time.

(g) Each entrance to the radiation room and each entrance to the area within the personnel access barrier of an underwater irradiator must be posted as required by Rule 64E-5.323, F.A.C. Panoramic irradiators also must be posted as required by Rule 64E-5.323, F.A.C. The sign can be removed, covered, or otherwise made inoperative when the sources are shielded fully.

(h) If the radiation room has roof plugs or other movable shielding, it must not be possible to operate the irradiator unless the shielding is in its proper location. This requirement can be met by interlocks which prevent operation if shielding is not placed properly or by an operating procedure requiring inspection of shielding before operating.

(2) Underwater irradiators must have a personnel access barrier around the pool which must be locked to prevent access when the irradiator is not attended. Only operators and facility management shall have access to keys to the personnel access barrier. There must be an intrusion alarm to detect unauthorized entry when the personnel access barrier is locked. Activation of the intrusion alarm must alert an individual, not necessarily on site, who is prepared to respond or summon assistance.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1506, Amended 10-8-00.

- 64E-5.1407 Shielding.

(1) The radiation dose rate in areas which normally are occupied during operation of a panoramic irradiator must not exceed 2 millirem (0.02 millisievert) per hour at 30 centimeters or more from the wall of the room when the sources are exposed. The dose rate must be averaged over an area not to exceed 100 square centimeters having no linear dimension greater than 20 centimeters. Areas where the radiation dose rates exceed 2 millirem (0.02 millisievert) per hour must be locked, roped off, or posted to prevent access and not entered without written approval or in the physical presence of the radiation safety officer or his designee.

(2) The radiation dose at 30 centimeters over the pool of a pool irradiator when the source is in the fully shielded position must not exceed 2 millirem (0.02 millisievert) per hour.

(3) The radiation dose rate at 1 meter from the shield of a dry-source-storage panoramic irradiator must not exceed 2 millirem (0.02 millisievert) per hour and at 5 centimeters from the shield must not exceed 20 millirem (0.02 millisievert) per hour.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1507.

- 64E-5.1408 Fire Protection.

(1) The radiation room at a panoramic irradiator must have heat and smoke detectors. The detectors must activate an audible alarm. The alarm must be capable of alerting a person who is prepared to summon assistance promptly. The sources must become fully shielded automatically and the air handling systems within the radiation room must be disabled automatically if a fire is detected.

(2) The radiation room at a panoramic irradiator must be equipped with a fire suppression or extinguishing system capable of extinguishing a fire without the entry of personnel into the room. The system for the radiation room must have a shut-off valve to control flooding into unrestricted areas.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1508.

- 64E-5.1409 Radiation Monitors.

(1) Irradiators with automatic product conveyor systems must have a radiation monitor with an audible alarm located to detect loose radiation sources which are carried toward the product exit. If the monitor detects a source, an alarm must sound and product conveyors must stop automatically. The alarm must be capable of alerting an individual in the facility who is prepared to summon assistance. Underwater irradiators in which the product moves within an enclosed stationary tube are exempt from the requirements of this paragraph.

(2) For pool irradiators, the licensee shall provide a means to detect radioactive contamination in pool water each day the irradiator operates. The means can be either an on-line radiation monitor on the pool water purification system or an analysis of pool water. If the licensee uses an on-line radiation monitor, the detection of above normal background radiation levels must activate the alarm. The alarm set-point must be set as low as practical but high enough to avoid false alarms. If a false alarm due to background radiation occurs, the alarm set-point must be increased. Activation of the alarm must cause the water purification system to shut off automatically. However, the licensee can reset the alarm set-point to a higher level if necessary to operate the pool purification system to clean up contamination in the pool as provided specifically in written emergency procedures.

(3) Underwater irradiators which are not in a shielded radiation room must have a radiation monitor over the pool to detect abnormal radiation levels. The monitor must have an audible alarm and a visible indicator at entrances to the personnel access barrier around the pool. The audible alarm can have a manual shut-off. The alarm must be capable of alerting an individual who is prepared to respond promptly.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1509.

- 64E-5.1410 Control of Source Movement.

(1) The mechanism which moves the sources of a panoramic irradiator must require a key to operate. Operation of the mechanism must cause an audible signal to indicate that the sources are leaving the shielded position. Only one key shall be available for use at any time and only operators or facility management can possess it. The key must be attached to a portable radiation survey meter by a chain or cable. The lock must be designed so that the key cannot be removed if the source is in an unshielded position. The door to the radiation room must require the same key. The key must be in the possession of a person entering the radiation room.

(2) The console of a panoramic irradiator must have a source position indicator which indicates when the sources are in the fully shielded position, when they are in transit, and when they are exposed.

(3) The control console of a panoramic irradiator must have a control which promptly returns the sources to the shielded position.

(4) Each control for a panoramic irradiator must be labeled clearly as to its function.

(5) Controls for a panoramic irradiator must be color-coded or illuminated as follows:

(a) Red represents emergency or critical information;

(b) Yellow or orange represents caution, no emergency but some function taking place to be aware of; and

(c) Green or blue represents normal or safe functioning or information.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1510.

- 64E-5.1411 Irradiator Pools.

(1) For licenses initially issued after 8-14-96, irradiator pools must possess a watertight stainless steel liner or a liner metallurgically compatible with other components in the pool or be constructed so that there is a low likelihood of substantial leakage and have a surface designed to facilitate decontamination and must include a means of safely storing sources during repairs of the pool.

(2) For licenses initially issued after 8-14-96, irradiator pools must have no penetration more than 0.5 meter below the normal low water level which could allow water to drain out of the pool. Pipes which have intakes more than 0.5 meter below the normal low water level must have siphon breakers to prevent the siphoning of the pool.

(3) A means must be provided to replenish water losses from the pool.

(4) An audible and visible indicator must be provided to indicate if the pool water level is below the normal low water level or above the normal high water level.

(5) Irradiator pools must be equipped with a purification system designed to maintain the water during normal operation at a level of conductance not exceeding 20 microsiemens per centimeter and with a clarity so the sources can be seen clearly.

(6) A physical barrier such as a railing or cover must be used around irradiator pools during normal operation to prevent personnel from accidentally falling into the pool. The barrier can be removed during maintenance, inspection, and service operations.

(7) If long-handled tools or poles are used in irradiator pools, the radiation dose rate on the handling areas of the tools must not exceed 2 millirem (0.02 millisievert) per hour.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1511.

- 64E-5.1412 Source Rack Protection.

If the product to be irradiated moves on a product conveyor system, the source rack and the mechanism which moves the rack must be protected by a barrier or guides to prevent products and product carriers from hitting or touching the rack or mechanism. A collision alarm system on the protective barrier must cause an alarm at the control console notifying the operator that a collision between the barrier and product conveyor has occurred.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1512.

- 64E-5.1413 Power Failures.

(1) If electrical power at a panoramic irradiator is lost for longer than 10 seconds, the sources must return automatically to the shielded position.

(2) The lock on the door of the radiation room of a panoramic irradiator must not be deactivated by a power failure.

(3) During a power failure, the area around the pool of an irradiator shall not be entered without using an operable and calibrated radiation survey meter.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1513.

- 64E-5.1414 Design Requirements.

Irradiators whose construction begins after 8-14-96 must meet the design requirements of this section. The requirements must be met before the start of the construction of the specific component, but do not have to be met before submitting a license application. After construction of the irradiator is completed, the licensee must submit to the department copies of as-built construction drawings signed by the architect and the licensee.

(1) Panoramic irradiators shall meet the following design requirements:

(a) Shielding. The licensee shall design shielding walls to meet generally accepted building code requirements for reinforced concrete and shall design the walls, wall penetrations, and entrance ways to meet the radiation shielding requirements of Rule 64E-5.1407, F.A.C. If the irradiator will use more than 2 × 1017 becquerels (5 million curies) of activity, the licensee shall evaluate the effects of heating of the shielding walls by the irradiator sources.

(b) Foundations. The licensee shall design the foundation with consideration given to soil characteristics to ensure it is adequate to support the weight of the facility.

(c) Source Rack. The licensee shall determine that source rack drops due to loss of power will not damage the source rack and that source rack drops due to failure of cables or alternate means of support will not cause loss of integrity of sealed sources. The licensee shall review the design of the mechanism which moves the sources to assure that the likelihood of a stuck source is low and that if the rack sticks it can be freed without causing radiation overexposure of personnel.

(d) Access Control. The licensee shall verify from the design and logic diagram that the access control system will meet the requirements of Rule 64E-5.1406, F.A.C.

(e) Fire Protection. The licensee shall verify that the number, design, locations and spacing of the smoke and heat detectors and extinguishing system are appropriate to detect fires and that the detectors are protected from mechanical and radiation damage. The licensee shall verify that the design of the fire extinguishing system provides the necessary discharge patterns, densities, and flow characteristics for complete coverage of the radiation room and that the system is protected from mechanical and radiation damage.

(f) Source Return. The licensee shall verify that the source rack will be returned automatically to the fully shielded position if off-site power is lost for more than 10 seconds. If a component of the return mechanism fails, the design must allow for accomplishing the return without causing radiation overexposures of personnel.

(g) Seismicity. For panoramic irradiators to be built in seismic areas, the licensee shall design the reinforced concrete radiation shields to retain their integrity in the event of an earthquake by designing to seismic requirements of an appropriate source such as American Concrete Institute Standard 318-89, “Building Code Requirements for Reinforced Concrete,” Chapter 21, “Special Provisions for Seismic Design,” or local building codes.

(h) Wiring. The licensee shall verify that electrical wiring and electrical equipment in the radiation room are selected to minimize failures due to prolonged exposure to radiation.

(2) Pool irradiators shall meet the following design requirements:

(a) Pool Integrity. The licensee shall design the pool to assure that it is leak-resistant, that it is strong enough to bear the weight of the pool water and shipping casks, that a dropped cask would not fall on sealed sources, that all penetrations meet the requirements of subsection 64E-5.1411(2), F.A.C., and that metal components are metallurgically compatible with other components in the pool.

(b) Water Handling System. The licensee shall design the water purification system to meet the requirements of subsection 64E-5.1411(5), F.A.C. The system must be designed so that water leaking from the system does not drain to unrestricted areas without being monitored. The licensee shall design the water chiller system so that it shall compensate adequately for the amount of heat generated by the sealed sources. The water handling system must have remote controls capable of safely operating a contaminated system.

(c) Source rack. The licensee shall verify that there are no crevices on the source or between the source and source holder that would promote corrosion on a critical area of the source. The lift mechanisms for the source rack and source transport cask must be of designed working and breaking strength to lift safely a source transport cask and sources into and out of the irradiator pool.

(3) All irradiators shall meet the following design requirements:

(a) Radiation Monitors. The licensee shall evaluate the location and sensitivity of the monitor to detect sources carried by the product conveyor system as required by subsection 64E-5.1409(1), F.A.C. The licensee shall verify that the product conveyor will stop before a source on the product conveyor causes a radiation overexposure. For pool irradiators, the licensee shall verify that the radiation monitor on the water purification system is located near the area in which elevated radiation levels will be expected.

(b) Product carriers. For irradiators using product carriers, the design of the carrier shall prevent the carrier from opening or coming into contact with the source rack or protective barrier. The design shall be submitted to the department for approval.

(c) Floor penetrations. No floor penetrations, including expansion joints, floor joints and drains, shall allow the uncontrolled release of water from the radiation room that has not been analyzed for its radioactive content.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1514.

- 64E-5.1415 Construction Control.

The requirements of this section must be met before loading sources.

(1) Panoramic irradiators shall meet the following construction requirements:

(a) Shielding. The licensee shall monitor the construction of the shielding to verify that it meets design specifications and generally accepted building code requirements for reinforced concrete.

(b) Foundations. The licensee shall monitor the construction of the foundations to verify that they meet design specifications.

(c) Source Rack. The licensee shall test the movement of the source racks for proper operation before source loading. Testing must include source rack lowering due to simulated loss of power. For all irradiators with product conveyor systems, the licensee shall observe and test the operation of the conveyor system to assure that the requirements in Rule 64E-5.1412 and paragraph 64E-5.1414(3)(b), F.A.C., are met for protection of the source racks and the mechanism which moves the rack. Testing must include any limit switches and interlocks used to protect the source rack and the mechanism which moves the rack from moving product carriers.

(d) Access Control. The licensee shall test the access control system to assure that it functions as designed and that all alarms, controls, and interlocks work properly.

(e) Fire Protection. The licensee shall verify the ability of the heat and smoke detectors to detect a fire, to activate alarms, and to cause the source rack to become fully shielded automatically. The licensee also shall verify the operability of the fire suppression or extinguishing system.

(f) Source Return. The licensee shall demonstrate that the source racks can be returned to their fully shielded position without off-site power.

(g) Computer Systems. If a computer is used to control the access control system, the licensee shall demonstrate that the computer and the access control system will operate as planned if off-site power is lost by attempting to defeat the access control system in as many ways as possible. The computer must have suitable security features which prevent an irradiator operator from commanding the computer to override the access control system when it is required to be operable.

(h) Wiring. The licensee shall verify that the electrical wiring and electrical equipment that were installed meet the design specifications.

(2) Pool irradiators shall meet the following construction requirements:

(a) Pool Integrity. The licensee shall test the integrity of the pool and verify that the pool meets the design specifications. The licensee shall verify that penetrations and water intakes meet the requirements of subsection 64E-5.1411(2), F.A.C.

(b) Water Handling System. The licensee shall verify that the water purification system, the conductivity meter and the water level alarms operate properly.

(3) Radiation Monitors. For all irradiators, the licensee shall verify the proper operation of the monitor to detect sources carried on the product conveyor system and related alarms and interlocks required by subsection 64E-5.1409(1), F.A.C. For pool irradiators, the licensee shall verify the proper operation of the radiation monitor on the water purification system and the related alarms and interlocks required by subsection 64E-5.1409(2), F.A.C. For underwater irradiators, the licensee shall verify the proper operation of the pool monitor, alarms, and interlocks required by subsection 64E-5.1409(3), F.A.C.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1515.

- 64E-5.1416 Training.

(1) Before an individual is permitted to operate an irradiator without a supervisor present, the individual must be instructed in the following:

(a) The fundamentals of radiation protection applied to irradiators, including:

1. The difference between external radiation and radioactive contamination;

2. Units of radiation dose;

3. The department’s dose limits;

4. Why large radiation doses must be avoided;

5. How shielding and access controls prevent large doses;

6. How an irradiator is designed to avoid contamination;

7. The use of survey meters and personnel dosimeters;

8. Other radiation safety features of an irradiator; and

9. The basic function of the irradiator.

(b) The requirements of this part and Part IX of these rules;

(c) The operation of the irradiator;

(d) Licensee operating and emergency procedures which the individual is responsible for performing; and

(e) Case histories of accidents or problems involving irradiators similar to those to be used by the individual.

(2) Before an individual is permitted to operate an irradiator without a supervisor present, the individual shall pass a written test on the instruction received, consisting primarily of questions based on the licensee’s operating and emergency procedures that the individual is responsible for performing and other operations necessary to operate the irradiator safely without supervision.

(3) Before an individual is permitted to operate an irradiator without a supervisor present, the individual must have received on-the-job or simulator training in the use of the irradiator as described in the license application. The individual also shall demonstrate the ability to perform those portions of the operating and emergency procedures that he or she is to perform.

(4) The licensee shall conduct safety reviews and emergency drills as described below for irradiator operators at least annually. The licensee shall give each operator a brief written test on the information. Each safety review must include, to the extent appropriate, each of the following:

(a) Changes in operating and emergency procedures since the last review;

(b) Changes in rules and license conditions since the last review;

(c) Reports on recent accidents, mistakes, or problems which have occurred at irradiators;

(d) Relevant results of inspection of operator safety performance;

(e) Relevant results of the facility’s inspection and maintenance checks; and

(f) A drill to practice an emergency or abnormal event procedure.

(5) The licensee shall evaluate the safety performance of each irradiator operator at least annually to ensure that rules, license conditions, and operating and emergency procedures are followed. The licensee shall discuss the results of the evaluation with the operator and shall instruct the operator on how to correct any mistakes or deficiencies observed.

(6) Individuals who will be permitted unescorted access to the irradiators but who have not received the training required for operators and the radiation safety officer shall be trained and tested in precautions they should take to avoid radiation exposure, procedures or parts of procedures in Rule 64E-5.1418, F.A.C., which they are expected to perform or comply with, and the proper response to alarms required in this part. Tests can be oral.

(7) Individuals who must be prepared to respond to alarms required by Rules 64E-5.1406, 64E-5.1408, 64E-5.1409, 64E-5.1411, and 64E-5.1412, F.A.C., shall be trained and tested on how to respond. Each individual shall be retested at least once a year. Tests can be oral.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1516.

- 64E-5.1417 Operating and Emergency Procedures.

(1) The licensee shall have and follow written operating procedures for the following:

(a) Operation of the irradiator, including entering and leaving the radiation room;

(b) Use of personnel dosimeters;

(c) Surveying the shielding of panoramic irradiators;

(d) Monitoring pool water for contamination while the water is in the pool and before release of pool water to unrestricted areas;

(e) Leak testing of sources;

(f) Inspection and maintenance checks required by Rule 64E-5.1422, F.A.C.;

(g) Loading, unloading, and repositioning sources, if to be performed by the licensee;

(h) Inspection of movable shielding required by paragraph 64E-5.1406(1)(h), F.A.C., if applicable; and

(i) Security precautions while sources are stored outside the radiation room. Sealed sources must be moved into the radiation room within 48 hours of receipt unless the department is notified in writing that extenuating circumstances do not allow for source loading within the prescribed 48 hour period.

(2) The licensee shall have and follow emergency or abnormal event procedures for the following:

(a) Sources stuck in the unshielded position;

(b) Personnel overexposures;

(c) A radiation alarm from the product exit portal monitor or pool monitor;

(d) Detection of leaking sources, pool contamination, or alarm caused by contamination of pool water;

(e) A low water level alarm, a high water level alarm, an abnormal water loss, or leakage from the source storage pool;

(f) A prolonged loss of electrical power;

(g) A fire alarm or explosion in the radiation room;

(h) An alarm indicating unauthorized entry into the radiation room, the area around the pool, or another alarmed area;

(i) Natural phenomena, including an earthquake, a hurricane, a tornado, flooding, sinkhole formation, or other phenomena; and

(j) The jamming of automatic conveyor systems or an alarm indicating a collision between the barrier and product conveyor.

(3) The licensee can revise operating and emergency procedures only with departmental approval.

(4) The licensee shall provide and coordinate current emergency procedures annually with the local police, fire department, and civil authorities, including notification of responsible individuals and places of emergency treatment.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1517.

- 64E-5.1418 Personnel Monitoring.

(1) Irradiator operators shall wear either a film badge, OSLD or a TLD while operating a panoramic irradiator or while in the area around the pool of an underwater irradiator. The film badge, OSLD, and TLD processor must be accredited by NVLAP for high energy photons in the normal and accident dose ranges. Each film badge, OSLD, and TLD must be assigned to and worn by only one individual. Film badges must be replaced at least monthly and OSLDs and TLDs must be replaced at least quarterly. After replacement, each film badge, OSLD, and TLD must be processed promptly.

(2) Other individuals who enter the radiation room of a panoramic irradiator shall wear a dosimeter, which can be a pocket dosimeter. For groups of visitors, only two people are required to wear dosimeters. Date of entry, all names and total dose must be recorded. If pocket dosimeters are used to meet the requirements of this paragraph, a check of their response must be done at least annually. Acceptable dosimeters must read within 30 percent of the true radiation dose.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1518, Amended 10-8-00.

- 64E-5.1419 Radiation Surveys.

(1) Before the facility starts operation, the following radiation surveys must be performed:

(a) A radiation survey of the area above the pool after the sources are loaded and in the shielded position; and

(b) A survey of the area outside the shielding of the radiation room of a panoramic irradiator with the sources in the exposed position.

(2) If the surveys indicate that radiation levels specified in Rule 64E-5.1407, F.A.C., are exceeded, the shielding must be repaired to comply with the dose rate requirement in Rule 64E-5.1407, F.A.C., before operation of the facility can start.

(3) Radiation surveys described in (1) above must be performed after new sources are loaded and after any modifications which might increase dose rates are made to the radiation room, shielding or structure and at intervals not to exceed 3 years.

(4) Portable radiation survey meters used to meet the requirements of subsections (1) and (3) of this section and the requirements of subsections 64E-5.1413(3) and 64E-5.1424(1), F.A.C., must be calibrated at least annually to an accuracy of 20 percent for the gamma energy of the sources in use. The calibration must be done at two points on each scale or, for digital instruments, at one point per decade over the range that will be used. Portable radiation survey meters must be of a type that does not saturate and read zero at high radiation dose rates.

(5) Water from the irradiator pool or other potentially contaminated liquids and sediments from pool vacuuming must be monitored for radioactive contamination before release to unrestricted areas. Radioactive concentrations must not exceed those specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table II, Column 2, or Table III, as applicable. The lower limit of detection for the measurements must be below those concentrations.

(6) Resins to be released for unrestricted use must be monitored before release in an area with a background level less than 0.05 millirem (0.0005 millisievert) per hour. The resins can be released only if the survey does not detect radiation levels above background radiation levels. The survey meter must be capable of detecting radiation levels of 0.05 millirem (0.0005 millisievert) per hour.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1519, Amended 12-26-13.

- 64E-5.1420 Detection of Leaking or Contaminated Sources.

(1) Each dry-source-storage sealed source must be tested for leakage at least every 6 months using a leak test kit or a method approved by the department, U.S. Nuclear Regulatory Commission, agreement state or licensing state. The analysis must be capable of detecting the presence of 0.005 microcurie (185 becquerels) of radioactive material and must be performed by a person approved by the department, U.S. Nuclear Regulatory Commission, agreement state or licensing state to perform the analysis.

(2) For pool irradiators, the pool water must be checked for contamination each day the irradiator operates. The check must be done by using an on-line radiation monitor on a pool water circulating system as described in subsection 64E-5.1410(2), F.A.C., or by analysis of pool water. If a check for contamination is done by analysis of pool water, the results of the analysis must be available within 24 hours. If the licensee uses a radiation monitor on a pool water circulating system, the detection above normal radiation levels must activate an alarm. The alarm set-point must be set as low as practical but high enough to avoid false alarms. The licensee can reset the alarm set-point to a higher level if necessary to operate the pool water purification system to clean up contamination in the pool if specifically provided for in written emergency procedures.

(3) The licensee shall have written procedures and equipment available for the detection, isolation and removal of leaking sources.

(4) If a leaking source is detected, the licensee shall remove the leaking source from service and have it decontaminated, repaired, or disposed of by a licensee of the Department, NRC, Agreement State or Licensing State authorized to perform these functions. The licensee shall check its personnel, equipment, facilities, and irradiated product promptly for radioactive contamination. No product shall be shipped until the product has been checked and found free of contamination. If a product has been shipped that could have been contaminated inadvertently, the licensee shall arrange to locate and survey that product for contamination. If any personnel are contaminated, decontamination must be performed promptly. If contaminated equipment, facilities, or products are found, the licensee shall have them decontaminated or disposed of by a licensee of the Department, NRC, Agreement State or Licensing State authorized to perform these functions. If the pool water is contaminated, the licensee shall clean the pool water until the contamination levels do not exceed the appropriate concentration in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table II, Column 2.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1520, Amended 12-26-13.

- 64E-5.1421 Inspection and Maintenance.

(1) The licensee shall perform inspection and maintenance checks that include, as a minimum, each of the following at the frequency specified in the license or license application:

(a) Operation of each aspect of the access control system required by Rule 64E-5.1406, F.A.C.;

(b) Functioning of the source position indicator required by subsection 64E-5.1410(2), F.A.C.;

(c) Operation of the radiation monitor on the pool water purification system using a radioactive check source to detect radioactive contamination in pool water as required by subsection 64E-5.1409(2), F.A.C.;

(d) Operation of the over-pool radiation monitor at underwater irradiators as required by subsection 64E-5.1409(3), F.A.C.;

(e) Operation of the product exit monitor required by subsection 64E-5.1409(1), F.A.C.;

(f) Operation of the emergency source return control required by subsection 64E-5.1410(3), F.A.C.;

(g) Leak-tightness of systems through which pool water circulates (visual inspection);

(h) Operation of the heat and smoke detectors and extinguisher system required by Rule 64E-5.1408, F.A.C., but without turning on extinguishers;

(i) Operation of the means of pool water replenishment required by subsection 64E-5.1411(3), F.A.C.;

(j) Operation of the visible indicator of low and high pool water level required by subsection 64E-5.1411(4), F.A.C.;

(k) Operation of the intrusion alarm required by subsection 64E-5.1406(2), F.A.C.;

(l) Functioning and wear on the system, mechanisms, and cables used to raise and lower sources;

(m) Condition of the barrier to prevent products from hitting the sources or source mechanism and the operation of the collision alarm system, as required by Rule 64E-5.1412, F.A.C.;

(n) Amount of water added to the pool to determine if the pool is leaking;

(o) Electrical wiring on required safety systems for radiation damage;

(p) Pool water conductivity measurements and analysis as required by Rule 64E-5.1422, F.A.C.;

(q) Condition of the product carriers; and

(r) Operation of the siphon breakers in the pool plumbing.

(2) Malfunction and defects found during inspection and maintenance checks must be repaired without undue delay.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1521.

- 64E-5.1422 Pool Water Purity.

(1) Pool water purification systems must be run sufficiently to maintain the conductivity of the pool water below 20 microsiemens per centimeter under normal circumstances. If pool water conductivity rises above 20 microsiemens per centimeter, the licensee shall take prompt actions to lower the pool water conductivity and shall take corrective actions to prevent future recurrences.

(2) The licensee shall measure the pool water conductivity frequently enough, but no less than weekly, to assure that the conductivity remains below 20 microsiemens per centimeter. Conductivity meters must be calibrated at least annually.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1522.

- 64E-5.1423 Attendance During Operation.

(1) An operator and at least one other individual trained and prepared to render or summon assistance promptly if the access control alarm sounds shall be present on site whenever a panoramic irradiator is operated using an automatic product conveyor system and whenever the product is moved into or out of the radiation room when the irradiator is operated in a batch mode.

(2) At an underwater irradiator, an operator must be present whenever the product is moved into or out of the pool. Individuals who move the product into or out of the pool of an underwater irradiator need not be qualified as irradiator operators; however, they must have received the training described in subsections 64E-5.1416(6) and (7), F.A.C. Static irradiation can be performed without a person present at the facility only if the personnel access barrier around the pool is locked to prevent unauthorized entry and all required alarms are operable.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1523.

- 64E-5.1424 Entering and Leaving the Radiation Room.

(1) Upon first entering the radiation room of a panoramic irradiator after an irradiation, the irradiator operator shall use a survey meter to determine that the source has returned to its fully shielded position. The operator shall check the functioning of the survey meter with a radiation check source before entry. The response of the survey meter to the radiation check source must be consistent and reproducible.

(2) Before exiting from and locking the door to the radiation room of a panoramic irradiator before a planned irradiation, the irradiator operator shall perform the following:

(a) Inspect visually the entire radiation room to verify that no one else is in it; and

(b) Activate a control in the radiation room which permits the sources to be moved from the shielded position only if the door to the radiation room is locked within a preset time after setting the control.

(c) During a power failure, the area around the pool of an underwater irradiator cannot be entered without using an operable and calibrated radiation survey meter unless the over-the-pool monitor required by subsection 64E-5.1409(3), F.A.C., is operating with backup power.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1524.

- 64E-5.1425 Irradiation of Explosive or Highly Flammable Materials.

(1) Irradiation of explosive materials is prohibited unless the licensee has received prior written authorization from the department. Authorization will not be granted unless the licensee can demonstrate that detonation of the explosive would not rupture the sealed sources, injure personnel, damage safety systems, or cause radiation overexposures of personnel.

(2) Irradiation of more than small quantities of highly flammable material with a flash point below 140 degrees Fahrenheit is prohibited in panoramic irradiators unless the licensee has received prior written authorization from the department. Authorization will not be granted unless the licensee can demonstrate that a fire in the radiation room could be controlled without damage to sealed sources or safety systems and without radiation overexposures of personnel.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1525.

- 64E-5.1426 Records and Retention Periods.

(1) The licensee shall maintain the following records at the irradiator for the periods specified:

(a) A copy of the license application and the license authorizing the licensee to operate the facility until a new license is issued;

(b) Records of each individual’s training, tests, and safety reviews provided to meet the requirements of Rule 64E-5.1416, F.A.C., until 3 years after the individual terminates work;

(c) Records of the annual evaluation of the safety performance of irradiator operators required by subsection 64E-5.1416(5), F.A.C., for 3 years after the evaluation;

(d) An up-to-date copy of the operating and emergency procedures required by Rule 64E-5.1417, F.A.C., until superseded or the department terminates the license;

(e) Personnel monitoring results required by Rule 64E-5.1418, F.A.C., until the department terminates the license;

(f) Records of radiation surveys required by Rule 64E-5.1420, F.A.C., for 3 years from the date of the survey;

(g) Records of radiation survey meter calibrations required by Rule 64E-5.1419, F.A.C., and pool water conductivity meter calibrations required by subsection 64E-5.1422(2), F.A.C., until 3 years from the date of the calibration;

(h) Records of the results of leak tests required by Rule 64E-5.1420, F.A.C., and the results of contamination checks required by subsection 64E-5.1420(2), F.A.C., for 3 years from the date of the leak tests;

(i) Records of inspection and maintenance checks required by Rule 64E-5.1421, F.A.C., for 3 years;

(j) Records of major malfunctions, significant defects, operating difficulties or irregularities, and major operating problems that involve required radiation safety equipment for 3 years after repairs are completed;

(k) Records of the receipt, transfer and disposal of all licensed sealed sources as required by Rule 64E-5.103, F.A.C.;

(l) An inventory of all licensed sealed sources until the irradiator is decommissioned. The inventory must include for each sealed source the following:

1. The date received;

2. The person from whom it was received;

3. The model of the source;

4. The serial number of the source;

5. The radionuclide in the source;

6. The activity of the source as supplied from the manufacturer and the date of the assigned activity;

7. An up-to-date location of the source;

8. Information on leaking or damaged sources and any actions taken to decontaminate or repair those sources;

9. The date disposed of; and

10. The person to whom the source was transferred.

(m) Records on the design checks required by Rule 64E-5.1414, F.A.C., and the construction control checks required by Rule 64E-5.1415, F.A.C., until the license is terminated. The records must be signed and dated. The title or qualification of the person signing must be included;

(n) Records of water added to the pool as required by paragraph 64E-5.1421(1)(n), F.A.C., for 3 years;

(o) Records related to decommissioning the irradiator as required by subparagraph 64E-5.214(4)(c)2., F.A.C.;

(p) Records of annual notification to local police, fire department and civil authorities of the current emergency procedures, responsible individuals and places of emergency treatment for 3 years;

(q) Records of conductivity meter calibration as required by subsection 64E-5.1422(2), F.A.C., for 3 years; and

(r) Records of written approval of access to restricted areas as required by subsection 64E-5.1407(1), F.A.C., for 3 years.

(2) Records required by this section shall be available for inspection by the department.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1526.

- 64E-5.1427 Reports and Notifications.

(1) In addition to the other reporting requirements in these rules, each licensee shall notify the department as soon as possible but not later than 4 hours after the discovery of the following events:

(a) Threats of violence or acts of terrorism against the operation of this facility;

(b) Fire or explosion in the radiation room;

(c) Detection of radiation by the product portal monitors;

(d) Detection of any radioactive contamination;

(e) Sources stuck in an unshielded position;

(f) Damage to source racks;

(g) Failure of the cable or drive mechanism used to move the source racks;

(h) Inoperability of the access control system;

(i) Structural damage to the pool liner or walls;

(j) Abnormal water loss or leakage from the source storage pool; or

(k) Pool water conductivity exceeding 100 microsiemens per centimeter.

(2) The licensee must submit a written report within 30 days for any reports required by subsection (1) above. The report must describe the event, what caused the event to the extent known, and corrective actions to prevent recurrence taken up to the time the report is made.

(3) The licensee shall notify individuals of their exposure to radiation or radioactive materials as required by Rule 64E-5.903, F.A.C.

(4) The licensee shall notify the department at least 14 days before a source loading or unloading.

Rulemaking Authority 404.051(4) FS. Law Implemented 404.051(1), (5), (6), 404.061, 404.081, 404.141 FS. History–New 8-14-96, Formerly 10D-91.1527.

- 64E-5.1501 Purpose and Scope.

(1) The packaging and transportation of radioactive material are also subject to the requirements of other agencies such as the U.S. Department of Transportation, the U.S. Nuclear Regulatory Commission and the U.S. Postal Service. The requirements of this part are in addition to, and not in substitution for, other requirements.

(2) Determinations and listings of A1 and A2 values are found in 10 C.F.R., Part 71, Appendix A, as published on 01/01/2012 which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03457 or at http://www.gpo.gov/fdsys/pkg/CFR-2012-title10-vol2/pdf/CFR-2012-title10-vol2-part71.pdf.

(3) The regulations in this part apply to any licensee authorized by specific or general license issued by the department to receive, possess, use, or transfer licensed material, if the licensee delivers that material to a carrier for transport, transports the material outside the site of usage as specified in the license, or transports that material on public highways. No provision of this part authorizes possession of licensed material.

(4) Definition of terms used in this part are those listed in Rule 64E-5.1502, F.A.C., as described in 49 C.F.R. and 10 C.F.R. 71.4, except that whenever a definition refers to evaluation or approval by the U.S. Department of Transportation or NRC, and such evaluation or approval is within the jurisdiction of the State of Florida as an Agreement State, the Department shall perform the evaluation or approval.

Rulemaking Authority 404.051, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.20(1) FS. History–New 7-17-85, Amended 5-15-96, Formerly 10D-91.2001, Amended 2-28-08, 12-26-13.

- 64E-5.1502 Transportation of Radioactive Material.

(1) No person shall deliver radioactive material to a carrier for transport or transport radioactive material except as authorized in a general license or specific license issued by the Department or as exempted in Rule 64E-5.1503, F.A.C.

(2) Each licensee who transports radioactive material outside of the confines of his facility or other place of use, or who offers radioactive material to a carrier for transport shall:

(a) Comply with the current applicable requirements, appropriate to the mode of transport, of 49 C.F.R. Parts 107, 171-180, 383, 390-397, published on 01/01/2012 which is herein incorporated by reference and can be obtained from the internet http://www.flrules.org/Gateway/reference.asp?No=Ref-03458, https://www.flrules.org/Gateway/reference.asp?No=Ref-03473, https://www.flrules.org/Gateway/reference.asp?No=Ref-03474, https://www.flrules.org/Gateway/reference.asp?No=Ref-03475, and https://www.flrules.org/Gateway/reference.asp?No=Ref-03476 or at http://www.gpo.gov/fdsys/browse/collectionCfr.action?collectionCode=CFR&searchPath=Title+49%2FSubtitle+B&oldPath=Title+49&isCollapsed=true&selectedYearFrom=2012&ycord=1546 and 10 C.F.R. Part 71 published on 01/01/2012 which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03459 or at http://www.gpo.gov/fdsys/pkg/CFR-2012-title10-vol2/pdf/CFR-2012-title10-vol2-part71.pdf, and 10 C.F.R. Parts 73.72 through 73.74 published on 01/01/2012 which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03460 or at http://www.gpo.gov/fdsys/browse/collectionCfr.action?collectionCode=CFR&searchPath=Title+10%2FChapter+I%2FPart+73%2FSubjgrp&oldPath=Title+10%2FChapter+I%2FPart+73%2FSubjgrp&isCollapsed=true&selectedYearFrom=2012&ycord=1772.

(b) Establish procedures for safely opening and closing packages in which radioactive material is transported and to assure that, prior to the delivery to a carrier for transport, each package is properly closed for transport; and

(c) Assure that any special instructions needed to safely open the package are sent to or have been made available to the consignee.

(d) The licensee shall comply with U.S. Department of Transportation and NRC regulations in the following areas:

1. Packaging, 49 C.F.R. part 173, subparts A, B, and I;

2. Marking and labeling, 49 C.F.R. part 172, subpart D, §§172.400 through 172.407, §§172.436 through 172.441 of subpart E;

3. Placarding, 49 C.F.R. part 172, subpart F, especially §§172.500 through 172.519 and 172.556, and appendices B and C;

4. Accident reporting, 49 C.F.R. part 171, §§171.15 and 171.16;

5. Shipping papers and emergency information, 49 C.F.R. part 172, subparts C and G;

6. Hazardous material employee training, 49 C.F.R. part 172, subpart H;

7. Security plans, 49 C.F.R. part 172, subpart I;

8. Hazardous material shipper/carrier registration, 49 C.F.R. part 107, subpart G;

9. Definitions, 10 C.F.R. 71.4;

10. Transportation of licensed material, 10 C.F.R. 71.5;

11. Exemptions for low level material, 10 C.F.R. 71.14(a);

12. General license, NRC-approved package, 10 C.F.R. 71.17;

13. Previously approved package, 10 C.F.R. 71.19(a) and (b);

14. General license, U.S. Department of Transportation specification container material, 10 C.F.R. 71.20;

15. General license, Use of foreign approved package, 10 C.F.R. 71.21;

16. General license, Fissile material, 10 C.F.R. 71.22;

17. External radiation standards for all packages, 10 C.F.R. 71.47;

18. Assumptions as to unknown properties, 10 C.F.R. 71.83;

19. Preliminary determinations, 10 C.F.R. 71.85;

20. Routine determinations, 10 C.F.R. 71.87;

21. Air transportation of plutonium, 10 C.F.R. 71.88;

22. Opening instructions, 10 C.F.R. 71.89;

23. Advance notification of shipment of irradiated reactor fuel and nuclear waste, 10 C.F.R. 71.97;

24. Quality assurance requirements, 10 C.F.R. 71.101(a), (b), (c), (f) and (g);

25. Quality assurance organization, 10 C.F.R. 71.103;

26. Quality assurance program, 10 C.F.R. 71.105;

27. Exemption of physicians, 10 C.F.R. 71.13;

28. Handling storage and shipping control, 10 C.F.R. 71.127;

29. Inspection tests and operating status, 10 C.F.R. 71.129;

30. Nonconforming materials parts or components, 10 C.F.R. 71.131;

31. Corrective action, 10 C.F.R. 71.13;

32. Quality assurances records, 10 C.F.R. 71.135;

33. Audits, 10 C.F.R. 71.137;

34. Appendix A to Part 71; and

35. General license plutonium beryllium special form material.

(e) The licensee shall also comply with U.S. Department of Transportation regulations pertaining to the following modes of transportation:

1. Rail, 49 C.F.R. part 174, subparts A through D and K;

2. Air, 49 C.F.R. part 175;

3. Vessel, 49 C.F.R. part 176, subparts A through F and M; and

4. Public Highway, 49 C.F.R. part 177 and parts 390 through 397.

(3) If U.S. Department of Transportation regulations are not applicable to a shipment of licensed material, the licensee shall conform to the standards and requirements of the U.S. Department of Transportation specified in subsection (2) of this section to the same extent as if the shipment or transportation were subject to U.S. Department of Transportation regulations. A request for modification, waiver, or exemption from those requirements, and any notification referred to in those requirements, must be filed with, or made to, the Department.

Rulemaking Authority 404.051, 404.061, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.061(2), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2003, Amended 10-8-00, 9-28-06, 2-28-08, 12-26-13.

- 64E-5.1503 Exemptions.

(1) Common and contract carriers, freight forwarders, and warehousemen who are subject to the requirements of the U.S. Department of Transportation in 49 C.F.R. Parts 170 through 189 or the U.S. Postal Service in the Postal Service Manual (Domestic Mail Manual), Section 124.3 incorporated by reference, 39 C.F.R. Part 111.1 (1974), are exempt from these regulations to the extent that they transport or store radioactive material in the regular course of their carriage for another or storage incident thereto. Common and contract carriers who are not subject to the requirements of the U.S. Department of Transportation or U.S. Postal Service are subject to Rule 64E-5.1501, F.A.C., and other applicable sections of these regulations.

(2) Any licensee is exempt from the requirements of this part to the extent that he delivers to a carrier for transport a package containing radioactive material having a Rulemaking activity not greater than 0.002 microcurie (74 Bq) per gram.

Rulemaking Authority 404.051, 404.061, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.061(2), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2004.

- 64E-5.1504 General Licenses for Carriers.

(1) A general license is hereby issued to any common or contract carrier not exempt under Rule 64E-5.1503, F.A.C., to receive, possess, transport and store radioactive material in the regular course of their carriage for another or storage incident thereto, provided the transportation and storage is in accordance with the applicable requirements, appropriate to the mode of transport, of the U.S. Department of Transportation insofar as such requirements relate to the loading and storage of packages, placarding of the transporting vehicle, and incident reporting. Any notification of incidents referred to in these U.S. Department of Transportation requirements shall also be filed with, or made to, the Department.

(2) A general license is hereby issued to any private carrier to transport radioactive material, provided the transportation is in accordance with the applicable requirements, appropriate to the mode of transport, of the U.S. Department of Transportation insofar as such requirements relate to the loading and storage of packages, placarding of the transporting vehicle, and incident reporting. Any notification of incidents referred to in these U.S. Department of Transportation requirements shall be filed with, or made to, the Department.

(3) Persons who transport radioactive material pursuant to the general license in subsection 64E-5.1504(1) or (2), F.A.C., are exempt from the requirements of Parts III and IX to the extent that they transport radioactive material.

Rulemaking Authority 404.051, 404.061, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.061(2), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2005.

- 64E-5.1505 Routine Determinations.

Prior to each shipment of radioactive material, the licensee shall ensure that the package with its contents satisfies the applicable requirements of this part and of the license. The licensee shall determine that:

(1) The package is proper for the contents to be shipped;

(2) The package is in an unimpaired physical condition except for superficial defects such as marks or dents;

(3) Each closure device of the packaging, including any required gasket, is properly installed and secured and free of defects;

(4) Any system for containing liquid is adequately sealed and has adequate space or other specified provision for expansion of the liquid;

(5) Any pressure relief device is operable and set in accordance with written procedures;

(6) The package has been loaded and closed in accordance with written procedures;

(7) Any structural part of the package which could be used to lift or tie down the package during transport is rendered inoperable for that purpose unless it satisfies design requirements specified by the U.S. Nuclear Regulatory Commission;

(8) The level of nonfixed or removable radioactive contamination on the external surfaces of each package presented for shipment is as low as reasonably achievable and shall not exceed the applicable levels specified in the regulations of the U.S. Department of Transportation, 49 C.F.R. 173.443; and

(9) The external radiation levels around the package and around the vehicle, if applicable, will not exceed the limits specified in the regulations of the U.S. Department of Transportation, 49 C.F.R. 173.441.

Rulemaking Authority 404.051, 404.061, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.061(2), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2006.

- 64E-5.1506 Advance Notification of Shipment of Certain Quantities of Radioactive Waste.

(1) In addition to the notification requirements of Rule 64E-5.1508, F.A.C., prior to the transport of certain quantities of radioactive waste outside of the confines of the licensee’s facility or other place of use or storage, or prior to the delivery of certain quantities of radioactive waste to a carrier for transport, each licensee shall provide advance notification of such transport in writing to the governor, or governor’s designee, of each state through which the waste will be transported. A list of the mailing addresses of the governors and governors’ designees is available upon request from the Director, Office of State Programs, U.S. Nuclear Regulatory Commission, Washington, D.C., 20555.

(2) Such advance notification is required only when:

(a) The radioactive waste is required to be in Type B packaging for transportation;

(b) The radioactive waste is being transported to, through, or across state boundaries to a disposal site or to a collection point for transport to a disposal site;

(c) The quantity of licensed material in a single package exceeds:

1. Five thousand curies (185 TBq) of special form radionuclides; or

2. Five thousand curies (185 TBq) of uncompressed gases of argon 41, krypton 85m, krypton 87, xenon 131m, or xenon 135; or

3. Fifty thousand curies (1.85 PBq) of argon 37, or of uncompressed gases of krypton 85 or xenon 133, or of hydrogen 3 as a gas, as luminous paint, or adsorbed on solid material; or

4. Twenty curies (740 GBq) of other nonspecial form radionuclides for which A02 is less than or equal to 4 curies (148 GBq); or

5. Two hundred curies (7.4 TBq) of other nonspecial form radionuclides for which A02 is greater than 4 curies (148 GBq); and

(d) The quantity of spent nuclear fuel is less than that subject to advance notification requirements of 10 C.F.R. Part 73.

(3) Each advance notification required by (1) shall contain the following information:

(a) The name, address, and telephone number of the generator, carrier and receiver of the radioactive waste shipment;

(b) A description of the radioactive waste contained in the shipment as required by the regulations of the U.S. Department of Transportation;

(c) The point of origin of the shipment and the 7-day period during which departure of the shipment is estimated to occur;

(d) The 7-day period during which arrival of the shipment at state boundaries is estimated to occur;

(e) The destination of the shipment, and the 7-day period during which arrival of the shipment is estimated to occur; and

(f) A point of contact with a telephone number for current shipment information.

(4) The notification required by subsection 64E-5.1506(1), F.A.C., shall be made in writing to the office of each appropriate governor or governor’s designee and to the Department. A notification delivered by mail must be postmarked at least 7 days before the beginning of the 7-day period during which departure of the shipment is estimated to occur. A notification delivered by messenger must reach the office of the governor, or governor’s designee, at least 4 days before the beginning of the 7-day period during which departure of the shipment is estimated to occur. A copy of the notification shall be retained by the licensee for 1 year for inspection by the Department.

(5) The licensee shall notify each appropriate governor, or governor’s designee, and the Department of any changes to schedule information provided pursuant to subsection 64E-5.1506(1), F.A.C. Such notification shall be by telephone to a responsible individual in the office of the governor, or governor’s designee, of the appropriate states. The licensee shall maintain for 1 year for inspection by the Department a record of the name of the individual contacted.

(6) Each licensee who cancels a radioactive waste shipment, for which advance notification has been sent, shall send a cancellation notice to the governor, or governor’s designee, of each appropriate state and to the Department. A copy of the notice shall be retained by the licensee for 1 year for inspection by the Department.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.061(2), 404.081(1), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2007.

- 64E-5.1507 Designation of Routes for Shipment of Radioactive Waste Requiring Advance Notification.

(1) The Department may designate routes within the State of Florida for all shipments requiring advanced notification under Rule 64E-5.1506, F.A.C. Factors that the Department will consider in the designation of routes for shipments requiring advanced notifications are:

(a) Population density in the vicinity of available highways;

(b) Accident rates of available highways;

(c) Transit time;

(d) Time and day of the week during which the shipment is to occur; and

(e) Routes that may have been previously designated by other states.

(2) In lieu of the Department’s designation of routes, the interstate highway system and limited access roadways are preferred for transportation of radioactive waste requiring advanced notification pursuant to Rule 64E-5.1506, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(2), 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2008.

- 64E-5.1508 Inspection of Low-Level Radioactive Waste Shipments.

(1)(a) All persons licensed by the Department to use, manufacture, produce, transfer, transport, receive, acquire, own, process or possess radioactive materials, as well as nuclear power plants licensed by the U.S. Nuclear Regulatory Commission, and who desire to ship radioactive waste, including radioactive waste requiring advance notification as specified in Rule 64E-5.1506, F.A.C., to a low-level radioactive waste treatment, storage or disposal facility, shall notify the Department no less than 48 hours prior to departure of the shipment. The notification shall be made in writing or by telephone.

(b) Each notification must contain the following information:

1. The name, address and telephone number of the generator;

2. The name and telephone number of the contact person, designated by the generator, with whom the Department may make arrangements for the inspection;

3. The name and telephone number of the carrier;

4. The location of departure, if different from the address of the generator;

5. The scheduled date and time of departure; and

6. The proposed route to the low-level radioactive waste facility.

(2) Upon notification from a licensee or nuclear power plant licensed by the U.S. Nuclear Regulatory Commission, the Department shall dispatch an authorized representative to the licensee’s facility to inspect the shipment of radioactive waste destined for a low-level radioactive waste treatment, storage or disposal facility. The inspection shall include:

(a) Surveys of the external radiation levels of the vehicle;

(b) Inspection of package integrity, bracing and blocking, if accessible;

(c) Verification of required package marking and labeling, if accessible;

(d) Verification of required vehicle placarding; and

(e) Examination of the shipping papers for compliance with the regulations of the U.S. Department of Transportation.

(3) Licensees or nuclear power plants licensed by the U.S. Nuclear Regulatory Commission shall also provide the Department’s representative the following information or material during the course of inspection of the low-level radioactive waste shipment:

(a) Time of departure of shipment;

(b) Proposed route of the shipment to the low-level radioactive waste facility;

(c) Estimated time of arrival of the shipment at the low-level radioactive waste facility;

(d) The carrier’s name;

(e) A complete and legible copy of the bill of lading; and

(f) A complete and legible copy of the radioactive shipment manifest.

(4) If the shipment of low-level radioactive waste is found to be in compliance with the regulations of the U.S. Department of Transportation, the Department’s representative shall affix his initials on the bill of lading and the shipment may then proceed to the low-level radioactive waste facility. If the shipment of low-level radioactive waste is found to be in violation of the regulations of the U.S. Department of Transportation by the Department’s representative, the licensee shall not allow the shipment to leave the boundaries of his facility until the violation is corrected and the Department’s representative affixes his initials on the bill of lading signifying the shipment is in compliance.

(5) Licensees or nuclear power plant licensees of the U.S. Nuclear Regulatory Commission shall, within 72 hours of receiving notice of the arrival of their shipment at its destination for unloading, notify the Department of such arrival. The licensee shall also forward to the Department within 2 weeks of receiving notice of the arrival of the shipment at a destination for unloading, records of receipt and any other records indicating that a shipment was found in violation of the low-level radioactive waste treatment, storage or disposal facility’s or host state’s rules or regulations.

(6) Each generator of radioactive waste whose shipment is inspected by the department’s representative will be billed quarterly by the department a fee of $1.95 per cubic foot (0.02832 cubic meter) of waste shipped or $150.00 per shipment inspected, whichever is greater. This quarterly billing will be paid to the Department within 30 days of receipt of the bill.

Rulemaking Authority 404.051, 404.061, 404.071, 404.20 FS. Law Implemented 404.022, 404.051, 404.061, 404.071, 404.20 FS. History–New 7-17-85, Amended 7-5-88, Formerly 10D-91.2009, Amended 8-16-07.

- 64E-5.1509 Permit Requirements.

(1) Any carrier who transports low-level radioactive waste destined for a low-level radioactive waste treatment, storage or disposal facility, prior to entrance into the State of Florida, shall obtain a permit from the Department for transporting such waste into the State.

(a) An application for a permit must contain the following information or material:

1. Name, address and telephone number of the carrier; and

2. Certification statement that the carrier will comply with this part and the regulations of the U.S. Department of Transportation.

(b) Each application for a permit must be accompanied by an annual fee of $100. Permits shall be valid for 365 days following the date of issue. Permit fees are not refundable. Permits may not be transferred or assigned to another carrier.

(2)(a) Before any shipment of low-level radioactive waste may be transported into or through the State, the permitted carrier shall give written or telephonic notice to the Department not less than 48 hours prior to the date of the arrival of the shipment at the borders of the State. The carrier must provide the Department with the following information in the notice:

1. The expected date and time the shipment will arrive at the borders of the State;

2. The estimated time the shipment will remain in the State;

3. An estimate of the radioisotopes contained within the shipment;

4. An estimate of the total activity, in curies, contained within the shipment;

5. An estimate of the total volume, in cubic feet, contained within the shipment; and

6. The proposed route over which the shipment will be transported.

(b) The carrier must immediately notify the Department of any cancellations or changes of information provided in the prior notification, such as changes in the date of shipment arrival, the length of time the shipment will remain in the State, or the description or quantity of the radioactive waste contained within the shipment.

(3) Any permit issued pursuant to subsection 64E-5.1509(1), F.A.C., may be suspended if the Department has reasonable cause to suspect that the continued shipment of low-level radioactive waste presents a hazard to the public health. Grounds for suspension of a permit may include failure to include the information requested pursuant to subsection 64E-5.1509(2), F.A.C., falsification of information submitted on the application for a permit, or violation of Florida law or Department regulations. Prior to the suspension of a permit, the holder of the permit shall be notified in writing that the permit will be suspended and that an opportunity for an administrative hearing will be provided, if requested in writing within 30 days of the receipt of the notice of the intent to suspend the permit. The Department may remove the suspension at any time if the Department determines that the suspected hazard no longer exists.

(4) All applications for permits and prior notifications of impending shipments shall be addressed to the Department as outlined in subsection 64E-5.1513(2), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.131, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.131(2), (3), 404.20(6), (7), (8) FS. History–New 7-17-85, Formerly 10D-91.2010.

- 64E-5.1510 Air Transport of Plutonium.

Notwithstanding the provisions of any general license and notwithstanding any exemptions stated directly in this part or included indirectly by citation of the U.S. Department of Transportation regulations, as may be applicable, the licensee shall assure that plutonium in any form is not transported by air or offered to a carrier for air transport unless:

(1) The plutonium is contained in a medical device designed for individual human application; or

(2) The plutonium is contained in a material in which the specific activity is not greater than 0.002 microcuries (74 Bq) per gram of material and in which the radioactivity is essentially uniformly distributed; or

(3) The plutonium is shipped in a single package containing no more than an A02 quantity of plutonium in any isotope or form and is shipped in accordance with Rule 64E-5.1502, F.A.C.; or

(4) The plutonium is shipped in a package specifically authorized for the shipment of plutonium by air in the Certificate of Compliance for that package issued by the U.S. Nuclear Regulatory Commission.

Rulemaking Authority 404.051, 404.061, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.061(2), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2011.

- 64E-5.1511 Notification in the Event of Suspected or Real Breach of Containment.

In addition to the reporting requirements of the U.S. Department of Transportation, any carrier transporting radioactive material in the State shall notify the Department immediately in the event the carrier suspects or knows of a breach in the containment of the radioactive material being transported. Notification shall be made as described in subsection 64E-5.1513(2), F.A.C.

Rulemaking Authority 404.051, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), (11), 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2012.

- 64E-5.1512 Inspections.

(1) A department representative is authorized to inspect any record of persons engaged in the transportation of a radioactive material where such records reasonably relate to packaging, preparing for shipment and transporting radioactive material.

(2) A department representative is authorized to enter upon and inspect the premises and transport vehicles of any person engaged in the transportation of radioactive material for the purpose of determining compliance with or violation of the provisions of Section 404.20, F.S., and these regulations.

(3) The Department may investigate the cause and circumstances of every event in which notification was made pursuant to Rule 64E-5.1511, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.022, 404.051(1), (4), (12), 404.061(2), 404.071(1), 404.20(1), (2), (7), (8) FS. History–New 7-17-85, Formerly 10D-91.2013.

- 64E-5.1513 Communications.

(1) All communications concerning this part should be addressed to: Department of Health, Bureau of Radiation Control, 4052 Bald Cypress Way, Bin C21, Tallahassee, Florida 32399-1741.

(2) All notifications required to be made pursuant to Rules 64E-5.1506, 64E-5.1508, 64E-5.1509 and 64E-5.1511, F.A.C., shall be addressed to: Department of Health, Office of Radiation Control, Post Office Box 680069, Orlando, Florida 32868-0069; telephone (407) 297-2095.

(3) Immediate notification as required by Rule 64E-5.1511, F.A.C., shall be made by telephone or telegraph.

Rulemaking Authority 404.042, 404.051, 404.20 FS. Law Implemented 404.042, 404.051(1), (4), (6), (11), 404.061(2), 404.081(1), 404.141, 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.2014.

- 64E-5.1601 Definitions.

The following definitions apply only in this part.

(1) “AAPM” means the American Association of Physicists in Medicine: www.aapm.org.

(2) “Authorized user” means a person who has met the requirements of subsection 64E-5.1603(1), F.A.C.

(3) “Authorized medical physicist” means a person who has met the requirements of subsection 64E-5.1603(2), F.A.C.

(4) “Authorized operator” means a person who has met the requirements of subsection 64E-5.1603(3), F.A.C.

(5) “Electronic brachytherapy” means a method of radiation therapy using electrically-generated x-rays to deliver a radiation dose at a distance of up to a few centimeters by intracavitary, intraluminal or interstitial application, or by applications with the source in contact with the body surface or very close to the body surface.

(6) “Electronic brachytherapy device” or “device” means the system used to produce and deliver therapeutic radiation including the x-ray tube, the control mechanism, the cooling system, and the power source.

(7) “Electronic brachytherapy source” or “source” means the x-ray tube component used in an electronic brachytherapy device.

(8) “Medical event” means any event, except for an event that results from patient intervention, in which the administration of radiation results in:

(a) A total dose delivered that differs from the prescribed dose by 20 percent or more;

(b) A fractionated dose delivered that differs from the prescribed dose, for a single fraction, by 50 percent or more; or

(c) A dose to the wrong individual or the wrong treatment site.

(9) “Mobile electronic brachytherapy device” means a device which is transported from one address to be used at another address.

(10) “Portable shielding” means shielding that can be easily moved into the primary or secondary beam in order to reduce the radiation exposure to the patient, occupational worker or a member of the public.

Rulemaking Authority 404.051(4), 404.22 FS. Law Implemented 404.031, 404.051, 404.22 FS. History–New 3-12-09.

- 64E-5.1602 Administrative Requirements.

(1) Registration and Notification.

(a) No electronic brachytherapy device may be used on a human without a current certificate of registration from the department.

(b) An electronic brachytherapy device that is not operational and that is under the control of a registered vendor prior to final installation is exempt from the registration and fee requirements of this section.

(c) A separate registration and radiation protection program are required for facilities for which one or more of the following applies:

1. The facilities are not at the same physical address;

2. The facilities are not under the same radiation safety program; or

3. The facilities are not under the same management.

(d) Each person who acquires an electronic brachytherapy device shall apply for registration of the radiation device with the Department within 30 days after acquisition. Application for registration shall be on DH Form 1107, 10/15, “Radiation Machine Facility Registration,” as incorporated in subparagraph 64E-5.511(2)(a)1., F.A.C. The application must include the following documents:

1. A list identifying the radiation safety officer and all authorized medical physicists, authorized operators, and authorized users except visiting authorized users, together with documentation of their training and education as described in Rule 64E-5.1603, F.A.C.;

2. A copy of the most current record of surveys, calculations and quality assurance checks on each device;

3. A current copy of the quality management program as described in subsection 64E-5.1604(3), F.A.C.;

4. A current copy of the quality assurance program as described in subsection 64E-5.1604(4), F.A.C.; and

5. A copy of the device manufacturer’s U.S. Food and Drug Administration certification; and

6. Facility design information, which at a minimum must include:

a. A diagram of the physical facility showing the location of the electronic brachytherapy treatment rooms;

b. Whether the facility is a new structure or a modification to an existing structure; and

c. The type and thickness of the portable shielding used for compliance and a procedure demonstrating the use of the shielding prior to treatment.

(e) The registrant shall update the registration on file with the department within 30 days of any change to any information reported in paragraph 64E-5.1602(1)(d), F.A.C.

(2) Installation, Maintenance or Repair.

(a) Only a manufacturer’s representative registered as a vendor under subsection 64E-5.511(3), F.A.C., shall install an electronic brachytherapy device.

(b) Only a manufacturer’s representative registered as a vendor under subsection 64E-5.511(3), F.A.C., or an authorized medical physicist shall adjust, repair, maintain, or service an electronic brachytherapy device in accordance with the manufacturer’s guidelines.

(c) A registrant shall retain a record of the installation, maintenance, adjustment, service and repair of an electronic brachytherapy device for 5 years.

(3) Fees. The registrant of an electronic brachytherapy device shall comply with the requirements of paragraph 64E-5.511(2)(b), F.A.C., and pay the fees for a medical accelerator unit.

Rulemaking Authority 404.051(4), 404.22 FS. Law Implemented 404.051, 404.081(1), 404.22 FS. History–New 3-12-09, Amended 6-3-15, 3-21-16.

- 64E-5.1603 Training and Education.

(1) Qualification of Authorized User.

(a) The registrant shall require the authorized user to be a physician who:

1. Is licensed by the department as a medical doctor or doctor of osteopathy;

2. Has completed a manufacturer’s device-specific training as specified in subsection 64E-5.1603(5), F.AC.; and

3. Is certified in:

a. Radiation oncology or therapeutic radiology by the American Board of Radiology;

b. Radiation oncology by the American Osteopathic Board of Radiology;

c. Radiology, with specialization in radiotherapy, as a British “Fellow of the Faculty of Radiology” or “Fellow of the Royal College of Radiology”; or

d. Therapeutic radiology by the Canadian Royal College of Physicians and Surgeons.

(b) A physician shall not act as an authorized user for any electronic brachytherapy device until such time as said physician’s training has been reviewed and approved by the department.

(2) Qualification of Authorized Medical Physicist.

(a) The registrant shall require the authorized medical physicist to be a person who:

1. Is currently licensed pursuant to Section 483.901, F.S., as a therapeutic radiological physicist; and

2. Has completed a manufacturer’s device-specific training as specified in subsection 64E-5.1603(5), F.A.C.

(b) A medical physicist shall not act as an authorized medical physicist for any electronic brachytherapy device until such time as said physicist’s training has been reviewed and approved by the department.

(3) Qualification of Authorized Operator. A person, other than an authorized user, who operates an electronic brachytherapy device to apply ionizing radiation to a human, shall be:

(a) Certified in accordance with the Chapter 468, Part IV, F.S., as a radiation therapy technologist; and

(b) Have completed a manufacturer’s device-specific training as specified in subsection 64E-5.1603(5), F.A.C.

(4) Qualification of Radiation Safety Officer. The registrant shall require the radiation safety officer to be a person who has completed a manufacturer’s device specific training as specified in subsection 64E-5.1603(5), F.A.C., and be:

(a) An authorized user or authorized medical physicist; or

(b) A person certified by:

1. The American Board of Radiology in Radiology, Diagnostic Radiology, Therapeutic Radiology, or Radiation Oncology;

2. The American Board of Health Physics in Comprehensive Health Physics;

3. The American Board of Radiology in Radiological Physics, Therapeutic Radiological Physics, or Medical Nuclear Physics;

4. The American Board of Nuclear Medicine;

5. The American Board of Science in Nuclear Medicine; or

6. The American Board of Medical Physicists; or

(c) A person who has completed classroom and laboratory training consisting of the following:

1. One hundred hours of radiation physics and instrumentation;

2. Thirty hours of radiation protection;

3. Twenty hours of mathematics pertaining to the use and measurement of radiation;

4. Twenty hours of radiation biology;

5. Thirty hours of medical therapy training; and

6. One year of full time experience in radiation safety at a medical institution under the supervision of the individual identified as the radiation safety officer.

(5) Manufacturer’s Training. The registrant shall require training in electronic brachytherapy device operation, safety procedures, and US Food & Drug Administration approved clinical uses. All training taken to satisfy this requirement must have been completed within the 7 years preceding the date of application. This training requirement must be approved by the department and must be satisfied by:

(a) Completion of a training program provided by the manufacturer; or

(b) Completion of a training program which is provided by an institution approved by the manufacturer; or

(c) Receiving training that is substantially equivalent to the manufacturer’s training program from an authorized user or authorized medical physicist who is authorized to use the device on a department registration.

(6) Annual Training.

(a) The registrant shall provide radiation safety training, initially and at least annually, to all personnel providing patient care and treatment planning to patients.

(b) The training should include device operation, safety procedures and clinical use updates.

(7) Training Records. The registrant shall retain for three years a record of each individual receiving initial manufacturer’s training and annual training.

Rulemaking Authority 404.051(4), 404.22 FS. Law Implemented 404.051, 404.081(1), 404.22 FS. History–New 3-12-09.

- 64E-5.1604 General Technical Requirements For Electronic Brachytherapy Facilities.

(1) Radiation Surveys.

(a) The registrant shall ensure that a survey, as defined in subsection 64E-5.101(151), F.A.C., of all new facilities and existing facilities not previously surveyed, is performed with an operable radiation measurement survey instrument according to the requirements of Part III of Chapter 64E-5, F.A.C.

(b) The survey shall be performed by, or under the direction of, an authorized medical physicist or radiation safety officer who shall determine and record whether radiation levels are in compliance with the dose limits of Part III of Chapter 64E-5, F.A.C. Portable shielding may be used to comply with these radiation dose limits. Such surveys shall be conducted with the electronic brachytherapy device controls, source position, portable shielding and site-specific scattering phantom all set so as to produce the highest radiation exposure level that could occur during treatment.

(c) The survey record shall include: the date of the measurements; the reason the survey is required; the manufacturer’s name, model number and serial number of the electronic brachytherapy device; the instrument(s) used to measure radiation levels; a plan of the areas surrounding the treatment room that were surveyed; the measured dose rate at several points in each area expressed in microsieverts or millirems per hour; the calculated maximum level of radiation over a period of 1 week for each restricted and unrestricted area; and the signature of the individual responsible for conducting the survey.

(d) A survey shall also be performed prior to any subsequent medical use, when:

1. Making any change in the portable shielding;

2. Making any change in the location where the electronic brachytherapy device is used within the treatment room; or

3. Relocating the electronic brachytherapy device.

(e) The registrant shall maintain the record of each survey for the duration of the registration.

(2) Dosimetry Equipment.

(a) For electronic brachytherapy devices, the calibration of the dosimetry system shall be for the source and energy or energies in use according to a current published protocol from a nationally recognized professional association with expertise in electronic brachytherapy, such as the AAPM. In the absence of such a published protocol, the manufacturer’s current protocol shall be followed.

(b) The registrant shall have available for use a dosimetry system for quality assurance check measurements. The quality assurance check system may be the same system used to meet the requirement for calibration.

(c) The registrant shall maintain a record of each dosimetry system calibration, intercomparison, and comparison for the duration of the registration. Each record shall include: the date; the model numbers and serial numbers of the instruments that were calibrated, inter-compared or compared; the names of the individuals who performed the calibration, intercomparison, or comparison, and; evidence that the intercomparison was performed by, or under the direct supervision and in the physical presence of, an authorized medical physicist of record.

(3) Quality Management Program.

(a) Each registrant under this part shall establish and maintain a written quality management program to provide a high confidence that electronic brachytherapy devices will be used as directed by the authorized user. The quality management program must include written policies and procedures to meet the following objectives:

1. Except where a delay to provide a written directive as defined in subsection 64E-5.101(173), F.A.C., would jeopardize the patient’s health as specified in subparagraphs 64E-5.1604(3)(a)2. and 3., F.A.C., a written directive is prepared prior to administration of a therapeutic radiation dose;

2. An oral directive is only acceptable when a delay to provide a written directive would jeopardize the patient’s health because of the emergent nature of the patient’s condition. The information contained in the oral directive must be documented immediately in the patient’s record and a written directive prepared within 24 hours of the oral directive;

3. An oral revision to an existing written directive is only acceptable when a delay to provide a written revision to an existing written directive would jeopardize the patient’s health. The oral revision must be documented immediately in the patient’s record and a revised written directive must be signed by the authorized user within 48 hours of the oral revision;

4. A written directive that changes an existing written directive for any therapeutic radiation procedure is only acceptable if the revision is dated and signed by an authorized user prior to the administration of the therapeutic electronic brachytherapy dose, or the next electronic brachytherapy fractional dose;

5. The patient’s identity is verified by more than one method as the individual named in the written directive prior to administration;

6. The final plans of treatment and related calculations agree with the respective written directives;

7. Each administration agrees with the written directive; and

8. Any unintended deviation from the written directive is identified and evaluated and appropriate action is taken.

(b) The registrant shall retain for 3 years each written directive in an auditable form.

(c) The registrant shall develop procedures for and conduct a review of the quality management program including an evaluation of the following:

1. A representative sample of patient administrations within the review period, as described in a procedure submitted to the Department;

2. All recordable events, as defined in subsection 64E-5.101(123), F.A.C., within the review period; and

3. All medical events within the review period to verify compliance with all aspects of the quality management program.

(d) The review of the quality management program shall be conducted at intervals not to exceed 12 months. The registrant shall maintain a record of each dated review for inspection by the department in an auditable form for 3 years and shall include evaluations and findings of the review.

(e) The registrant shall evaluate each of these reviews to determine the effectiveness of the quality management program and make modifications to meet the objectives of the program.

(f) The registrant may make modifications to the quality management program to increase the program’s efficiency as long as the program’s effectiveness is not diminished. The registrant is required to submit any modifications to the quality management program to the department within 30 days after the modifications have been made.

(g) Within 30 days of discovery of each recordable event, as defined in subsection 64E-5.101(123), F.A.C., the registrant shall:

1. Assemble the relevant facts including the cause;

2. Identify and implement any corrective action required to prevent recurrence; and

3. Retain a record in an auditable form for 3 years of the relevant facts and any corrective action taken.

(h) Each registrant shall maintain records and reports of medical events until the termination of the registration.

(4) Quality Assurance Program.

(a) Each registrant shall develop and administer a written quality assurance program as a method of minimizing deviations from facility procedures and to document preventative measures taken prior to serious patient injury or medical event. The quality assurance program must include written procedures for performing:

1. Treatment planning, chart and treatment field parameters;

2. Patient simulation, verification of catheter placement and device exchange;

3. Dose calculation and review; and

4. Review of daily treatment records.

(b) Deviations from the prescribed treatment or from the facility’s quality assurance and operating procedures shall be investigated and brought to the attention of the authorized user, authorized medical physicist and radiation safety officer.

(c) A review of the quality assurance program shall be conducted at intervals not to exceed 3 months and shall include all the deviations from the prescribed treatment. A signed record of each dated review shall be maintained for inspection by the department in an auditable form for 3 years and shall include evaluations and findings of the review.

(5) Authority and Responsibilities.

(a) Radiation Safety Officer.

1. A registrant shall appoint a radiation safety officer responsible for implementing the radiation safety program. The registrant, through the radiation safety officer, shall ensure that radiation safety activities are performed in accordance with approved procedures and regulatory requirements in the daily operation of the electronic brachytherapy devices.

2. The radiation safety officer shall promptly investigate and implement corrective actions as necessary regarding:

a. Incidents as defined in Rule 64E-5.344, F.A.C.;

b. Reportable events as defined in Rule 64E-5.345, F.A.C.; and

c. Other deviations from approved radiation safety practice. A written report of these investigations and the corrective actions taken shall be given to management within 30 days of the incident, event or deviation.

3. The radiation safety officer shall implement written policies and procedures to:

a. Use electronic brachytherapy devices safely;

b. Perform radiation surveys whenever necessary;

c. Perform checks of survey instruments and other safety equipment;

d. Train personnel who work in or frequent areas where radiation is present; and

e. Keep a copy of all records and reports required by department regulations, a copy of these regulations, and a copy of each registration correspondence to the department, and the written policies and procedures required by the regulations.

4. The radiation safety officer shall review at least every 3 months the occupational radiation exposure records of all personnel working with radiation therapy devices.

(b) Authorized User. Authorized users shall:

1. Be physically present during the initiation of each patient treatment;

2. Be physically present during the continuation of each patient treatment or identify in writing a physician under the supervision of the authorized user who is trained in the operation of and emergency response for the device who will be physically present during the continuation of each patient treatment;

3. Personally review the patient’s case to assure that the therapeutic radiation procedure is appropriate; and

4. Review the progress of the patient receiving therapy and modify the originally prescribed dose, if needed.

(c) Visiting Authorized User.

1. A registrant may permit any visiting authorized user to use an electronic brachytherapy device for medical use under the terms of the registrant’s registration and radiation protection program for 60 days each year if:

a. The visiting authorized user has the prior written permission of the registrant’s management;

b. The registrant has a copy of an electronic brachytherapy device registration issued by the department or another state that identifies the visiting authorized user by name as an authorized user for medical use of an electronic brachytherapy device; and

c. The visiting authorized user performs only those procedures for which he is specifically authorized by the registration described in sub-subparagraph 64E-5.1604(5)(c)1.b., F.A.C.

2. A registrant shall retain copies of the records specified in subparagraph 64E-5.1604(5)(c)1., F.A.C., for 5 years after the last visit of the visiting authorized user.

(d) Authorized Medical Physicist. The authorized medical physicist shall:

1. Be physically present during the initiation and continuation of each patient treatment.

2. Evaluate the output from the electronic brachytherapy source;

3. Generate the necessary dosimetry information;

4. Review treatment calculations prior to initial treatment of any treatment site;

5. Establish the quality assurance spot checks and review the data from those checks as required by the submitted procedures;

6. Consult with the authorized user in treatment planning, as needed; and

7. Perform calculations and assessments regarding patient treatments that may constitute medical events.

(6) Operating Procedures. The registrant shall ensure compliance with the following procedures.

(a) An electronic brachytherapy device shall only be used as approved by the U.S. Food and Drug Administration for human use.

(b) When not in operation, the electronic brachytherapy device shall be secured from unauthorized use.

(c) When a patient must be held in position for radiation therapy, mechanical supporting or restraining devices shall be used.

(d) A copy of the current operating and emergency procedures shall be kept in close proximity to the electronic brachytherapy device and easily accessible to the operator.

(e) No individual other than the patient shall be exposed during the treatment.

(f) The radiation safety officer or his/her designee, and an authorized user, shall be notified as soon as possible but no later than 24 hours after a patient’s, or human research subject’s, medical emergency or death;

(g) Only individuals approved by the authorized user, radiation safety officer, or authorized medical physicist shall be present in the treatment room during treatment and a written log shall be kept of all personnel present during treatment;

(h) Simultaneous operation of more than one radiation-producing device in a treatment room shall be prohibited; and

(i) The registrant shall develop, implement, and maintain written procedures for responding to any situation in which the operator is unable to complete the treatment in compliance with the written directive. These procedures must include:

1. Instructions for responding to equipment failures and the names of the individuals responsible for implementing corrective actions;

2. The process for restricting access to, and posting of, the treatment area to minimize the risk of inadvertent exposure; and,

3. The names and telephone numbers of the authorized users, the authorized medical physicist, and the radiation safety officer to be contacted if the device operates abnormally.

(7) Possession of a Survey Instrument. Each facility location authorized to use an electronic brachytherapy device shall possess portable monitoring equipment. At a minimum, such equipment shall include a portable radiation measurement survey instrument capable of measuring dose rates over the range 0.1 microsievert (0.01 millirem) per hour to 10 millisievert (1000 millirem) per hour. All survey instruments shall be operable and calibrated annually.

(8) Calibration.

(a) Validation of the electronic brachytherapy source output shall be performed by an authorized medical physicist.

(b) Calibration validation measurements shall be made for each x-ray tube, or after any repair affecting the x-ray beam generation, or when indicated by the spot checks.

(c) Calibration validation must include determination of:

1. The output within 2% of the expected value, or determination of the output if there is no expected value;

2. Timer accuracy and linearity over the typical range of use;

3. Proper operation of back-up exposure control devices;

4. Evaluation that the relative dose distribution about the source is within 5% of that expected; and

5. Source positioning accuracy to within 1 millimeter within the applicator;

(d) The validation of the output shall use a dosimetry system as described by the facility’s procedures to measure the output. Such procedures shall use a current published protocol from a nationally recognized professional association with expertise in electronic brachytherapy, such as the AAPM. In the absence of such a published protocol, the manufacturer’s current protocol shall be followed.

(e) The registrant shall make calibration measurements required by this section according to a current published protocol from a nationally-recognized professional association with expertise in electronic brachytherapy, such as the AAPM. In the absence of such a published protocol, the manufacturer’s current testing protocol shall be followed.

(9) Routine and Day-Of-Use Periodic Spot Checks for Electronic Brachytherapy Devices and Dosimetry Equipment.

(a) A registrant authorized to use electronic brachytherapy devices shall have a program to perform spot checks on each unit:

1. At the beginning of each day of use of an electronic brachytherapy unit;

2. Each time the unit is moved to a new room or site; and

3. After each x-ray tube installation.

(b) The authorized medical physicist shall:

1. Establish written procedures for performing the spot checks;

2. Supervise the making of the spot checks and review the spot check results within 2 days of completion; and

3. Notify the registrant in writing of any failures detected during the spot checks, within 24 hours of the identification of the spot check failure.

(c) The authorized user will prevent the clinical use of a malfunctioning device until the malfunction identified in the spot check has been evaluated and corrected or, if necessary, the equipment repaired.

(d) The spot checks must, at a minimum, assure proper operation of:

1. Radiation exposure indicator lights on the electronic brachytherapy device and on the control console; and

2. The integrity of all cables, catheters or parts of the device.

(e) Spot checks of dosimetry must include checks that the output of the electronic brachytherapy source falls within 3% of expected values, which include:

1. Output as a function of time, or output as a function of setting on a monitor chamber; and

2. Verification of the consistency of the dose distribution to within 3% of that found during calibration;

3. Validation of the operation of positioning methods to assure that the treatment dose exposes the intended location within 1 mm; and

4. Inspection of all treatment components (e.g., connecting guide tubes, transfer tubes, transfer-tube-applicator interfaces, treatment spacers) on the day of use for any imperfections.

(f) A registrant shall retain a record of each spot check for 3 years. The record shall include:

1. The date of the check;

2. The manufacturer’s name, model number, and serial number of the electronic brachytherapy source;

3. Notations indicating the operability of electronic brachytherapy source exposure indicator lights, applicators, source-transfer tubes, transfer tube-applicator interfaces, and source-positioning accuracy; and

4. The name and signature of the individual who performed the check.

(10) Mobile Electronic Brachytherapy Devices. A registrant providing mobile electronic brachytherapy services shall:

(a) Check all survey instruments before medical use at each address of use and on each day of use;

(b) Account for the x-ray tube in the device before departure from the client’s address; and

(c) Perform, at each location, all of the required periodic spot checks specified in subsection 64E-5.1604(9), F.A.C., to assure proper operation of the device.

(11) Treatment Planning.

(a) The authorized medical physicist shall perform acceptance testing on the treatment planning system of therapy related computer systems according to a current published protocol from a nationally recognized professional association with expertise in electronic brachytherapy, such as the AAPM. In the absence of such a published protocol, the manufacturer’s current protocol shall be followed. At a minimum, the acceptance testing shall include verification of:

1. The electronic brachytherapy source specific input parameters required by the dose-calculation algorithm;

2. The accuracy of dose, dwell-time, and treatment-time calculations at representative points;

3. The accuracy of isodose plots and graphic displays;

4. The accuracy of the software used to determine source positions from images; and

5. If the treatment-planning system is different from the treatment delivery system, the accuracy of electronic transfer of the treatment-delivery parameters to the treatment delivery unit from the treatment planning system.

(b) The authorized medical physicist shall compare the position indicators in the applicator to the actual position of the source or planned dwell positions at the time of commissioning.

(c) Prior to each patient treatment regimen, the authorized medical physicist shall confirm the accuracy of the treatment parameters and dose.

Rulemaking Authority 404.051(4), 404.22 FS. Law Implemented 404.051, 404.081(1), 404.22 FS. History–New 3-12-09.

- 64E-5.201 Licensing of Radioactive Material.

(1) This part provides for the licensing of radioactive material. No person shall receive, possess, use, transfer, own or acquire radioactive material except as authorized in a specific or general license issued pursuant to this part or as otherwise provided in this part. Unless otherwise specified in the license or these rules, no licensee shall use radioactive materials:

(a) In or on human beings;

(b) In field applications where radioactive material is released to the environment;

(c) In products distributed to the public;

(d) In animals, plants, or their products which will be used for human consumption; or

(e) In plants or animals where their products are released to the environment.

(2) In addition to the requirements of this part, all licensees are subject to the requirements of Parts I, III, IX and XV. Licensees engaged in industrial radiographic operations are also subject to the requirements of Part IV, licensees using radionuclides in the healing arts are subject to the requirements of Part VI and licensees engaged in wireline and subsurface tracer studies are subject to the requirements of Part XI.

(3) The Procedures for Radioactive Materials Enforcement Actions, May 2000, which is available from the department and which is herein incorporated by reference, will be used to determine enforcement actions to be taken.

(4) Any license may be revoked, suspended or modified, in whole or in part, for any material false statement in the application or any statement of fact required under provisions of the law, or because of conditions revealed by such application or statement of fact on any report, record or inspection or other means which would warrant the Department to refuse to grant a license on an original application, or for violation of, or failure to observe any of the terms and conditions of the law or of the license, a rule, or an order of the Department.

Rulemaking Authority 404.051(4), 404.061(2), 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), 404.061(2), 404.081(1), 404.091, 404.141, 404.161, 404.162, 404.20(1) FS. History–New 7-17-85, Amended 8-25-91, 5-12-93, 5-15-96, Formerly 10D-91.301, Amended 10-8-00.

- 64E-5.202 Source Material ‒ Exemptions.

(1) Any person is exempt from this part to the extent that such person receives, possesses, uses, owns or transfers source material in any chemical mixture, compound, solution or alloy in which the source material is by weight less than 1/20 of 1 percent of the mixture, compound, solution or alloy.

(2) Any person is exempt from this part to the extent that such person receives, possesses, uses or transfers unrefined and unprocessed or containing source material; provided that, except as authorized in a specific license, such person shall not refine or process such ore.

(3) Any person is exempt from this part to the extent that such person receives, possesses, uses or transfers:

(a) Any quantities of thorium contained in:

1. Incandescent gas mantles;

2. Vacuum tubes;

3. Welding rods;

4. Electric lamps for illuminating purposes, provided that each lamp does not contain more than 50 milligrams of thorium;

5. Germicidal lamps, sunlamps and lamps for outdoor or industrial lighting provided that each lamp does not contain more than 2 grams of thorium;

6. Rare earth metals and compounds, mixtures and products containing not more than 0.25 percent by weight thorium, uranium, or any combination of these; or

7. Personnel neutron dosimeters, provided that each dosimeter does not contain more than 50 milligrams of thorium;

(b) Source material contained in the following products:

1. Glazed ceramic tableware, provided that the glaze contains not more than 20 percent by weight source material;

2. Glassware containing not more than 10 percent by weight source material, but not including commercially manufactured glass brick, ceramic tile or other glass, or ceramic used in construction;

3. Glass enamel or glass enamel frit containing not more than 10 percent by weight source material imported or ordered for importation into the United States, or initially distributed by manufacturers in the United States, before July 25, 1983; or

4. Piezoelectric ceramic containing not more than 2 percent by weight source material;

(c) Photographic film, negatives, and prints containing uranium or thorium;

(d) Any finished product or part fabricated of, or containing, tungsten-thorium or magnesium-thorium alloys, provided that the thorium content of the alloy does not exceed 4 percent by weight and that this exemption shall not be deemed to authorize the chemical, physical, or metallurgical treatment or processing of any such product or part;

(e) Uranium contained in counterweights installed in aircraft, rockets, projectiles, and missiles, or stored or handled in connection with installation or removal of such counterweights, provided that:

1. The counterweights are manufactured in accordance with a specific license issued by the U.S. Nuclear Regulatory Commission, authorizing distribution by the licensee pursuant to 10 C.F.R. Part 40;

2. Each counterweight has been impressed with the following legend clearly legible through any plating or other covering: “Depleated Uranium” or “Caution – Radioactive Material – Uranium” if manufactured prior to December 31, 1969;

3. Each counterweight is durably and legibly labeled or marked with the identification of the manufacturer and the statement: “Unauthorized Alterations Prohabited”, unless manufactured prior to December 31, 1969, and impressed with the legend “Caution – Radioactive Material – Uranium”.

4. This exemption shall not be deemed to authorize the chemical, physical, or metallurgical treatment or processing of any such counterweights other than repair or restoration of any plating or other covering;

(f) Natural or depleted uranium metal used as shielding constituting part of any shipping container provided that the shipping container is conspicuously and legibly impressed with the legend “Caution – Radioactive shielding – Uranium”; and the uranium metal is encased in mild steel or equally fire resistant metal of minimum wall thickness of one-eighth inch (3.2 mm).

(g) Thorium contained in finished optical lenses, provided that each lens does not contain more than 30 percent by weight of thorium, and that this exemption shall not be deemed to authorize either:

1. The shaping, grinding, or polishing of such lens or manufacturing processes other than the assembly of such lens into optical systems and devices without any alterations of the lens; or

2. The receipt, possession, use or transfer of thorium contained in contact lenses, or in spectacles, or in eyepieces in binoculars or other optical instruments;

(h) Uranium contained in detector heads for use in fire detection units, provided that each detector head contains not more than 0.005 microcurie of uranium; or

(i) Thorium contained in any finished aircraft engine part containing nickel-thoria alloy, provided that:

1. The thorium is dispersed in the nickel-thoria alloy in the form of finely divided thoria or thorium dioxide; and

2. The thorium content in the nickel-thoria alloy does not exceed 4 percent by weight.

(4) The exemptions in this section do not authorize the manufacture of any of the products described.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (4), 404.141 FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.302.

- 64E-5.203 Radioactive Material Other Than Source Material - Exemptions.

(1) Exempt Concentrations.

(a)1. Except as provided in this section, any person is exempt from this part to the extent that such person receives, possesses, uses, transfers, owns or acquires products containing radioactive material introduced in concentrations not in excess of those listed in Schedule A.

2. A manufacturer, processor, or producer of a product or material is exempt from the requirements for a license to the extent that this person transfers radioactive material contained in a product or material in concentrations not in excess of those specified in Schedule A and introduced into the product or material by a licensee holding a specific license issued by the NRC expressly authorizing such introduction. This exemption does not apply to the transfer of radioactive material contained in any food, beverage, cosmetic, drug, or other commodity or product designed for ingestion or inhalation by, or application to, a human being.

(b) No person may introduce radioactive material into a product or material knowing or having reason to believe that it will be transferred to persons exempt under paragraph (1)(a), above, or equivalent regulations of the NRC, an Agreement State or Licensing State, except in accordance with a specific license issued by the NRC pursuant to 10 CFR section 32.11.

(2) Exempt Quantities.

(a) Except as provided in paragraphs (2)(b) through (e), below, any person is exempt from these regulations to the extent that such person receives, possesses, uses, transfers, owns or acquires radioactive material in individual quantities each of which does not exceed the applicable quantity set forth in Schedule B.

(b) This paragraph does not authorize the production, packaging or repackaging of radioactive material for purposes of commercial distribution, or the incorporation of radioactive material into products intended for commercial distribution.

(c) No person may, for purposes of commercial distribution, transfer radioactive material in the individual quantities set forth in Schedule B, knowing or having reason to believe that such quantities of radioactive material will be transferred to persons exempt under this section or equivalent regulations of the U.S. Nuclear Regulatory Commission, an Agreement State or Licensing State, except in accordance with a specific license issued by the U.S. Nuclear Regulatory Commission, pursuant to Section 32.18 of 10 C.F.R. Part 32, or by the Department, pursuant to subsection 64E-5.210(2), F.A.C., which license states that the radioactive material may be transferred by the licensee to persons exempt under this subsection or the equivalent regulations of the U.S. Nuclear Regulatory Commission, an Agreement State or Licensing State. Authority to transfer possession or control by the manufacturer, processor, or producer of any equipment, device, commodity, or other product containing byproduct material whose subsequent possession, use, transfer, and disposal by all other persons are exempted from regulatory requirements may be obtained only from the U.S. Nuclear Regulatory Commission, Washington, D.C. 20555.

(d) No person may, for purposes of producing an increased radiation level, combine quantities of byproduct material covered by this exemption so that the aggregate quantity exceeds the limits set forth in Schedule B, except for radioactive material combined within a device placed in use before May 3, 1999, or as otherwise permitted by the regulations in this part.

(e)1. Any person, who possesses radioactive material received or acquired before September 25, 1971, under the then existing general license issued to transfer, receive, acquire, own, possess, use and import quantities of radioactive materials listed in subparagraph 64E-5.203(2)(e)2., F.A.C., Table of General Licensed Quantities prior to September 25, 1971 below, or similar general license of a State, or provided that no person shall at any one time possess or use, pursuant to the general license provisions of this section, more than a total of ten such quantities.

2. Below is the Table of General Licensed Quantities prior to September 25, 1971:

|  | Radioactive material | Column No. I Not as a sealed source (microcuries) | Column No. III As a sealed source (microcuries) |
| --- | --- | --- | --- |
| a. | Antimony (Sb 124) | 1 | 10 |
| b. | Arsenic 76 (As 76) | 10 | 10 |
| c. | Arsenic 77 (As 77) | 10 | 10 |
| d. | Barium 140 – Lanthanum 140 (Ba La 140) | 1 | 10 |
| e. | Beryllium 7 (Be 7) | 50 | 50 |
| f. | Cadmium 109 – Silver 109 (Cd Ag 109) | 10 | 10 |
| g. | Calcium 45 (Ca 45) | 10 | 10 |
| h. | Carbon 14 (C 14) | 50 | 50 |
| i. | Cerium 144 – Praseodymium (Ce Pr 144) | 1 | 10 |
| j. | Cesium – Barium 137 (Cs Ba 137) | 1 | 10 |
| k. | Chlorine 36 (Cl 36) | 1 | 10 |
| l. | Chromium 51 (Cr 51) | 50 | 50 |
| m. | Cobalt 60 (Co 60) | 1 | 10 |
| n. | Copper 64 (Cu 64) | 50 | 50 |
| o. | Europium 154 (Eu 154) | 1 | 10 |
| p. | Fluorine 18 (F 18) | 50 | 50 |
| q. | Gallium 72 (Ga 72) | 10 | 10 |
| r. | Germanium 71 (Ge 71) | 50 | 50 |
| s. | Gold 198 (Au 198) | 10 | 10 |
| t. | Gold 199 (Au 199) | 10 | 10 |
| u. | Hydrogen 3 (Tritium) (H 3) | 250 | 250 |
| v. | Indium 114 (In 114) | 1 | 10 |
| w. | Iodine 131 (I-131) | 10 | 10 |
| x. | Iridium 192 (Ir 192) | 10 | 10 |
| y. | Iron 55 (Fe 55) | 50 | 50 |
| z. | Iron 59 (Fe 59) | 1 | 10 |
| aa. | Lanthanum 140 (La 140) | 10 | 10 |
| bb. | Manganese 52 (Mn 52) | 1 | 10 |
| cc. | Manganese 56 (Mn 56) | 50 | 50 |
| dd. | Molybdenum 99 (Mo 99) | 10 | 10 |
| ee. | Nickel 59 (Ni 59) | 1 | 10 |
| ff. | Nickel 63 (Ni 63) | 1 | 10 |
| gg. | Niobium 95 (Nb 95) | 10 | 10 |
| hh. | Palladium 109 (Pd 109) | 10 | 10 |
| ii. | Palladium 103 – Rhodium 103 (Pd-Rh 103) | 50 | 50 |
| jj. | Phosphorus 32 (P 32) | 10 | 10 |
| kk. | Polonium 210 (Po 210) | 0.1 | 1 |
| ll. | Potassium 42 (K 42) | 10 | 10 |
| mm. | Praseodymium 143 (Pr 143) | 10 | 10 |
| nn. | Promethium 147 (Pm 147) | 10 | 10 |
| oo. | Rhenium 186 (Re 186) | 10 | 10 |
| pp. | Rhodium 105 (Rh 105) | 10 | 10 |
| qq. | Rubidium 86 (Rb 86) | 10 | 10 |
| rr. | Ruthenium 106 – Rhodium 106 (Ru Rh 106) | 1 | 10 |
| ss. | Samarium 153 (Sm 153) | 10 | 10 |
| tt. | Scandium 46 (Sc 46) | 1 | 10 |
| uu. | Silver 105 (Ag 105) | 1 | 10 |
| vv. | Silver 111 (Ag 111) | 10 | 10 |
| ww. | Sodium 22 (Na 22) | 10 | 10 |
| xx. | Sodium 24 (Na 24) | 10 | 10 |
| yy. | Strontium 89 (Sr 89) | 1 | 10 |
| zz. | Strontium 89 – Yttrium 90 (Sr Y 90) | 0.1 | 1 |
| aaa. | Sulfur 35 (S 35) | 50 | 50 |
| bbb. | Tantalum 182 (Ta 182) | 10 | 10 |
| ccc. | Technetium 96 (Tc 96) | 1 | 10 |
| ddd. | Technetium 99 (Tc 99) | 1 | 10 |
| eee. | Tellurium 127 (Te 127) | 10 | 10 |
| fff. | Tellurium 129 (Te 129) | 1 | 10 |
| ggg. | Thallium 204 (Tl 204) | 50 | 50 |
| hhh. | Tin 112 (Sn 113) | 10 | 10 |
| iii. | Tungsten 185 (W 185) | 10 | 10 |
| jjj. | Vanadium 48 (V 48) | 1 | 10 |
| kkk. | Yttrium 90 (Y 90) | 1 | 10 |
| lll. | Yttrium 91 (Y 91) | 1 | 10 |
| mmm. | Zinc 65 (Zn 65) | 10 | 10 |
| nnn. | Beta or Gamma emitting radioactive material not listed above | 1 | 10 |

(3) Exempt Items.

(a) Certain Items Containing Radioactive Material. Except for persons who apply radioactive material to, or persons who incorporate radioactive material into, the following products, any person is exempt from these regulations to the extent that he receives, possesses, uses, transfers, owns or acquires the following products. Authority to transfer possession or control by the manufacturer, processor, or producer of any equipment, device, commodity, or other product containing byproduct material whose subsequent possession, use, transfer and disposal by all other persons are exempted from regulatory requirements may be obtained only from the NRC, Washington, D.C. 20555:

1. Timepieces, hands or dials containing not more than the following specified quantities of radioactive material and not exceeding the following specified amount of radioactive material or dose rate, as applicable:

a. Twenty-five millicuries (925 MBq) of tritium per timepiece;

b. Five millicuries (185 MBq) of tritium per hand;

c. Fifteen millicuries (555 MBq) of tritium per dial; bezels when used shall be considered as part of the dial;

d. One hundred microcuries (3.7 MBq) of promethium 147 per watch or two hundred microcuries (7.4 MBq) of promethium 147 per any other timepiece;

e. Twenty microcuries (0.74 MBq) of promethium 147 per watch hand or 40 microcuries (1.48 MBq) of promethium 147 per other timepiece hand;

f. Sixty microcuries (2.22 MBq) of promethium 147 per watch dial or 120 microcuries (4.44 MBq) of promethium 147 per other timepiece dial; bezels, when used, shall be considered as part of the dial; and

g. The radiation dose rate from hands and dials containing promethium 147 or radium 226 will not exceed, when measured through 50 milligrams per square centimeter of absorber:

(I) For wrist watches, 0.1 millirad (1 uGy) per hour at 10 centimeters from any surface;

(II) For pocket watches, 0.1 millirad (1 uGy) per hour at 1 centimeter from any surface. Radium shall not be used for pocket watches; and

(III) For any other timepiece, 0.2 millirad (2 uGy) per hour at 10 centimeters from any surface;

h. One microcurie (37 kBq) of radium 226 per timepiece in intact timepieces manufactured prior to November 30, 2007.

2. Ionization chamber smoke detectors containing not more than 1 microcurie (µCi) of americium-241 per detector in the form of a foil and designed to protect life and property from fires.

3. Precision balances containing not more than 1 millicurie (37 MBq) of tritium per balance or not more than 0.5 millicurie (18.5 MBq) of tritium per balance part manufactured before December 17, 2007.

4. Marine compasses containing not more than 750 millicuries (27.8 GBq) of tritium gas and other marine navigational instruments containing not more than 250 millicuries (9.25 GBq) of tritium gas manufactured before December 17, 2007.

5. Electron tubes, including spark gap tubes, power tubes, gas tubes including glow lamps, receiving tubes, microwave tubes, indicator tubes, pick-up tubes, radiation detection tubes and any other completely sealed tube that is designed to conduct or control electrical currents; provided, that the radiation dose rate from each electron tube containing radioactive material shall not exceed 1 millirad (10 uGy) per hour at 1 centimeter from any surface when measured through 7 milligrams per square centimeter of absorber, and that each tube does not contain more than one of the following specified quantities of radioactive material:

a. One hundred fifty millicuries (5.55 GBq) of tritium per microwave receiver protector tube or 10 millicuries (370 MBq) of tritium per any other electron tube.

b. One microcurie (37 kBq) of cobalt 60.

c. Five microcuries (185 kBq) of nickel 63.

d. Thirty microcuries (1.11 MBq) of krypton 85.

e. Five microcuries (185 kBq) of cesium 137.

f. Thirty microcuries (1.11 MBq) of promethium 147.

6. Ionizing radiation measuring instruments containing, for purposes of internal calibration or standardization, one or more sources of radioactive material, provided that:

a. Each source contains no more than one exempt quantity set forth in Schedule B, and

b. Each instrument contains no more than 10 exempt quantities. For purposes of this requirement, an instrument’s source may contain either one or different types of radionuclides and an individual exempt quantity may be composed of fractional parts of one or more of the exempt quantities in Schedule B, provided that the sum of such fractions shall not exceed unity.

c. For americium 241, 0.05 microcurie (1.85 kBq) is considered an exempt quantity under this subparagraph.

(b) Self-Luminous Products Containing Radioactive Material.

1. Tritium, Krypton 85 or Promethium 147. Except for persons who manufacture, process or produce self-luminous products containing tritium, krypton 85 or promethium 147, any person is exempt from these regulations to the extent that such person receives, possesses, uses, transfers, owns or acquires tritium, krypton 85 or promethium 147 in self-luminous products manufactured, or processed, produced, imported or transferred in accordance with a specific license issued by the U.S. Nuclear Regulatory Commission pursuant to Section 32.22 of 10 C.F.R. Part 32, which license authorizes the transfer of the product to persons who are exempt from regulatory requirements. The exemption in this paragraph does not apply to tritium, krypton 85 or promethium 147 used in products for frivolous purposes or in toys or adornments.

2. Radium 226. Any person is exempt from these regulations to the extent that such person receives, possesses, uses, transfers or owns articles containing less than 0.1 microcurie (3.7 kBq) of radium 226 which were acquired prior to December 1980.

(c) Gas and Aerosol Detectors Containing Radioactive Material.

1. Except for persons who manufacture, process, produce, or initially transfer for sale or distribution gas and aerosol detectors containing radioactive material, any person is exempt from these regulations to the extent that such person receives, possesses, uses, transfers, owns or acquires radioactive material in gas and aerosol detectors designed to protect life or property from fires and airborne hazards, provided that detectors containing radioactive material shall have been manufactured, processed, produced, imported or initially transferred in accordance with a specific license issued by the NRC pursuant to Section 32.26 of 10 C.F.R. Part 32. Authority to transfer possession or control by the manufacturer, processor or producer of any equipment, device, commodity or other product containing byproduct material whose subsequent possession, use, transfer and disposal by all other persons are exempted from regulatory requirements may be obtained only from the NRC, Washington, D.C., 20555.

2. This exemption also covers gas and aerosol detectors manufactured or distributed before November 30, 2007 in accordance with a specific license issued by a State under comparable 10 CFR, section 32.26 authorizing distribution to persons exempt from regulatory requirements.

(4) Radioactive drug: capsules containing carbon 14 urea for in vivo diagnostic use for humans.

(a) Except as provided in paragraphs (b) and (c) of this section, any person is exempt from the requirements for a license set forth in these regulations if such person receives, possesses, uses, transfers, owns, or acquires capsules containing 1 microcurie (37 kBq) carbon 14 urea each, allowing for nominal variation that can occur during the manufacturing process, for in vivo diagnostic use for humans.

(b) Any person who desires to use the capsules for research involving human subjects shall apply for and receive a specific license as specified in these regulations.

(c) Any person who desires to manufacture, prepare, process, produce, package, repackage, or transfer for commercial distribution such capsules shall apply for and receive a specific license as specified in 10 CFR Part 32, Sec. 32.21.

(d) Nothing in this section relieves a person from complying with applicable FDA, other Federal, and State requirements governing receipt, administration, and use of drugs.

Rulemaking Authority 404.051, 404.061, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (10), 404.141 FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.303, Amended 10-8-00, Amended 12-26-13.

- 64E-5.204 Types of Licenses.

Licenses for radioactive materials are of two types: general and specific.

(1) Some general licenses provided in this part may be effective without the filing of applications with the department or the issuance of licensing documents to the particular persons, although the filing of a certificate with the department for general licenses pursuant to subsection 64E-5.206(7) or (8), F.A.C., shall be required of the particular general licensee prior to the receipt of radioactive material and the department requires registration of certain general licenses described in subsection 64E-5.206(4), F.A.C. The payment of a fee is also required by all persons possessing general licensed material described in paragraph (1)(c), below. The general licensee is subject to all other applicable portions of these regulations and any limitations of the general license.

(a) The annual registration fee set forth in this section for general licenses shall be payable every July 1, for as long as the license remains in effect.

(b) The annual fee for a general license set forth in Rule 64E-5.216, F.A.C., under reciprocal agreement shall be paid before the first entrance into the state and on each anniversary date thereafter, if applicable. Manufacturers, manufacturer’s representatives, distributors, and waste treatment, storage or disposal companies servicing Florida radioactive materials license applicants or licensees are not exempt from this fee.

(c) Payment of the indicated annual fee pursuant to paragraph (1)(a), above, is required for the following types of devices held or activities performed under a general license:

| 1. Static elimination devices as described in paragraph 64E-5.206(1)(a), F.A.C. | $30.00 per unit |
| --- | --- |
| 2. Measuring, gauging, and control devices as described in subsection 64E-5.206(4), F.A.C. | $30.00 per unit |
| 3. In Vivo testing as described in subsection 64E-5.206(7), F.A.C. | $150.00 per license |
| 4. In Vitro testing as described in subsection 64E-5.206(8), F.A.C. | $150.00 per license |
| 5. Depleted uranium as described in subsection 64E-5.205(4), F.A.C. | $150.00 per license |

(d) Those persons who hold a specific license from the U.S. Nuclear Regulatory Commission, an Agreement State or Licensing State and conduct activities under a reciprocal agreement with this State shall meet the requirements of subsection 64E-5.216(1), F.A.C., and pay the annual fee as specified in paragraph (2)(e), below.

(2) Specific licenses require the submission of an application to the Department and the issuance of a licensing document by the Department. The licensee is subject to all applicable portions of these regulations as well as any limitations specified in the licensing document. The licensee is subject to the payment of fees as authorized under Section 404.131, F.S., and as outlined below:

(a) The requirements of this part apply to a person who is an applicant for, or holder of a specific radioactive materials license issued pursuant to Subpart C, and for a special review of safety designs of sealed sources and devices, whether or not in conjunction with a license application on file or which may be filed.

(b) All communications concerning the requirements of this part should be addressed to or delivered in person to the Department of Health, Bureau of Radiation Control, 4052 Bald Cypress Way, Bin C21, Tallahassee, Florida 32399-1741.

(c) No additional fees shall be required for amendments to licenses.

(d) Payment of fees.

1. Application fees. Each application for a specific license for which a fee is prescribed shall be accompanied by a remittance in the full amount of the fee. No application will be processed prior to payment of the fee specified herein. The application fee is not refundable except in those cases where the Department has determined that a license is not required. The Department will consider any application abandoned if the Department does not receive a reply within 90 days to its most recent request for additional information. In such cases, the applicant must submit a new application with the application fee specified herein.

2. Annual fees. All current specific licenses that were in effect on January 1, 1979, are subject to payment of the annual fee prescribed herein and on every January 1, thereafter, as long as the license remains in effect. All specific licenses issued after January 1, 1979, are subject to payment of the annual fee specified in this section within 60 days of issuance of the license and on each anniversary date thereafter. The annual fee is not refundable except in those cases where the Department has determined that the fee is not required.

3. Method of payment. Checks, drafts or money orders for payment of fees shall be payable to DOH, Office of Radiation Control; and sent to: Department of Health, Bureau of Radiation Control, 4052 Bald Cypress Way, Bin C21, Tallahassee, Florida 32399-1741.

(e) Below is the schedule of fees for specific radioactive materials licenses:

|  | Application | Annual |
| --- | --- | --- |
|  | Fee | Fee |
| 1. Source Material: |  |  |
| a. Licenses for concentration of uranium from phosphate ores for the production of uranium as “yellow cake” or powdered solid; | $8,288 | $14,330 |
| b. Licenses for concentration of uranium from phosphate ores for the production of “green cake” or equivalent, moist or solid; | $4,522 | $8,927 |
| c. All other specific source material licenses excluding depleted uranium used as shielding and counterweights. | $653 | $275 |
| 2. Special Nuclear Material (SNM): |  |  |
| a. Licenses for use of SNM in sealed sources contained in devices used in measuring systems; | $784 | $622 |
| b. Licenses for use of SNM not sufficient to form a critical mass, except as in paragraph 2.a., above, and paragraphs 2.c. and 5.e., below; | $1,608 | $2,333 |
| c. Licenses for use of SNM to be used as calibration and reference sources. | $246 | $131 |
| 3. Byproduct, naturally occurring or accelerator produced material: |  |  |
| a. Licenses for processing or manufacturing for commercial distribution or industrial uses; | $3,508 | $3,362 |
| b. Licenses for processing or manufacturing and distribution of radiopharmaceuticals. This category includes radiopharmacies; | $3,072 | $4,608 |
| c. Licenses for industrial radiography performed only in an approved shielded radiography installation; | $1,870 | $2,593 |
| d. Licenses for industrial radiography performed only at the address indicated in the license, or at temporary job sites of the licensee; | $1,972 | $3,188 |
| e. Licenses for possession and use of radioactive materials in sealed sources for irradiation of materials where the source is not removed from the shield and is less than 10,000 curies; | $726 | $726 |
| f.(I) Licenses for possession and use of radioactive materials in sealed sources for irradiation of materials when the source is not removed from the shield and is greater than 10,000 curies and less than 100,000 curies, or where the source is less than 100,000 curies and is removed from the shield; | $1,697 | $1,956 |
| (II) Licenses for possession and use of radioactive materials in sealed sources for irradiation of materials when the source is equal to or greater than 100,000 curies and less than 1,000,000 curies; | $4,391 | $4,753 |
| (III) Licenses for possession and use of radioactive materials in sealed sources for irradiation of materials when the source is equal to or greater than 1,000,000 curies; | $11,736 | $5,278 |
| g. Licenses issued to distribute items containing radioactive materials to persons under a general license; | $1,972 | $2,580 |
| h. Fixed gauging devices; | $726 | $1,159 |
| i. Well logging; |  |  |
| (I) Sealed sources or sub-surface tracer studies; | $1,362 | $1,798 |
| (II) Sub-surface tracer studies and sealed sources; | $1,723 | $1,913 |
| j. Nuclear laundry; | $3,840 | $6,781 |
| k. Industrial or medical research and development; | $1,421 | $1,769 |
| l.(I) Portable gauging devices; | $726 | $1,159 |
| (II) In Vitro and clinical laboratory; | $870 | $1,102 |
| (III) Academic; | $1,174 | $1,405 |
| (IV) Possession of uranium or thorium, or their decay products, as a result of mining or processing; | $1,174 | $1,044 |
| (V) All other specific licenses except as otherwise noted; | $870 | $1,202 |
| m. Licenses of broad scope; |  |  |
| (I) Academic; | $3,840 | $8,815 |
| (II) Medical; | $3,840 | $6,569 |
| (III) Industrial or Research and Development; | $3,840 | $5,482 |
| n. Gas chromatography devices; | $521 | $377 |
| o. Reference or calibration sources equal to or less than 1 millicurie total; | $377 | $158 |
| p. Nuclear service licenses, such as, leak testing, instrument calibration, etc. | $622 | $492 |
| 4. Waste disposal or processing: |  |  |
| a. Commercial waste disposal or treatment facilities, including burial or incineration; | $331,010 | $300,666 |
| b. All other commercial facilities involving compaction, repackaging, storage or transfer; | $32,501 | $29,965 |
| c. Commercial treatment of radioactive materials for release to unrestricted areas. | $6,913 | $6,882 |
| 5. Medical use: |  |  |
| a.(I) Teletherapy or gamma stereotactic radiosurgery including gamma knife devices; | $1,838 | $1,791 |
| (II) High, medium, low or pulsed dose rate remote afterloader devices; | $1,697 | $1,654 |
| (III) High, medium, low or pulsed dose rate remote afterloader devices and gamma stereotactic radiosurgery including gamma knife devices or teletherapy devices; | $1,838 | $1,791 |
| b. Medical institutions, including hospitals, except category 5.a.(I), 5.a.(II), 5.a.(III), 5.e. and 5.f.; | $1,972 | $2,290 |
| c. Private practice physicians except category 5.a.(I), 5.a.(II), 5.a.(III), 5.d. and 5.f.; | $1,421 | $1,608 |
| d. Private practice physicians using only strontium 90 eye applicators, or materials authorized by Rule 64E-5.630, F.A.C., or materials authorized by Rule 64E-5.631, F.A.C.; | $726 | $898 |
| e. Nuclear powered pacemakers; | $521 | $319 |
| f.(I) Mobile nuclear medicine services. | $1,697 | $1,950 |
| (II) Mobile high, medium, low or pulsed dose rate remote afterloader device when the treatment is only performed on the mobile vehicle. | $2,970 | $3,308 |
| 6. Civil defense. | $653 | $985 |
| 7. Device, product, or sealed source safety evaluation: |  |  |
| a. Safety evaluation of devices or products containing radioactive material, except reactor fuel devices, for commercial distribution or in accordance with the unique specifications of, and for use by, a single applicant; per device remaining in active status. Devices or products in inactive status more than 5 years must submit another application fee and be re-evaluated; | $4,500 | $2,570 |
| b. Safety evaluation of sealed sources containing radioactive material, for commercial distribution or in accordance with the unique specifications of, and for use by, a single applicant; per source remaining in active status. Sources in inactive status more than 5 years must submit another application fee and be re-evaluated. | $2,400 | $2,900 |

Rulemaking Authority 404.051, 404.061, 404.131 FS. Law Implemented 404.031, 404.051, 404.061, 404.131, 404.141 FS. History–New 7-17-85, Amended 9-9-90, 8-25-91, 5-12-93, 11-6-94, Formerly 10D-91.304, Amended 5-18-98, 9-28-06, 8-16-07, 12-26-13.

- 64E-5.205 General Licenses – Source Material.

(1) A general license is hereby issued authorizing commercial and industrial firms, research, educational and medical institutions, and State and local governmental agencies to use and transfer not more than 15 pounds (6.82 kg) of source material at any given time for research, development, educational, commercial or operational purposes. A person authorized to use or transfer source material, pursuant to this general license, may not receive more than a total of 150 pounds (68.2 kg) of source material in any calendar year.

(2) Persons who receive, possess, use or transfer source material pursuant to the general license issued in subsection (1), above, are exempt from the provisions of Parts III and IX to the extent that such receipt, possession, use or transfer is within the terms of such general license; provided, however, this exemption shall not be deemed to apply to any such person who is also in possession of source material under a specific license issued pursuant to this part.

(3) A general license is hereby issued authorizing the receipt of title to source material without regard to quantity. This general license does not authorize any person to receive, possess, use or transfer source material.

(4) Depleted Uranium in Industrial Products and Devices.

(a) A general license is hereby issued to receive, acquire, possess, use or transfer, in accordance with the provisions of paragraphs (4)(b), (c), (d) and (e), below, depleted uranium contained in industrial products or devices for the purpose of providing a concentrated mass in a small volume of the product or device.

(b) The general license in paragraph (4)(a), above, applies only to industrial products or devices which have been manufactured either in accordance with a specific license issued to the manufacturer of the products or devices pursuant to Rule 64E-5.210, F.A.C., or in accordance with a specific license issued to the manufacturer by the U.S. Nuclear Regulatory Commission or an Agreement State which authorizes manufacture of the products or devices for distribution to general licensees of the U.S. Nuclear Regulatory Commission or an Agreement State.

(c)1. Persons who receive, acquire, possess or use depleted uranium pursuant to the general license established by paragraph (4)(a), above, shall submit the information requested on DOH Form 1619, entitled “General License for Depleted Uranium”, which is herein incorporated by reference effective 7-17-85, with the Department. The form shall be submitted within 30 days after the first receipt or acquisition of such depleted uranium. The registrant shall submit a fee as required in paragraph 64E-5.204(1)(c), F.A.C.

2. The registrant possessing or using depleted uranium under the general license established by paragraph (4)(a), above, shall report in writing to the Department any changes in information furnished by him in the “Registration Certificate – Use of Depleted Uranium Under General License” form. The report shall be submitted within 30 days after the effective date of such change.

(d) A person who receives, acquires, possesses or uses depleted uranium pursuant to the general license established by paragraph (4)(a), above:

1. Shall not introduce such depleted uranium, in any form, into a chemical, physical, or metallurgical treatment or process, except a treatment or process for repair or restoration of any plating or other covering of the depleted uranium;

2. Shall not abandon such depleted uranium;

3. Shall transfer or dispose of such depleted uranium only by transfer in accordance with the provisions of Rule 64E-5.215, F.A.C. In the case where the transferee receives the depleted uranium pursuant to the general license established by (4)(a), above, the transferor shall furnish the transferee a copy of this regulation and a copy of the “Registration Certificate – Use of Depleted Uranium Under General License”. In the case where the transferee receives the depleted uranium pursuant to a general license contained in the U.S. Nuclear Regulatory Commission’s or Agreement State’s regulation equivalent to paragraph (4)(a), above, the transferor shall furnish the transferee a copy of this regulation and a copy of the “Registration Certificate – Use of Depleted Uranium Under General License” accompanied by a note explaining that use of the product or device is regulated by the U.S. Nuclear Regulatory Commission or Agreement State under requirements substantially the same as in this regulation;

4. Within 30 days of any transfer, shall report in writing to the Department the name and address of the person receiving the depleted uranium pursuant to such transfer; and

5. Shall not export such depleted uranium except in accordance with a license issued by the U.S. Nuclear Regulatory Commission pursuant to 10 C.F.R. Part 110.

(e) Any person receiving, acquiring, possessing, using or transferring depleted uranium pursuant to the general license established by paragraph (4)(a), above, is exempt from the requirements of Parts III and IX with respect to the depleted uranium covered by that general license.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (8), (9), (10), 404.061(2), 404.081(1), 404.141 FS. History–New 7-17-85, Formerly 10D-91.305.

- 64E-5.206 General Licenses – Radioactive Material Other Than Source Material.

(1) Certain Devices and Equipment. A general license is hereby issued to transfer, receive, acquire, own, possess and use radioactive material incorporated in the following devices or equipment which have been manufactured, tested and labeled by the manufacturer in accordance with a specific license issued to the manufacturer by the U.S. Nuclear Regulatory Commission for use pursuant to Section 31.3 of 10 C.F.R. Part 31. This general license is subject to the provisions of Rules 64E-5.103 through 64E-5.104, paragraph 64E-5.203(1)(b), Rules 64E-5.213, 64E-5.215, Part III, Part IX and Part XV, F.A.C.

(a) Static Elimination Devices. Devices designed for use as static eliminators which contain, as a sealed source or sources, radioactive material consisting of a total of not more than 500 microcuries (18.5 MBq) of polonium 210 per device; and

(b) Ion Generating Tubes. Devices designed for ionization of air which contain, as a sealed source or sources, radioactive material consisting of a total of not more than 500 microcuries (18.5 MBq) of polonium 210 per device or a total of not more than 50 millicuries (1.85 GBq) of tritium per device.

(2) Reserved.

(3) Reserved.

(4) Certain Measuring, Gauging and Controlling Devices.

(a) A general license is hereby issued to commercial and industrial firms and to research, educational and medical institutions, individuals in the conduct of their businesses, and state or local government agencies to own, receive, acquire, possess, use or transfer in accordance with the provisions of paragraphs (4)(b), (c) and (d), below, radioactive material, excluding special nuclear material, contained in devices designed and manufactured for the purpose of detecting, measuring, gauging or controlling thickness, density, level, interface location, radiation, leakage, or qualitative or quantitative chemical composition, or for producing light or an ionized atmosphere.

(b)1. The general license in paragraph (4)(a), above, applies only to radioactive material contained in devices which have been manufactured or initially transferred and labeled in accordance with the specifications contained in a specific license issued by the Department pursuant to subsection 64E-5.210(4), F.A.C., or in accordance with the specifications contained in a specific license issued by the NRC, or an Agreement State, which authorizes distribution of devices to persons granted a general license by the NRC, or an Agreement State. Regulations under the Federal Food, Drug, and Cosmetic Act authorizing the use of radioactive control devices in food production require certain additional labeling thereon which is found in 21 C.F.R. Part 179, section 179.21, April 1, 2013 edition, and is herein incorporated by reference and may be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-05471 or at http://www.gpo.gov/fdsys/pkg/CFR-2013-title21-vol3/pdf/CFR-2013-title21-vol3-part179-subpartB.pdf.

2. The devices must have been received from one of the specific licenses described in subparagraph (b)1., above or through a transfer made under subparagraph 64E-5.206(4)(c)8., F.A.C.

(c) Any person who owns, receives, acquires, possesses, uses, or transfers radioactive material in a device pursuant to the general license in paragraph (4)(a), above;

1. Shall assure that all labels affixed to the device at the time of receipt, and bearing a statement that removal of the label is prohibited, are maintained thereon and shall comply with all instructions and precautions provided by such labels;

2. Shall assure that the device is tested for leakage of radioactive material and proper operation of the on-off mechanism and indicator, if any, at no longer than 6-month intervals or at such other intervals as are specified in the label. However,

a. Devices containing only krypton need not be tested for leakage of radioactive material; and,

b. Devices containing only tritium or not more than 100 microcuries (3.7 MBq) of other beta- and/or gamma-emitting material or 10 microcuries (0.37 MBq) of alpha-emitting material and devices held in storage in the original shipping container prior to initial installation need not be tested for any purpose;

3. Shall assure that other testing, installation, servicing and removal from installation involving the radioactive material, its shielding or containment, are performed:

a. In accordance with the instructions provided by the labels, or

b. By a person holding an applicable specific license from the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State to perform such activities;

4. Shall maintain records showing compliance with the requirements of subparagraphs (4)(c)2. and 3., above. The records shall show the results of tests. The records also shall show the dates of performance of, and the names of persons performing testing, installation, servicing and removal from installation concerning the radioactive material, its shielding or containment. Records of tests for leakage of radioactive material required by subparagraph (4)(c)2., above, shall be maintained for at least three years after the next required leak test is performed or until the transfer or disposal of the sealed source. Records of tests of the on-off mechanism and indicator required by subparagraph (4)(c)2., above, shall be maintained for at least three years after the next required test of the on-off mechanism and indicator is performed or until the sealed source is transferred or disposed. Records which are required by subparagraph (4)(c)3., above, shall be maintained for a period of at least 3 years from the date of the recorded event or until the transfer or disposal of the device.

5. Upon the occurrence of a failure of or damage to, or any indication of a possible failure of or damage to, the shielding of the radioactive material or the on-off mechanism or indicator, or upon the detection of 0.005 microcurie (185 Bq) or more removable radioactive material, shall immediately suspend operation of the device until it has been repaired by the manufacturer or other person holding an applicable specific license from the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State to repair such devices, or disposed of by transfer to a person authorized by an applicable specific license to receive the radioactive material contained in the device and, within 30 days, furnish to the Department a report containing a brief description of the event and the remedial action taken; and in the case of removable radioactive materials or failure of or damage to a source likely to result in contamination of the premises or the environment, a plan for ensuring the premises and environment are acceptable for unrestricted use using the criteria described in Rule 64E-5.222, F.A.C.

6. Shall not abandon the device containing radioactive material;

7. Except as provided in subparagraph (4)(c)8., below, shall transfer or dispose of the device containing radioactive material only by export as provided by subparagraph 15. below, transfer to a specific licensee of the Department, the NRC, or an Agreement State, whose specific license authorizes him to receive the device, and within 30 days after transfer of a device to a specific licensee or export, shall furnish to the Department a report containing identification of the device by manufacturer’s or initial transferor’s name and model number and serial number, the name, address, license number, where applicable, of the person receiving the device, and the date of the transfer;

8. Shall transfer the device by export as provided by paragraph 15 below, or to another general licensee only:

a. Where the device remains in use at a particular location. In such case the transferor shall give the transferee a copy of this section, a copy of Rules 64E-5.103, 64E-5.343, and 64E-5.344, F.A.C., and any safety documents identified in the label on the device and within 30 days of the transfer, report to the Department the manufacturer’s or initial transferor’s name and model number and serial number of device transferred, the transferor’s name and mailing address for the location of use, and the name, title and phone number of the responsible individual identified by the transferee in accordance with subparagraph 64E-5.206(4)(c)11., F.A.C., to have knowledge of and authority to take actions to ensure compliance with these regulations; or

b. Where the device is held in storage in the original shipping container at its intended location of use prior to initial use by a general licensee; and

9. Shall comply with the provisions of Rules 64E-5.343 and 64E-5.344, F.A.C., for reporting radiation incidents, theft or loss of licensed material, but shall be exempt from the other requirements of Parts III and IX.

10. Shall be required to obtain written Department authorization before transferring the device to any other specific license not specifically identified in subparagraph 64E-5.206(4)(c)7., F.A.C. A holder of a specific license may transfer a device for possession and use under its own specific license without prior approval if the specific license holder satisfies the following requirements:

a. Verifies that the specific license authorizes the possession and use, or applies for and obtains an amendment to the license authorizing the possession and use;

b. Removes, alters, covers, or clearly and unambiguously augments the existing label otherwise required by subparagraph 64E-5.206(4)(c)1., F.A.C., so that the device is labeled in compliance with Rule 64E-5.325, F.A.C., provided the manufacturer, model number, and serial number is retained;

c. Obtains manufacturer’s or initial transferor’s information concerning maintenance that would be applicable under the specific license such as leak testing procedures;

d. Reports the transfer under subparagraph 64E-5.206(4)(c)7., F.A.C.

11. Shall appoint an individual responsible for having knowledge of the appropriate regulations and requirements and the authority for taking required actions to comply with appropriate regulations and requirements. The general licensee, through this individual, shall ensure the day-to-day compliance with the appropriate regulations and requirements. This appointment does not relieve the general licensee of any of its responsibility in the regard.

12.a. Shall register, in accordance with sub-subparagraphs 64E-5.206(4)(c)12.b. and 64E-5.206(4)(c)12.c., F.A.C., all devices except exit signs containing tritium. Each address for a location of use as described in sub-sub-subparagraph 64E-5.206(4)(c)12.c.(IV), F.A.C., represents a separate general license and requires a separate registration.

b. Shall annually register with the Department the possession of a device meeting the criteria in sub-subparagraph 64E-5.206(4)(c)12.a., F.A.C. Registration must be done by verifying, correcting or adding to the information provided in a request for registration received from the Department. The registration information must be submitted to the Department within 30 days of the date of the request for registration or as otherwise indicated in the request. In addition, the general licensee holding devices that meet the criteria of sub-subparagraph 64E-5.206(4)(c)12.a., F.A.C., is subject to the bankruptcy notification requirements in subsection 64E-5.213(3), F.A.C.

c. Shall provide the following information and any other information requested by the Department:

(I) Name and mailing address of the general licensee;

(II) For each device, the manufacturer’s name or initial transferor name, model number, serial number, the radioisotope and activity as identified on the label;

(III) Name, title, and telephone number of the responsible person designated a representative of the general licensee under subparagraph 64E-5.206(4)(c)11., F.A.C.;

(IV) Address or location at which the device(s) are used or stored. For portable devices, the address of the primary place of storage;

(V) Certification by the responsible representative of the general licensee that the information concerning the devices(s) has been verified through a physical inventory and checking the label information; and

(VI) Certification by the responsible representative of the general licensee that they are aware of the requirements of the general license.

d. Persons generally licensed by other Agreement States, Licensing States, or the U.S. Nuclear Regulatory Commission with respect to devices meeting the criteria in 10 CFR 31.5(c)(13)(i) are not subject to registration requirements if the devices are used in areas subject to the Department jurisdiction for less than 180 days in any calendar year. The Department will not request registration from such licensees.

13. Shall report to the Department changes in the general licensee name and the mailing address for each location of use within 30 days of the effective date of the change. For a portable device, a report of address change is required for a change in the device’s primary place of storage.

14. Shall not hold devices that are not in use longer than 2 years. If the devices with shutters are not being used, the shutters must be locked in the closed position. The testing required by subparagraph 64E-5.206(4)(c)2., F.A.C., need not be performed during the period of storage only. However, when devices are put back into service or transferred to another person, and have not been tested within the required test interval, they must be tested for leakage before use or transfer and the shutter trested before use. Devices kept in standby for future use are excluded from the two year time limit if the general licensee performs physical inventories at intervals not to exceed three months while they are in standby.

15. Shall not export the device containing radioactive material except in accordance with 10 C.F.R. Part 110;

16. Shall respond to written requests from the Department to provide information relating to the general license within 30 calendar days of the date of the request, or other time specified in the request. If the general licensee cannot provide the requested information within the allotted time, it shall, within that same time period, request a longer period to supply the information by providing the Department, a written justification for the request for extension of time.

(d) The general license in paragraph (4)(a), above, does not authorize the manufacture of devices containing radioactive material.

(e) The general license provided in paragraph (4)(a), above, is subject to the provisions of Rules 64E-5.103 through 64E-5.104, 64E-5.213, 64E-5.215, and Part XV, F.A.C.

(5) Luminous Safety Devices for Aircraft.

(a) A general license is hereby issued to own, receive, acquire, possess and use tritium or promethium 147 contained in luminous safety devices for use in aircraft, provided:

1. Each device contains not more than 10 curies (370 GBq) of tritium or 300 millicuries (11.1 GBq) of promethium 147; and

2. Each device has been manufactured, assembled or imported in accordance with a specific license issued by the U.S. Nuclear Regulatory Commission, or each device has been manufactured or assembled in accordance with the specifications contained in a specific license issued by the Department or any Agreement State to the manufacturer or assembler of such device pursuant to licensing requirements equivalent to those in Section 32.53 of 10 C.F.R. Part 32.

(b) Persons who own, receive, acquire, possess or use luminous safety devices pursuant to the general license in paragraph (5)(a), above, are exempt from the requirements of Parts III and IX except that they shall comply with the provisions of Rules 64E-5.343 and 64E-5.344, F.A.C.

(c) This general license does not authorize the manufacture, assembly or repair of luminous safety devices containing tritium or promethium 147.

(d) This general license does not authorize the ownership, receipt, acquisition, possession or use of promethium 147 contained in instrument dials.

(e) This general license is subject to the provisions of Rules 64E-5.103 through 64E-5.104, 64E-5.213, 64E-5.215, and Part XV, F.A.C.

(6) Calibration and Reference Sources.

(a) A general license is hereby issued to those persons listed below to own, receive, acquire, possess, use and transfer, in accordance with the provisions of paragraphs (6)(d) and (e), below, americium 241 in the form of calibration or reference sources:

1. Any person who holds a specific license issued by the Department which authorizes him to receive, possess, use and transfer radioactive material; and

2. Any person who holds a specific license issued by the U.S. Nuclear Regulatory Commission which authorizes him to receive, possess, use and transfer special nuclear material.

(b) A general license is hereby issued to own, receive, possess, use and transfer plutonium in the form of calibration or reference sources in accordance with the provisions of paragraphs (6)(d) and (e), below, to any person who holds a specific license issued by the Department which authorizes him to receive, possess, use and transfer radioactive material.

(c) A general license is hereby issued to own, receive, possess, use and transfer radium 226 in the form of calibration or reference sources in accordance with the provisions of paragraphs (6)(d) and (e), below, to any person who holds a specific license issued by the Department which authorizes him to receive, possess, use and transfer radioactive material.

(d) The general licenses in paragraph (6)(a), (b) and (c), above, apply only to calibration or reference sources which have been manufactured in accordance with the specifications contained in a specific license issued to the manufacturer or importer of the sources by the U.S. Nuclear Regulatory Commission pursuant to Section 32.57 of 10 C.F.R. Part 32 or Section 70.39 of 10 C.F.R. Part 70 or which have been manufactured in accordance with the specifications contained in a specific license issued to the manufacturer by the Department, an Agreement State or Licensing State pursuant to licensing requirements equivalent to those contained in Section 32.57 of 10 C.F.R. Part 32 or Section 70.39 of 10 C.F.R. Part 70.

(e) The general licenses provided in paragraphs (6)(a), (b) and (c), above, are subject to the provisions of Rules 64E-5.103 through 64E-5.104, 64E-5.213, 64E-5.215, Parts III, IX and XV, F.A.C. In addition, persons who own, receive, acquire, possess, use or transfer one or more calibration or reference sources pursuant to these general licenses:

1. Shall not possess at any given time, at any single location of storage or use, more than 5 microcuries (185 kBq) of americium 241, 5 microcuries (185 kBq) of plutonium, or 5 microcuries (185 kBq) of radium 226 in such sources;

2. Shall not receive, possess, use or transfer such source unless the source, or the storage container, bears a label which includes one of the following statements, as appropriate, or a substantially similar statement which contains the information called for in one of the following statements, as appropriate:

a. The receipt, possession, use and transfer of this source, model ___, serial no. ___, are subject to a general license and the regulations of the U.S. Nuclear Regulatory Commission or of a state with which the Commission has entered into an agreement for the exercise of regulatory authority. Do not remove this label.

CAUTION – RADIOACTIVE MATERIAL – THIS SOURCE CONTAINS (AMERICIUM 241) (PLUTONIUM). DO NOT TOUCH RADIOACTIVE PORTION OF THIS SOURCE.

____________________________

Name of manufacturer or importer

b. The receipt, possession, use and transfer of this source, model ___, serial no. ___, are subject to a general license and the regulations of a Licensing State. Do not remove this label.

CAUTION – RADIOACTIVE MATERIAL – THIS SOURCE CONTAINS RADIUM 226. DO NOT TOUCH RADIOACTIVE PORTION OF THIS SOURCE.

____________________________

Name of manufacturer or importer

3. Shall not transfer, abandon or dispose of such source except by transfer to a person authorized by a license from the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State to receive the source;

4. Shall store such source, except when the source is being used, in a closed container adequately designed and constructed to contain americium 241, plutonium or radium 226, which might otherwise escape during storage; and

5. Shall not use such source for any purpose other than the calibration of radiation detectors or the standardization of other sources.

(f) These general licenses do not authorize the manufacture of calibration or reference sources containing americium 241, plutonium or radium 226.

(7) Medical Diagnostic Uses.

(a) A general license shall be issued to any physician to receive, possess, transfer or use radioactive material set forth below for the stated diagnostic uses, provided, however, that the use is in accordance with the provision of (7)(b), (c) and (d), below, the radioactive material is in the form of capsules, disposable syringes or other prepackaged individual doses; and the radioactive material has been manufactured in accordance with a specific license issued by the Department pursuant to subsection 64E-5.210(7), F.A.C., or by the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State pursuant to equivalent regulations authorizing distribution to persons under a general license pursuant to this subsection or its equivalent:

1. Chromium 51 as sodium radiochromate for determination of red blood cell volumes and studies of red blood cell survival time;

2. Cobalt 57 for the measurement of intestinal absorption of cyanocobalamin;

3. Cobalt 58 for the measurement of intestinal absorption of cyanocobalamin;

4. Cobalt 60 for the measurement of intestinal absorption of cyanocobalamin;

5. Iodine 125 as iodinated human serum albumin (IHSA) for determinations of blood and blood plasma volume;

6. Iodine 131 as sodium iodide for measurement of thyroid uptake; and

7. Iodine 131 as iodinated human serum albumin (IHSA) for determinations of blood and blood plasma volume.

(b) No physician shall receive, possess, use or transfer radioactive material pursuant to the general license established by paragraph (7)(a), above, until he has submitted the original and one copy of the completed form DH 361, 10/12 and received from the Department a validated copy of this form with a certification number assigned. DH 361 10/12, entitled, “Certificate – Medical Use of Radioactive Material under General License,” is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03450 or at http://www.doh.state.fl.us/environment/radiation/matform.htm.

(c) A physician who receives, possesses or uses a pharmaceutical containing radioactive material pursuant to the general license established by paragraph (7)(a), above, shall comply with the following:

1. The physician shall not possess at any given time, pursuant to the general license in paragraph (7)(a), above, more than

a. Two hundred microcuries (7.4 MBq) of iodine 131,

b. Two hundred microcuries (7.4 MBq) of iodine 125,

c. Five microcuries (185 kBq) of cobalt 57,

d. Five microcuries (185 kBq) of cobalt 58,

e. Five microcuries (185 kBq) of cobalt 60,

f. Two hundred microcuries (7.4 MBq) of chromium 51;

2. The physician shall store the pharmaceutical in the original shipping container until administered, or in a container providing equivalent radiation protection;

3. The physician shall use the pharmaceutical only for the uses authorized by paragraph (7)(a), above;

4. The physician shall not transfer the radioactive material to a person who is not authorized to receive it pursuant to a license issued by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or Licensing State, or in any manner other than in the unopened, labeled shipping container as received from the supplier, except by administering it to a patient.

(d) The general licensed physician possessing or using radioactive material under the general license of paragraph (7)(a), above, shall report in duplicate to the Department any changes in the information furnished by him on Form 361. The report shall be submitted within 30 days after the effective date of such change.

(e) Any person using radioactive material pursuant to the general license of paragraph (7)(a), above, is exempt from the requirements of Parts III and IX with respect to the radioactive material covered by the general license.

(f) Manufacturers of radiopharmaceuticals which are under the general license in this subsection are required to affix a certain identifying label to the container, and in the leaflet or brochure which accompanies the radiopharmaceutical, pursuant to subsection 64E-5.210(7), F.A.C.

(8) General License for Use of Radioactive Material for Certain In Vitro Clinical or Laboratory Testing.

(a) A general license is hereby issued to any physician, veterinarian, clinical laboratory or hospital to receive, acquire, possess, transfer or use, for any of the following stated tests, in accordance with the provisions of paragraphs (8)(b), (c), (d), (e) and (f), below, the following radioactive materials in prepackaged units for use in in vitro clinical or laboratory tests not involving internal or external administration of radioactive material, or the radiation therefrom, to human beings or animals:

1. Carbon 14, in units not exceeding 10 microcuries (370 kBq) each.

2. Cobalt 57, in units not exceeding 10 microcuries (370 kBq) each.

3. Hydrogen 3 (tritium), in units not exceeding 50 microcuries (1.85 MBq) each.

4. Iodine 125, in units not exceeding 10 microcuries (370 kBq) each.

5. Mock Iodine 125 reference or calibration sources, in units not exceeding 0.05 microcuries (1.85 kBq) of iodine 129 and 0.005 microcuries (0.185 kBq) of americium 241 each.

6. Iodine 131, in units not exceeding 10 microcuries (370 kBq) each.

7. Iron 59, in units not exceeding 20 microcuries (740 kBq) each.

8. Selenium 75, in units not exceeding 10 microcuries (370 kBq) each.

(b) No person shall receive, acquire, possess, use or transfer radioactive material pursuant to the general license established by paragraph (8)(a), above, until he has submitted the original and one copy of the completed form DH 360 10/12, and received from the Department a validated copy of this form with a certification number assigned. DH 360 10/12 entitled, “Certificate – In Vitro Testing with Radioactive Material under General License” is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03451 or at http://www.doh.state.fl.us/environment/radiation/matform.htm.

(c) A person who receives, acquires, possesses or uses radioactive material pursuant to the general license established by paragraph (8)(a), above, shall comply with the following:

1. The general licensee shall not possess at any given time, pursuant to the general license in paragraph (8)(a), above, at any single location of storage or use, a combined total amount of iodine 125, iodine 131, selenium 75, iron 59 or cobalt 57 in excess of 200 microcuries (7.4 MBq).

2. The general licensee shall store the radioactive material, until used, in the original shipping container or in a container providing equivalent radiation protection.

3. The general licensee shall use the radioactive material only for the uses authorized by paragraph (8)(a), above.

4. The general licensee shall not transfer the radioactive material to a person who is not authorized to receive it pursuant to a license issued by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or Licensing State, nor transfer the radioactive material in any manner other than in the unopened, labeled shipping container as received from the supplier.

5. The general licensee shall dispose of the mock iodine 125 reference or calibration sources described in paragraph (8)(a), above, as required by Rule 64E-5.328, F.A.C.

(d) The general licensee shall not receive, acquire, possess or use radioactive material pursuant to paragraph (8)(a), above:

1. Except as prepackaged units which are labeled in accordance with the provisions of an applicable specific license issued pursuant to subsection 64E-5.210(8), F.A.C., or in accordance with the provisions of a specific license issued by the U.S. Nuclear Regulatory Commission, an Agreement State or Licensing State which authorizes the manufacture and distribution of iodine 125, iodine 131, carbon 14, hydrogen 3 (tritium), iron 59, selenium 75, cobalt 57 or mock iodine 125 to persons under a general license described in this subsection or its equivalent, and

2. Unless one of the following statements, as appropriate or a substantially similar statement which contains the information called for in one of the following statements, appears on a label affixed to each prepackaged unit or appears in a leaflet or brochure which accompanies the package:

a. This radioactive material shall be received, acquired, possessed and used only by physicians, veterinarians, clinical laboratories or hospitals and only for in in vitro clinical or laboratory tests not involving internal or external administration of the material, or the radiation therefrom, to human beings or animals. Its receipt, acquisition, possession, use and transfer are subject to the regulations and a general license of the U.S. Nuclear Regulatory Commission or of a state with which the U.S. Nuclear Regulatory Commission has entered into an agreement for the exercise of regulatory authority.

|  | __________________ |
| --- | --- |
|  | Name of manufacturer |

b. This radioactive material shall be received, acquired, possessed and used only by physicians, veterinarians, clinical laboratories or hospitals and only for in vitro clinical or laboratory tests not involving internal or external administration of the material, or the radiation therefrom, to human beings or animals. Its receipt, acquisition, possession, use and transfer are subject to the regulations and a general license of a Licensing State.

|  | __________________ |
| --- | --- |
|  | Name of manufacturer |

(e) The physician, veterinarian, clinical laboratory or hospital possessing or using radioactive material under the general license of paragraph (8)(a), above, shall report in writing to the Department any changes in the information furnished by him in the “Certificate – In Vitro Testing with Radioactive Material Under General License”. The report shall be furnished within 30 days after the effective date of such change.

(f) Any person using radioactive material pursuant to the general license of paragraph (8)(a), above, is exempt from the requirements of Parts III and IX with respect to radioactive material covered by that general license, except that such persons using the mock iodine 125 described in subparagraph (8)(a)5., above, shall comply with the provisions of Rules 64E-5.328, 64E-5.343, and 64E-5.344, F.A.C.

(g) The New Drug provisions of the Federal Food, Drug and Cosmetic Act also govern the availability and use of any specific diagnostic drugs in interstate commerce.

(9) Ice Detection Devices.

(a) A general license is hereby issued to own, receive, acquire, possess, use and transfer strontium 90 contained in ice detection devices, provided each device contains not more than 50 microcuries (1.85 MBq) of strontium 90 and each device has been manufactured or imported in accordance with a specific license issued by the U.S. Nuclear Regulatory Commission or each device has been manufactured in accordance with the specifications contained in a specific license issued by the Department or an Agreement State to the manufacturer of such device pursuant to licensing requirements equivalent to those in Section 32.61 of 10 C.F.R. Part 32.

(b) Persons who own, receive, acquire, possess, use or transfer strontium 90 contained in ice detection devices pursuant to the general license in paragraph (9)(a), above:

1. Shall, upon occurrence of visually observable damage, such as a bend or crack or discoloration from overheating to the device, discontinue use of the device until it has been inspected, tested for leakage and repaired by a person holding a specific license from the U.S. Nuclear Regulatory Commission or an Agreement State to manufacture or service such devices; or shall dispose of the device pursuant to the provisions of Rule 64E-5.328, F.A.C.;

2. Shall assure that all labels affixed to the device at the time of receipt, and which bear a statement which prohibits removal of the labels, are maintained thereon; and

3. Are exempt from the requirements of Parts III and IX except that such persons shall comply with the provisions of Rules 64E-5.328, 64E-5.343 and 64E-5.344, F.A.C.

(c) This general license does not authorize the manufacture, assembly, disassembly or repair of strontium 90 in ice detection devices.

(d) This general license is subject to the provisions of Rules 64E-5.103 through 64E-5.104, 64E-5.213, 64E-5.215 and Part XV, F.A.C.

(10) Ownership of Radioactive Material. A general license is hereby issued to own radioactive material without regard to quantity. Notwithstanding any other provisions of this part, this general license does not authorize the manufacture, production, transfer, receipt, possession or use of radioactive material.

(11) Certain Items and Self-Luminous Products Containing Radium-226.

(a) A general license is hereby issued to any person to acquire, receive, possess, use, or transfer, in accordance with the provisions of subsections 64E-5.206(11)(b), (c), and (d), F.A.C., radium-226 contained in the following products manufactured prior to November 30, 2007.

1. Antiquities originally intended for use by the general public. For the purposes of this paragraph, antiquities mean products originally intended for use by the general public and distributed in the late 19th and early 20th centuries, such as radium emanatory jars, revigators, radium water jars, radon generators, refrigerator cards, radium bath salts, and healing pads.

2. Intact timepieces containing greater than 0.037 megabecquerel (1 microcurie), nonintact timepieces, and timepiece hands and dials no longer installed in timepieces.

3. Luminous items installed in air, marine, or land vehicles.

4. All other luminous products, provided that no more than 100 items are used or stored at the same location at any one time.

5. Small radium sources containing no more than 0.037 megabecquerel (1 microcurie) of radium-226. For the purposes of this paragraph, ‘‘small radium sources’’ means discrete survey instrument check sources, sources contained in radiation measuring instruments, sources used in educational demonstrations (such as cloud chambers and spinthariscopes), electron tubes, lightning rods, ionization sources, static eliminators, or as designated by the NRC.

(b) Persons who acquire, receive, possess, use, or transfer byproduct material under the general license issued in paragraph 64E-5.206(11)(a), F.A.C., of this section are exempt from the provisions of Parts III and IX, to the extent that the receipt, possession, use, or transfer of radioactive materials is within the terms of the general license. This exemption shall not apply to any such person specifically licensed under Chapter 64E-5, F.A.C.

(c) Any person who acquires, receives, possesses, uses, or transfers byproduct material in accordance with the general license in subsection 64E-5.206(11)(a), F.A.C., must also comply with the following requirements:

1. Shall notify the Department should there be any indication of possible damage to the product so that it appears it could result in a loss of the radioactive material. A report containing a brief description of the event, and the remedial action taken, must be furnished to the Department within 30 days;

2. Shall not abandon products containing radium-226. The product, and any radioactive material from the product, may only be disposed of according to Rule 64E-5.328, F.A.C., or by transfer to a person authorized by a specific license to receive the radium-226 in the product or as otherwise approved by the Department;

3. Shall not export products containing radium-226 except in accordance with 10 C.F. R. Part 110, 1-1-14 edition which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-05472 or at http://www.gpo.gov/fdsys/pkg/CFR-2014-title10-vol2/pdf/CFR-2014-title10-vol2-part110.pdf;

4. Shall dispose of products containing radium-226 at a disposal facility authorized to dispose of radioactive material in accordance with any Federal or State solid or hazardous waste law, including the Solid Waste Disposal Act, as authorized under the Energy Policy Act of 2005, by transfer to a person authorized to receive radium-226 by a specific license issued under Part III, or equivalent regulations of an Agreement State or the NRC, as otherwise approved by the Department.

5. Shall respond to written requests from the Department to provide information relating to the general license within 30 calendar days of the date of the request, or other time specified in the request. If the general licensee cannot provide the requested information within the allotted time, it shall, within that same time period, request a longer period to supply the information by providing the Department a written justification for the request.

(d) Except for disassembly and repair of timepieces described in subparagraph 64E-5.206(11)(a)2., F.A.C., the general license in paragraph 64E-5.206(11)(a), F.A.C., does not authorize the manufacture, assembly, disassembly, repair, or import of products containing radium-226.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.051(1), (4), (6), (8), (9), (10), (11), 404.061(2), 404.071(1), (3), 404.081(1), 404.141 FS. History–New 7-17-85, Amended 4-4-89, 1-1-94, Formerly 10D-91.306, Amended 9-28-06, 2-28-08, 12-26-13, 7-1-15.

- 64E-5.207 Filing Application for Specific Licenses.

(1) An original and one copy of an application for specific licenses, license renewals, and license amendments shall be filed with the department on Application for Radioactive Materials License Non-Human Use, DH Form 1054 12/09 or Application for Radioactive Materials Human Use, DH Form 1322 12/09, which are herein incorporated by reference.

(2) The Department may at any time after the filing of the original application, and before the expiration of the license, require further statements in order to enable the Department to determine whether the application should be granted or denied or whether a license should be modified or revoked.

(3) An existing license shall not expire until final action by the department if a licensee has filed an application for renewal in proper form not less than 30 days before expiration of his existing license or for a new license authorizing the same activities.

(4) Applications for license amendments are not required to be submitted on DOH forms but shall specify the respects in which the licensee desires the license to be amended and the grounds for such amendment.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), (11), 404.061(2), 404.141 FS. History–New 7-17-85, Amended 4-4-89, 5-12-93, 5-15-96, Formerly 10D-91.307, Amended 2-11-10.

- 64E-5.208 General Requirements for the Issuance of Specific Licenses.

A license application for a new, amended, or renewed license will be approved if the department determines that:

(1) The applicant is qualified by reason of training and experience to use the material in question for the purpose requested in accordance with these rules in such a manner as to minimize danger to public health and safety or property;

(2) The applicant’s proposed equipment, facilities and procedures are adequate to minimize danger to public health and safety or property.

Rulemaking Authority 404.051, 404.061, 404.071, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), 404.141 FS. History–New 7-17-85, Amended 5-12-93, 5-15-96, Formerly 10D-91.308.

- 64E-5.209 Special Requirements for Specific Licenses of Broad Scope.

This section prescribes requirements for the issuance of specific licenses of broad scope for radioactive material and certain regulations governing holders of such licenses. Authority to transfer possession or control by the manufacturer, processor or producer of any equipment, device, commodity or other product containing byproduct material whose subsequent possession, use, transfer and disposal by all other persons are exempted from regulatory requirements may be obtained only from the U.S. Nuclear Regulatory Commission, Washington, D.C. 20555.

(1) The different types of broad scope licenses are set forth below:

(a) A Type A specific license of broad scope is a specific license authorizing receipt, acquisition, ownership, possession, use and transfer of any chemical or physical form of the radioactive material specified in the license, but not exceeding quantities specified in the license, for any authorized purpose. The quantities specified are usually in the multicurie range.

(b) A Type B specific license of broad scope is a specific license authorizing receipt, acquisition, ownership, possession, use and transfer of any chemical or physical form of radioactive material specified in Schedule D of this part, for any authorized purpose. The possession limit for a Type B license of broad scope, if only one radionuclide is possessed thereunder, is the quantity specified for that radionuclide in Schedule D, Column I. If two or more radionuclides are possessed thereunder, the possession limit for each is determined as follows: For each radionuclide, determine the ratio of the quantity possessed to the applicable quantity specified in Schedule D, Column I, for that radionuclide. The sum of the ratios for all radionuclides possessed under the license shall not exceed unity.

(c) A Type C specific license of broad scope is a specific license authorizing receipt, acquisition, ownership, possession, use and transfer of any chemical or physical form of radioactive material specified in Schedule D of this part, for any authorized purpose. The possession limit for a Type C license of broad scope, if only one radionuclide is possessed thereunder, is the quantity specified for that radionuclide in Schedule D, Column II. If two or more radionuclides are possessed thereunder, the possession limit is determined for each as follows: For each radionuclide, determine the ratio of the quantity possessed to the applicable quantity specified in Schedule D, Column II, for that radionuclide. The sum of the ratios for all radionuclides possessed under the license shall not exceed unity.

(2) An application for a Type A specific license of broad scope will be approved if:

(a) The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.;

(b) The applicant has engaged in more than one type of activity involving the use of radioactive material; and

(c) The applicant has established administrative controls and provisions relating to organization and management, procedures, record keeping, material control and accounting, and management review that are necessary to assure safe operations, including:

1. The establishment of a radiation safety committee composed of such persons as a radiation safety officer, a representative of management, and persons trained and experienced in the safe use of radioactive material;

2. The appointment of a radiation safety officer who is qualified by training and experience in radiation protection, and who is available for advice and assistance on radiation safety matters; and

3. The establishment of appropriate administrative procedures to assure:

a. Control of procurement and use of radioactive material;

b. Completion of safety evaluations of proposed uses of radioactive material which take into consideration such matters as the adequacy of facilities and equipment, training and experience of the users, and the operating or handling procedures; and

c. Review, approval and recording by the radiation safety committee of safety evaluations of proposed uses prepared in accordance with sub-subparagraph (2)(c)3.b., above, prior to use of the radioactive material.

(3) An application for a Type B specific license of broad scope will be approved if:

(a) The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.; and

(b) The applicant has established administrative controls and provisions relating to organization and management, procedures, record keeping, material control and accounting, and management review that are necessary to assure safe operations, including:

1. The appointment of a radiation safety officer who is qualified by training and experience in radiation protection, and who is available for advice and assistance on radiation safety matters, and

2. The establishment of appropriate administrative procedures to assure,

a. Control of procurement and use of radioactive material,

b. Completion of safety evaluations of proposed uses of radioactive material which take into consideration such matters as the adequacy of facilities and equipment, training and experience of the user, and the operating or handling procedures, and

c. Review, approval and recording by the radiation safety officer of safety evaluations of proposed uses prepared in accordance with sub-subparagraph (3)(b)2.b., above, prior to use of the radioactive material.

(4) An application for a Type C specific license of broad scope will be approved if:

(a) The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.;

(b) The applicant submits a statement that radioactive material will be used only by, or under the direct supervision of, individuals who have received:

1. A college degree at the bachelor level, or equivalent training and experience, in the physical or biological sciences or in engineering, and

2. At least 40 hours of training and experience in the safe handling of radioactive material, and in the characteristics of ionizing radiation, units of radiation dose and quantities, radiation detection instrumentation and biological hazards of exposure to radiation appropriate to the type and forms of radioactive material to be used; and

(c) The applicant has established administrative controls and provisions relating to procurement of radioactive material, procedures, record keeping, material control and accounting, and management review necessary to assure safe operations.

(5) Specific licenses of broad scope are subject to the following conditions:

(a) Unless specifically authorized, persons licensed pursuant to this section shall not:

1. Conduct tracer studies in the environment involving direct release of radioactive material;

2. Receive, acquire, own, possess, use or transfer devices containing 100,000 curies (3.7 PBq) or more of radioactive material in sealed sources used for irradiation of materials;

3. Conduct activities for which a specific license issued by the Department under Rule 64E-5.210 or 64E-5.211, F.A.C., is required; or

4. Add or cause the addition of radioactive material to any food, beverage, cosmetic, drug or other product designed for ingestion or inhalation by, or application to, a human being.

(b) Each Type A specific license of broad scope issued under this part shall be subject to the condition that radioactive material possessed under the license may only be used by, or under the direct supervision of, individuals approved by the licensee’s radiation safety committee.

(c) Each Type B specific license of broad scope issued under this part shall be subject to the condition that radioactive material possessed under the license may only be used by, or under the direct supervision of, individuals approved by the licensee’s radiation safety officer.

(d) Each Type C specific license of broad scope issued under this part shall be subject to the condition that radioactive material possessed under the license may only be used by, or under the direct supervision of, individuals who satisfy the requirements of (4), above.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (8), (9), (10), (11), 404.061(2), 404.071(1), (3), 404.081(1), 404.141 FS. History–New 7-17-85, Formerly 10D-91.310.

- 64E-5.210 Special Requirements for a Specific License to Manufacture, Assemble, Repair or Distribute Commodities, Products or Devices Which Contain Radioactive Material.

(1) Licensing the Distribution of Radioactive Material in Exempt Concentrations. Authority to transfer possession or control by the manufacturer, processor, or producer of any equipment, device, commodity or other product containing by-product material whose subsequent possession, use, transfer and disposal by all other persons are exempted from regulatory requirements may be obtained only from the NRC, Washington, D.C. 20555. No person may introduce radioactive materials into a product or material knowing or having reason to believe that it will be transferred to persons exempt under subsection 64E-5.203(1), F.A.C., NRC or equivalent regulations of an Agreement State, except in accordance with a license issued by the NRC under 10 CFR section 32.11.

(2) Licensing the Distribution of Radioactive Material in Exempt Quantities. Authority to transfer possession or control by the manufacturer, processor, or producer of any equipment, device, commodity or other product containing byproduct material whose subsequent possession, use, transfer and disposal by all other persons are exempted from regulatory requirements may be obtained only from the NRC, Washington, D.C. 20555. No person may introduce radioactive materials into a product or material knowing or having reason to believe that it will be transferred to persons exempt under subsection 64E-5.203(1), F.A.C., NRC or equivalent regulations of an Agreement State, except in accordance with a license issued by the NRC under 10 CFR section 32.11.

(3) Licensing the Distribution of Radioactive Material in Exempt Items. Authority to transfer possession or control by the manufacturer, processor, or producer of any equipment, device, commodity or other product containing by-product material whose subsequent possession, use, transfer and disposal by all other persons are exempted from regulatory requirements may be obtained only from the NRC, Washington, D.C. 20555. No person may introduce radioactive materials into a product or material knowing or having reason to believe that it will be transferred to persons exempt under subsection 64E-5.203(1), F.A.C., NRC or equivalent regulations of an Agreement State, except in accordance with a license issued by the NRC under 10 CFR section 32.11.

(4) Licensing the Manufacture and Distribution of Devices to General Licensees Under subsection 64E-5.206(4), F.A.C.

(a) An application for a specific license to manufacture or distribute devices containing radioactive material, excluding special nuclear material, to persons possessing a general license under subsection 64E-5.206(4), F.A.C., or equivalent regulations of the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State will be approved if:

1. The applicant satisfies the general requirements of Rule 64E-5.208, F.A.C.;

2. The applicant submits sufficient information relating to the design, manufacture, prototype testing, quality control, labels, proposed uses, installation, servicing, leak testing, operating and safety instructions, and potential hazards of the device to provide reasonable assurance that:

a. The device can be safely operated by persons not having training in radiological protection,

b. Under ordinary conditions of handling, storage and use of the device, the radioactive material contained in the device will not be released or inadvertently removed from the device, and it is unlikely that any person will receive a dose in excess of 10 percent of the limits specified in Rule 64E-5.304, F.A.C., and

c. Under accident conditions such as fire and explosion associated with handling, storage and use of the device, it is unlikely that any person would receive an external radiation dose or dose commitment in excess of the following organ doses:

(I) Whole body; head and trunk; active blood-forming organs; gonads; or lens of eye – 15 rems (150 mSv)

(II) Hands and forearms; feet and ankles; localized areas of skin averaged over areas no larger than 1 square centimeter – 200 rems (2 Sv)

(III) Other organs – 50 rems (500 mSv); and

3. Each device bears a durable, legible, clearly visible label or labels approved by the Department which contain in a clearly identified and separate statement:

a. Instructions and precautions necessary to assure safe installation, operation and servicing of the device; documents such as operating and service manuals may be identified in the label and used to provide this information.

b. The requirement, or lack of requirement, for leak testing, or for testing any on-off mechanism and indicator, including the maximum time interval for such testing, and the identification of radioactive material by isotope, quantity of radioactivity, and date of determination of the quantity, and

c. The information called for in one of the following statements, as appropriate, in the same or substantially similar form. The model, serial number, and name of the manufacturer or distributor may be omitted from this label provided the information is elsewhere specified in labeling affixed to the device:

(I) The receipt, possession, use and transfer of this device, model ___, serial no. ___, are subject to a general license or the equivalent and the regulations of the U.S. Nuclear Regulatory Commission or a state with which the U.S. Nuclear Regulatory Commission has entered into an agreement for the exercise of regulatory authority. This label shall be maintained on the device in a legible condition. Removal of this label is prohibited.

CAUTION – RADIOACTIVE MATERIAL

_____________________________

Name of manufacturer or distributor

(II) The receipt, possession, use and transfer of this device, model ___, serial no. ___, are subject to a general license or the equivalent, and the regulations of a Licensing State. This label shall be maintained on the device in a legible condition. Removal of this label is prohibited.

CAUTION – RADIOACTIVE MATERIAL

_____________________________

Name of manufacturer or distributor

4. Each device having a separable source housing that provides the primary shielding for the source also bears, on the source housing, a durable label containing the device model number and serial number, the radioisotope and quantity, the words “Caution Radioactive Materials,” the radiation symbol described in Rule 64E-5.322, F.A.C., the name of the manufacturer or initial distributor.

5. Each device containing at least 10 millicuries (370 MBq) of cesium-137, 0.1 millicuries (3.7 MBq) of strontium-90, 1 millicurie (37 MBq) of cobalt-60, or 1 millicurie (37 MBq) of americium-241 or any other element with atomic numbers greater than 92, based on the activity indicated on the label, must bear a permanent label affixed to the source housing if separable, or the device if the source housing is not separable, that includes the words “Caution Radioactive Materials,” and if practical, the radiation symbol described in Rule 64E-5.322, F.A.C. Example of a permanent label include labels that are embossed, etched, stamped or engraved to the source housing or device as applicable.

(b) In the event the applicant desires that the device be required to be tested at intervals longer than 6 months, either for proper operation of the on-off mechanism and indicator, if any, or for leakage of radioactive material or for both, the applicant shall include in the application sufficient information to demonstrate that such longer interval is justified by performance characteristics of the device or similar devices and by design features which have a significant bearing on the probability or consequences of leakage of radioactive material from the device or failure of the on-off mechanism and indicator. In determining the acceptable interval for the test for leakage of radioactive material, the Department will consider the following information:

1. Primary containment or source capsule;

2. Protection of primary containment;

3. Method of sealing containment;

4. Containment construction material;

5. Form of contained radioactive material;

6. Maximum temperature withstood during prototype tests;

7. Maximum pressure withstood during prototype tests;

8. Maximum quantity of contained radioactive material;

9. Radiotoxicity of contained radioactive material; and

10. Operating experience with identical devices or similarly designed and constructed devices.

(c) In the event the applicant desires that the general licensee under Rule 64E-5.206, F.A.C., or under equivalent regulations of the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State be authorized to install the device, collect the sample to be analyzed by a specific licensee for leakage of radioactive material, service the device, test the on-off mechanism and indicator, or remove the device from installation, the applicant shall include in the application written instructions to be followed by the general licensee, estimated calendar quarter doses associated with such activity or activities, and bases for such estimates. The submitted information shall demonstrate that performance of such activity or activities by an individual untrained in radiological protection, in addition to other handling, storage, and use of devices under the general license, is unlikely to cause that individual to receive a dose in excess of 10 percent of the limits specified in Rule 64E-5.304, F.A.C.

(d) If a device containing radioactive material is transferred for use under the general license described in subsection 64E-5.206(4), F.A.C., each person that is licensed under subsection 64E-5.210(4), F.A.C., shall provide the information specified in this section to each person to whom a device is to be transferred. This information must be provided before the device may be transferred. In the case of a transfer through an intermediate person, the information must also be provided to the intended user prior to the initial transfer to the intermediate person. The required information includes the following:

1. A copy of the general license contained in subsection 64E-5.206(4); subparagraphs 64E-5.206(4)(c)2., 3., and 4. or subparagraph 64E-5.206(4)(c)12., F.A.C., do not apply to the particular device, those paragraphs may be omitted;

2. A copy of Rules 64E-5.103, 64E-5.343, and 64E-5.344, F.A.C.;

3. A list of services that can only be performed by a specific licensee;

4. Information on acceptable disposal options including costs of disposal; and

5. An indication that department policy is to issue high civil penalties for improper disposal.

(e) If a device containing radioactive material is transferred for use under an equivalent general license of an Agreement State or the NRC, each person that is licensed under subsection 64E-5.210(4), F.A.C., shall provide the information specified in this section to each person to whom a device is to be transferred. This information must be provided before the device may be transferred.  In the case of a transfer through an intermediate person, the information must also be provided to the intended user prior to the initial transfer to the intermediate person.  The required information includes the following:

1. A copy of the Agreement State or NRC equivalent to Rules 64E-5.103, 64E-5.343, and 64E-5.344, F.A.C. If a copy of the NRC regulations is provided to a prospective general licensee in lieu of the Agreement State’s regulations, it shall be accompanied by a note explaining that the use of the device is regulated by the Agreement State. If certain parts of the regulations do not apply to the particular device, those regulations may be omitted;

2. A list of services that can only be performed by a specific licensee;

3. Information on acceptable disposal options including costs of disposal; and

4. The name or title, address, and phone number of the contact at the Agreement State regulatory agency or U.S. Nuclear Regulatory Commission, as applicable, from which additional information may be obtained.

(f) Each device that is transferred must meet the labeling requirements in subparagraphs 64E-5.210(4)(d)3. through 5., F.A.C.

(g) If a notification of bankruptcy has been made under subsection 64E-5.213(3), F.A.C., or the license is to be terminated, each person licensed under subsection 64E-5.210(4), F.A.C., shall provide, upon request, to the department, U.S. Nuclear Regulatory Commission and to any appropriate Agreement State, records of final disposition required under paragraph 64E-5.210(4)(j), F.A.C.

(h) Each person licensed under subsection 64E-5.210(4), F.A.C., shall comply with the following reporting and record keeping requirements.

1. Report all transfers of devices to persons for use under the general license described in subsection 64E-5.206(4), F.A.C., and all receipts of devices from persons licensed under subsection 64E-5.206(4), F.A.C., to the Department. This report must be submitted at intervals not to exceed 3 months and contain all of the information described in “Transfers of Industrial Devices Report 04/2007” which is herein incorporated by reference and is available at the address listed in paragraph 64E-5.204(2)(b), F.A.C., or can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03452 or at http://www.doh.state.fl.us/environment/radiation/regs/64e-5tab.htm.

2. This report must be clear and legible and contain the following data:

a. The identity of each general licensee by name and mailing address for the location of use; if no mailing address for the location of use, an alternative address for the general licensee shall be submitted along with information on the actual location of use;

b. The name, title, and phone number of the person identified by the general licensee as having knowledge of and authority to take required actions to ensure compliance with the appropriate regulations and requirements;

c. The date of transfer;

d. The type, model number, and serial number of the device transferred; and

e. The quantity and type of radioactive materials contained in the device.

3. If one or more intermediate persons will temporarily possess the device at the intended place of use before its possession by the user, the report must include the same information for both the intended user and each intermediate person and clearly designate the intermediate person(s).

4. For devices received from a subsection 64E-5.206(4), F.A.C., general licensee, the report must include the identity of the general licensee by name and address, the type, model number, and serial number of the device received, the date of receipt, and, in the case of devices not initially transferred by the reporting licensee, the name of the manufacturer or initial transferor.

5. If the licensee makes changes to the device possessed by a subsection 64E-5.206(4), F.A.C., general licensee, such that the label must be changed to update required information, this report must identify the general licensee, the device, and the changes to information on the device label.

6. The report must clearly identify the specific licensee submitting the report and include the licenses number of the specific licensee.

7. If no transfers have been made to or from persons generally licensed under subsection 64E-5.206(4), F.A.C., during the reporting period, the report must so indicate.

(i) Each person licensed under subsection 64E-5.210(4), F.A.C., shall comply with the following additional reporting and record keeping requirements for transfers and receipt of devices to Agreement States or the NRC.

1. Report all transfers of devices to persons for use under the general license in an Agreement State or the NRC, that are equivalent to subsection 64E-5.206(4), F.A.C., and all receipts of devices from persons licensed under a general license in Agreement State or the NRC jurisdiction to the responsible Agreement State or the NRC agency. This report must contain all of the information described in “Transfers of Industrial Devices Report 04/2007.”

2. The report must be clear and legible and contain the following data:

a. The identity of each general licensee by name and mailing address for the location of use; if no mailing address for the location of use, an alternative address for the general licensee shall be submitted along with information on the actual location of use;

b. The name, title, and phone number of the person identified by the general licensee as having knowledge of and authority to take required actions to ensure compliance with the appropriate regulations and requirements;

c. The date of transfer;

d. The type, model number, and serial number of the device transferred; and

e. The quantity and type of radioactive materials contained in the device.

3. If one or more intermediate persons will temporarily possess the device at the intended place of use before its possession by the user, the report must include the same information for both the intended user and each intermediate person and clearly designate the intermediate person(s).

4. For devices received from a general licensee, the report must include the identity of the general licensee by name and address, the type, model number, and serial number of the device received, the date of receipt, and, in the case of devices not initially transferred by the reporting licensee, the name of the manufacturer or initial transferor.

5. If the licensee makes changes to the device possessed by a general licensee, such that the label must be changed to update required information, this report must identify the general licensee, the device, and the changes to information on the device label.

6. The report must clearly identify the specific licensee submitting the report and include the license number of the specific licensee.

7. If no transfers have been made to or from a particular Agreement State or the NRC during the reporting period, this information shall be reported to the responsible Agreement State or the NRC agency upon request of the agency.

8. The report must cover each calendar quarter and must be filed within 30 days of the end of the calendar quarter and must clearly indicate the period covered by the report.

(j) The persons shall maintain all information concerning transfers and receipts of devices that supports the reports required by subsection 64E-5.210(4), F.A.C. Records and reports described in subsection 64E-5.210(4), F.A.C., shall be maintained for inspection by the department for a period of 3 years following the date of the recorded event.

(5) Special Requirements for the Manufacture, Assembly or Repair of Luminous Safety Devices for Use in Aircraft. An application for a specific license to manufacture, assemble or repair luminous safety devices containing tritium or promethium 147 for use in aircraft, for distribution to general licensees under subsection 64E-5.206(5), F.A.C., will be approved if the requirements of Sections 32.53, 32.54, 32.55, 32.56 and 32.101 of 10 C.F.R., Part 32, or their equivalent and the general requirements specified in Rule 64E-5.208, F.A.C., are satisfied.

(6) Special Requirements for License to Manufacture Calibration Sources Containing Americium 241, Plutonium or Radium 226 for Distribution to Persons Generally Licensed Under subsection 64E-5.206(6), F.A.C. An application for a specific license to manufacture calibration and reference sources containing americium 241, plutonium or radium 226 to general licensees under subsection 64E-5.206(6), F.A.C., will be approved if the requirements of Sections 32.57, 32.58, 32.59 and 32.102 of 10 C.F.R., Part 32 and Section 70.39 of 10 C.F.R., Part 70, or their equivalent and the general requirements of Rule 64E-5.208, F.A.C., are satisfied.

(7) Manufacture and Distribution of Radioactive Material for Medical Use Under General License. In addition to requirements set forth in Rule 64E-5.208, F.A.C., a specific license authorizing the distribution of radioactive material for use by physicians under the general license in subsection 64E-5.206(7), F.A.C., will be issued if:

(a) The applicant submits evidence that the radioactive material is to be manufactured, labeled and packaged in accordance with a new drug application which the Commissioner of Food and Drugs, Food and Drug Administration, has approved, or in accordance with a license for a biological product issued by the Secretary, U.S. Department of Health and Human Services; and

(b) One of the following statements, as appropriate, or a substantially similar statement which contains the information called for in one of the following statements, appears on the label affixed to the container or appears in the leaflet or brochure which accompanies the package:

1. This radioactive drug may be received, possessed and used only by physicians licensed by the State of Florida to dispense drugs in the practice of medicine. Its receipt, possession, use and transfer are subject to the regulations and a general license or its equivalent of the U.S. Nuclear Regulatory Commission or of a state with which the Commission has entered into an agreement for the exercise of regulatory authority.

____________________

Name of Manufacturer

2. This radioactive drug may be received, possessed and used only by physicians licensed by the State of Florida to dispense drugs in the practice of medicine. Its receipt, possession, use and transfer are subject to the regulations and a general license or its equivalent of a Licensing State.

____________________

Name of Manufacturer

(8) Manufacture and Distribution of Radioactive Material for Certain in In Vitro Clinical or Laboratory Testing Under General License. An application for a specific license to manufacture or distribute radioactive material for use under the general license of subsection 64E-5.206(8), F.A.C., will be approved if:

(a) The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.

(b) The radioactive material is to be prepared for distribution in prepackaged units of:

1. Carbon 14 in units not exceeding 10 microcuries (370 kBq) each.

2. Cobalt 57 in units not exceeding 10 microcuries (370 kBq) each.

3. Hydrogen 3 (tritium) in units not exceeding 50 microcuries (1.85 MBq) each.

4. Iodine 125 in units not exceeding 10 microcuries (370 kBq) each.

5. Mock iodine 125 in units not exceeding 0.05 microcurie (1.85 kBq) of iodine 129 and 0.005 microcurie (185 Bq) of americium 241 each.

6. Iodine 131 in units not exceeding 10 microcuries (370 kBq) each.

7. Iron 59 in units not exceeding 20 microcuries (740 kBq) each.

8. Selenium 75 in units not exceeding 10 microcuries (370 kBq) each.

(c) Each prepackaged unit bears a durable, clearly visible label:

1. Identifying the radioactive contents as to chemical form and radionuclide, and indicating that the amount of radioactivity does not exceed 10 microcuries (370 kBq) of iodine 125, iodine 131, carbon 14, cobalt 57 or selenium 75; 50 microcuries (1.85 MBq) of hydrogen 3 (tritium); 20 microcuries (740 kBq) of iron 59; or mock iodine 125 in units not exceeding 0.05 microcurie (1.85 kBq) of iodine 129 and 0.005 microcurie (185 Bq) of americium 241 each; and

2. Displaying the radiation caution symbol described in subsection 64E-5.322(1), F.A.C., and the words, “CAUTION, RADIOACTIVE MATERIAL”, and “Not for Internal or External Use in Humans or Animals”.

(d) One of the following statements, as appropriate, or a substantially similar statement which contains the information called for in one of the following statements, appears on a label affixed to each prepackaged unit or appears in a leaflet or brochure which accompanies the package:

1. This radioactive material may be received, acquired, possessed and used only by physicians, veterinarians, clinical laboratories or hospitals and only for in in vitro clinical or laboratory tests not involving internal or external administration of the material, or the radiation therefrom, to human beings or animals. Its receipt, acquisition, possession, use and transfer are subject to the regulations and a general license of the U.S. Nuclear Regulatory Commission or of a state with which the Commission has entered into an agreement for the exercise of regulatory authority.

____________________

Name of manufacturer

2. This radioactive material may be received, acquired, possessed and used only by physicians, veterinarians, clinical laboratories or hospitals and only for in vitro clinical or laboratory tests not involving internal or external administration of the material, or the radiation therefrom, to human beings or animals. Its receipt, acquisition, possession, use and transfer are subject to the regulations and a general license of a Licensing State.

____________________

Name of manufacturer

(e) The label affixed to the unit, or the leaflet or brochure which accompanies the package, contains adequate information as to the precautions to be observed in handling and storing such radioactive material. In the case of the mock iodine 125 reference or calibration source, the information accompanying the source must also contain directions to the licensee regarding the waste disposal requirements set out in Part III.

(f) The applicant satisfies the requirements specified in paragraph 64E-5.210(10)(b), F.A.C.

(9) Licensing the Manufacture and Distribution of Ice Detection Devices. An application for a specific license to manufacture and distribute ice detection devices to general licensees under subsection 64E-5.206(9), F.A.C., will be approved if:

(a) The applicant satisfies the general requirements of Rule 64E-5.208, F.A.C.; and

(b) The criteria of Sections 32.61, 32.62, and 32.103 of 10 C.F.R., Part 32, are met.

(10) Manufacture and Distribution of Radiopharmaceuticals Containing Radioactive Material for Medical Use. An application for a specific license to manufacture and distribute radiopharmaceuticals containing radioactive material for use by persons licensed pursuant to Part VI for the uses listed in Rules 64E-5.626, 64E-5.627, 64E-5.630 and 64E-5.664, F.A.C., will be approved if:

(a) The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.;

(b) The applicant submits evidence that:

1. The applicant is registered or licensed with the U.S. Food and Drug Administration as a drug manufacturer; or

2. The applicant is registered or licensed as a drug manufacturer as specified in Chapter 499, Florida Statutes; or

3. The applicant has a nuclear pharmacy permit and only authorized nuclear pharmacists compound or dispense radiopharmaceuticals as specified in Section 465.0193, F.S.

(c) The applicant submits information on the radionuclide, chemical and physical form, the maximum activity per vial, syringe, generator, or other container of the radioactive drug, and shielding provided by the packaging of the radioactive material which is appropriate for safe handling and storage of radiopharmaceuticals by medical use licensees;

(d) The applicant satisfies the following labeling requirements:

1. The label affixed to each transport radiation shield of any material of a radioactive drug transferred for commercial distribution includes the radiation symbol and the words “Caution, Radioactive Material” or “Danger, Radioactive Material”; the name of the radioactive drug or its abbreviation; and the quantity of the radioactive material at a specified date and time. The time can be omitted for radioactive drugs with a half life greater than 100 days.

2. A label affixed to each syringe, vial, or other container used to hold a radioactive drug transferred for commercial distribution includes the words “Caution, Radioactive Material” or “Danger, Radioactive Material” and an identifier that correlates the syringe, vial, or other container with the information on the transport radiation shield label; and

(e) A licensee shall possess and use instruments to measure the radioactivity of radioactive drugs. The licensee shall have procedures for use of the instruments. The licensee shall measure by direct measurements or by combination of measurements and calculations the amount of radioactivity in doses of alpha-emitting, beta-emitting, or photon-emitting radioactive drugs before transfer for commercial distribution. In addition, the licensee shall:

1. Perform tests before initial use, periodically, and following repair on each instrument for accuracy, linearity, and geometry dependence appropriate for the use of the instrument and make adjustments when needed; and

2. Check each instrument for constancy and proper operation at the beginning of each day of use.

(f) An application from a medical facility, educational institution, or Federal facility to produce Positron Emission Tomography (PET) radiopharmaceuticals containing for noncommercial transfer to licensees in its consortium licensed for medical pursuant to Part VI, or equivalent Agreement State, or NRC rules will be approved if:

1. The requirements of paragraphs 64E-5.210(10)(a), (b), and (e), F.A.C., are satisfied;

2. The information required of paragraphs 64E-5.210(10)(c) and (d), F.A.C., indicates the PET drugs to be noncommercially transferred to members of its consortium.

(11) Manufacture and Distribution of Generators or Reagent Kits for Preparation of Radiopharmaceuticals Containing Radioactive Material. An application for a specific license to manufacture and distribute generators or reagent kits containing radioactive material for preparation of radiopharmaceuticals by persons licensed pursuant to Part VI for the uses listed in Rule 64E-5.627 or 64E-5.664, F.A.C., will be approved if:

(a) The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.;

(b) The applicant submits evidence that:

1. The generator or reagent kit is to be manufactured, labeled and packaged in accordance with the Federal Food, Drug and Cosmetic Act or the Public Health Service Act, such as a new drug application (NDA) approved by the Food and Drug Administration (FDA), or a “Notice of Claimed Investigational Exemption for a New Drug” (IND) that has been accepted by the FDA; or

2. The manufacture and distribution of the generator or reagent kit are not subject to the Federal Food, Drug and Cosmetic Act and the Public Health Service Act;

(c) The applicant submits information on the radionuclide, chemical and physical form, packaging including maximum activity per package, and shielding provided by the packaging of the radioactive material contained in the generator or reagent kit;

(d) The label affixed to the generator or reagent kit contains information on the radionuclide, quantity and date of assay; and

(e) The label affixed to the generator or reagent kit, or the leaflet or brochure which accompanies the generator or reagent kit, contains:

1. Adequate information pertaining to radiation safety on the procedures to be followed and the equipment and shielding to be used in eluting the generator or processing radioactive material with the reagent kit, and

2. A statement that this generator or reagent kit, as appropriate, is approved for use by persons licensed by the department pursuant to Part VI for uses listed in Rule 64E-5.627, F.A.C., or under equivalent licenses of the U.S. Nuclear Regulatory Commission, an agreement state or a licensing state. The labels, leaflets, or brochures required by this section are in addition to the labeling required by the Food and Drug Administration (FDA) and they may be separate from or, with the approval of FDA, may be combined with the labeling required by FDA.

(12) Manufacture and Distribution of Sources or Devices Containing Radioactive Material for Medical Use.

(a) An application for a specific license to manufacture and distribute sources and devices containing radioactive material to persons licensed pursuant to Part VI for use as a calibration, transmission or reference source or for the uses listed in Rule 64E-5.631, 64E-5.634, 64E-5.664 or 64E-5.632, F.A.C., will be approved if:

1. The applicant satisfies the general requirements in Rule 64E-5.208, F.A.C.;

2. The applicant submits sufficient information regarding each type of source or device pertinent to an evaluation of its radiation safety, including:

a. The radioactive material contained, its chemical and physical form, and amount,

b. Details of design and construction of the source or device,

c. Procedures for, and results of, prototype tests to demonstrate that the source or device will maintain its integrity under stresses likely to be encountered in normal use and accidents,

d. For devices containing radioactive material, the radiation profile of a prototype device,

e. Details of quality control procedures to assure that production sources and devices meet the standards of the design and prototype tests,

f. Procedures and standards for calibrating sources and devices,

g. Legend and methods for labeling sources and devices as to their radioactive content, and

h. Instructions pertaining to radiation safety for handling and storing the source or device; these instructions are to be included on a durable label attached to the source or device or attached to a permanent storage container for the source or device; provided, that instructions which are too lengthy for such label may be summarized on the label and printed in detail on a brochure which is referenced on the label;

3. The label affixed to the source or device, or to the permanent storage container for the source or device, contains information on the radionuclide, quantity, date of assay, and a statement that the name of source or device is licensed by the department for distribution to persons licensed pursuant to Part VI or under equivalent licenses of the U.S. Nuclear Regulatory Commission, an agreement state or a licensing state, provided, that such labeling for sources which do not require long term storage may be on a leaflet or brochure which accompanies the source;

(b) In the event the applicant desires that the source or device be required to be tested for leakage of radioactive material at intervals longer than 6 months, he shall include in his application sufficient information to demonstrate that such longer interval is justified by performance characteristics of the source or device or similar sources or devices and by design features that have a significant bearing on the probability or consequences of leakage of radioactive material from the source; and

(c) In determining the acceptable interval for test of leakage of radioactive material, the Department will consider the following information:

1. Primary containment or source capsule,

2. Protection of primary containment,

3. Method of sealing containment,

4. Containment construction materials,

5. Form of contained radioactive material,

6. Maximum temperature withstood during prototype tests,

7. Maximum pressure withstood during prototype tests,

8. Maximum quantity of contained radioactive material,

9. Radiotoxicity of contained radioactive material, and

10. Operating experience with identical sources or devices or similarly designed and constructed sources or devices.

(13) Requirements for License to Manufacture and Distribute Industrial Products Containing Depleted Uranium for Mass-Volume Applications.

(a) An application for a specific license to manufacture industrial products and devices containing depleted uranium for use pursuant to subsection 64E-5.205(4), F.A.C., or equivalent regulations of the U.S. Nuclear Regulatory Commission or an Agreement State will be approved if:

1. The applicant satisfies the general requirements specified in Rule 64E-5.208, F.A.C.;

2. The applicant submits sufficient information relating to the design, manufacture, prototype testing, quality control procedures, labeling or marking, proposed uses, and potential hazards of the industrial product or device to provide reasonable assurance that possession, use or transfer of the depleted uranium in the product or device is not likely to cause any individual to receive in any period of 1 calendar quarter a radiation dose in excess of 10 percent of the limits specified in Subpart III A of these rules; and

3. The applicant submits sufficient information regarding the industrial product or device and the presence of depleted uranium for a mass-volume application in the product or device to provide reasonable assurance that unique benefits will accrue to the public because of the usefulness of the product or device.

(b) In the case of an industrial product or device whose unique benefits have not been demonstrated, the Department will approve an application for a specific license under this subsection only if the product or device is found to combine a high degree of utility and low probability of uncontrolled disposal and dispersal of significant quantities of depleted uranium into the environment.

(c) Each person licensed pursuant to paragraph (13)(a), above, shall:

1. Maintain the level of quality control required by the license in the manufacture of the industrial product or device and in the installation of the depleted uranium into the product or device;

2. Label or mark each unit to:

a. Identify the manufacturer of the product or device and the number of the license under which the product or device was manufactured, the fact that the product or device contains depleted uranium and the quantity of depleted uranium in each product or device; and

b. State that receipt, possession, use and transfer of the product or device are subject to a general license or the equivalent and the regulations of the U.S. Nuclear Regulatory Commission or an Agreement State;

3. Assure that the depleted uranium before being installed in each product or device has been impressed with the following legend clearly legible through any plating or other covering: “Depleted Uranium”;

4.a. Furnish a copy of the general license described in subsection 64E-5.205(4), F.A.C., to each person to whom he transfers depleted uranium in a product or device for use pursuant to the general license described in subsection 64E-5.205(4), F.A.C., or

b. Furnish a copy of the general license certificate of the U.S. Nuclear Regulatory Commission’s or an Agreement State’s, or alternatively, furnish a copy of the general license described in subsection 64E-5.205(4), F.A.C., to each person to whom he transfers depleted uranium in a product or device for use pursuant to the general license of the U.S. Nuclear Regulatory Commission or an Agreement State, with a note explaining that use of the product or device is regulated by the U.S. Nuclear Regulatory Commission or an Agreement State under requirements substantially the same as those in subsection 64E-5.205(4), F.A.C.;

5. Report to the Department all transfers of industrial products or devices to persons for use under the general license described in subsection 64E-5.205(4), F.A.C. Such report shall identify each general licensee by name and address, an individual by name or position who may constitute a point of contact between the Department and the general licensee, the type and model number of device transferred, and the quantity of depleted uranium contained in the product or device. The report shall be submitted within 30 days after the end of each calendar quarter in which such a product or device is transferred to the general licensee. If no transfers have been made to general licensees under subsection 64E-5.205(4), F.A.C., during the reporting period, the report shall so indicate;

6.a. Report to the U.S. Nuclear Regulatory Commission all transfers of industrial products or devices to persons for use under the U.S. Nuclear Regulatory Commission general license in Section 40.25 of 10 C.F.R., Part 40,

b. Report to the responsible state agency all transfers of devices manufactured and distributed pursuant to subsection 64E-5.210(3), F.A.C., for use under a general license in that state’s rules equivalent to subsection 64E-5.205(4), F.A.C.,

c. Such report shall identify each general licensee by name and address, an individual by name or position who may constitute a point of contact between the agency and the general licensee, the type and model number of the device transferred, and the quantity of depleted uranium contained in the product or device. The report shall be submitted within 30 days after the end of each calendar quarter in which such product or device is transferred to the general licensee.

d. If no transfers have been made to U.S. Nuclear Regulatory Commission licensees during the reporting period, this information shall be reported to the U.S. Nuclear Regulatory Commission, and

e. If no transfers have been made to general licensees within a particular Agreement State during the reporting period, this information shall be reported to the responsible Agreement State agency upon the request of that agency; and

7. Keep records showing the name, address and point of contact for each general licensee to whom he transfers depleted uranium in industrial products or devices for use pursuant to the general license provided in subsection 64E-5.205(4), F.A.C., or equivalent regulations of the U.S. Nuclear Regulatory Commission or an Agreement State. The records shall be maintained for a period of 2 years and shall show the date of each transfer, the quantity of depleted uranium in each product or device transferred and compliance with the report requirements of this section.

(14) A licensee, manufacturer or an initial distributor of a sealed source or device containing a sealed source whose product contains exempt NARM or is intended for use under a general or specific license must submit a request for an evaluation of the sealed source or device containing a sealed source and obtain a registration from the Department.

(a) The request for review of a sealed source or device must be made in triplicate and include information about the design, manufacture, prototype testing, quality control and assurance program, labeling, leak testing and proposed uses. The licensee shall inform customers of current reasonable disposal options for the radioactive material.

(b) The request for review of a device must include information about installation, service and maintenance, operating and safety instructions, and its potential hazards. The information shall provide reasonable assurance that the radiation safety properties of the source or device are adequate to protect public health, safety and property.

(c) The Department shall use criteria and standards sufficient to ensure that the radiation safety properties of the sealed source or device are adequate to protect public health, safety and property. Criteria and standards used by the Department in evaluating a sealed source or device include:

1. U.S. Department of Health and Human Services Publication FDA 81-8025 June 1981, Guides for Naturally Occurring and Accelerator-Produced Radioactive Materials (NARM), which is herein incorporated by reference and can be obtained from the internet at https://www.flrules.org/gateway/reference.asp?No=Ref-03508 or at http://www.doh.state.fl.us/environment/radiation/matform.htm which is available from the department.

2. NRC Regulatory Guide 10.10 March 1987, Guide for the Preparation of Applications for Radiation Safety Evaluations and Registration of Devices Containing Byproduct Material, which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03453 or at http://pbadupws.nrc.gov/docs/ML0037/ML003740220.pdf or at http://www.doh.state.fl.us/environment/radiation/.

3. NRC Regulatory Guide 10.11 June 1987, Guide for the Preparation of Applications for Radiation Safety Evaluations of Sealed Sources Containing Byproduct Material, which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03454 or at http://pbadupws.nrc.gov/docs/ML0037/ML003740233.pdf or at http://www.doh.state.fl.us/environment/radiation/.

4. American National Standards Institute (ANSI) Standard, ANSI-HPS N43.8-2008, Classification of Industrial Ionizing Radiation Gauging Devices, which is herein incorporated by reference and can be obtained from the internet at http://hps.org/hpssc/documents/ansi_standards_order_form.pdf.

5. ANSI Standard, ANSI-HPS N43.4-2005, Classification of Radioactive Self-Luminous Light Sources, which is herein incorporated by reference and can be obtained from the internet at http://hps.org/hpssc/documents/ansi_standards_order_form.pdf.

6. ANSI Standard N432-1980, NBS Handbook 136, as issued in January 1981, Radiological Safety for the Design and Construction of Apparatus for Gamma Radiography, which is herein incorporated by reference and can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-03455 or at http://pbadupws.nrc.gov/docs/ML0508/ML050840139.pdf or at http://www.doh.state.fl.us/environment/radiation/matform.htm which is available from the department.

7. ANSI Standard, ANSI-HPS N43.6-2007, Sealed Radioactive Sources Classification, which is herein incorporated by reference and can be obtained from the internet at http://hps.org/hpssc/documents/ansi_standards_order_form.pdf. The Ansi publications referenced in this rule section: ANSI-HPS N43.8-2008; ANSI-HPS N43.4-2005; ANSI-HPS N43.6-2007; are copyrighted materials. These materials are available for public inspection and examination at the Florida Department of State, Room 701, The Capitol, Tallahassee, Florida 32399-0250, and at the Florida Department of Health, Bureau of Radiation Control, 4042 Bald Cypress Way, Tallahassee, Florida 32399-1741.

(d) The licensee or applicant shall not distribute devices or products containing sealed sources unless the devices or sealed sources are manufactured and distributed in accordance with the registration and as authorized by a specific radioactive materials license issued by the department for such manufacture or distribution.

(e) The department shall not perform registration of devices or products containing sealed sources for persons outside the state.

(15) Each licensee who manufactures a nationally tracked source after February 6, 2007 shall assign a unique serial number to each nationally tracked source. Serial numbers must be composed only of alpha-numeric characters.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.051, 404.061, 404.081, 404.141 FS. History–New 7-17-85, Amended 8-25-91, 5-12-93, 1-1-94, 5-15-96, Formerly 10D-91.311, Amended 8-6-01, 9-28-06, 8-16-07, 2-28-08, 2-11-10, 12-26-13.

- 64E-5.211 Special Requirements for Issuance of Specific Licenses for Source Material Milling.

In addition to the requirements set forth in Rule 64E-5.208, F.A.C., a specific license for source material milling will be issued if the applicant submits to the Department an application as described herein and meets the other conditions specified below:

(1) An application for a license to receive title to, receive, possess and use source material for milling or by product material as defined in Part I shall address the following:

(a) Description of the proposed project or action;

(b) Area or site characteristics including geology, topography, hydrology and meteorology;

(c) Radiological and nonradiological impacts of the proposed project or action, including waterway and groundwater impacts;

(d) Environmental effects of accidents;

(e) Long-term impacts including decommissioning, decontamination and reclamation; and

(f) Site and project alternatives.

(2) The applicant shall not commence construction of the project until the Department has weighed the environmental, economic, technical and other benefits against the environmental costs and has concluded that the issuance of the license is appropriate.

(3) At least 1 full year prior to any major site construction, a preoperational monitoring program shall be conducted to provide complete baseline data on a milling site and its environs. Throughout the construction and operating phases of the mill, an operational monitoring program shall be conducted to measure or evaluate compliance with applicable standards and regulations; to evaluate performance of control systems and procedures; to evaluate environmental impacts of operation; and to detect potential longterm effects.

(4) Prior to issuance of the license, the applicant shall establish financial surety arrangements consistent with the requirements of Rule 64E-5.217, F.A.C.

(a) The amount of funds to be insured by financial surety arrangements shall be based on cost estimates which are furnished by the licensee and which the Department shall evaluate to determine that the cost estimates are reasonably comparable to other decontamination or decommissioning estimates in a plan for decontamination and decommissioning of mill buildings and the milling site to levels which would allow unrestricted use of these areas upon decommissioning, and the reclamation of tailings and waste disposal areas. The licensee shall submit this plan in conjunction with an environmental report that addresses the expected environmental impacts of the milling operation, decommissioning and tailings reclamation, and that evaluates alternatives for mitigating these impacts. In establishing specific surety arrangements, the licensee’s cost estimates shall take into account total costs that would be incurred if an independent contractor were hired to perform the decommissioning and reclamation work. In order to avoid unnecessary duplication and expense, the Department may accept financial sureties that have been consolidated with financial surety arrangements established to meet requirements of other Federal or state agencies or local governing bodies for such decommissioning, decontamination, reclamation and long-term site surveillance, provided such arrangements are considered adequate to satisfy these requirements and that portion of the surety which covers the decommissioning and reclamation of the mill, mill tailings site and associated areas, and the long-term funding charge are clearly identified. The licensee’s surety mechanism will be reviewed annually by the Department to assure that sufficient funds will be available for completion of the reclamation plan if the work had to be performed by an independent contractor. The amount of surety liability should be adjusted to recognize any increases or decreases resulting from inflation, changes in engineering plans, activities performed and any other conditions affecting costs. Regardless of whether reclamation is phased through the life of the operation or takes place at the end of operations, an appropriate portion of surety liability shall be retained until final compliance with the reclamation plan is determined. This will yield a surety that is at least sufficient at all times to cover the costs of decommissioning, decontamination and reclamation of the areas that are expected to be disturbed before the next license renewal. The term of the surety mechanism must be open ended, unless it can be demonstrated that another arrangement would provide an equivalent level of assurance. This assurance could be provided with a surety instrument which is written for a specified period of time which must be automatically renewed unless the surety agent notifies the beneficiary, the Department and the licensee prior to the renewal date of their intention not to renew. In such a situation, the surety requirement still exists and the licensee would be required to submit an acceptable replacement surety within a brief period of time to allow at least 60 days for the Department to collect.

(b) The total amount of funds for reclamation or long term surveillance and control shall be transferred to the United States if title and custody of such material and its disposal site is transferred to the United States upon termination of a license. Such funds include sums collected for long term surveillance and control. Such funds do not, however, include monies held as surety where no default has occurred and the reclamation or other bonded activity has been performed.

(5) The applicant shall provide procedures describing the means employed to meet the following requirements during the operational phase of any project.

(a) Milling operations shall be conducted so that all effluent releases are below the limits of Part III and are as low as is reasonably achievable.

(b) The mill operator shall conduct daily inspections of any tailings or waste retention systems. Such inspections shall be conducted by a licensed engineer. Records of such inspections shall be maintained for review by the Department.

(c) The mill operator shall immediately notify the Department of the following:

1. Any failure in a tailings or waste retention system which results in a release of tailings or waste into unrestricted areas, and

2. Any unusual condition not contemplated in the design of the retention system which, if not corrected, could lead to failure of the system and result in a release of tailings or waste into unrestricted areas.

(6) Continued Surveillance Requirements for Source Material Mills Having Reclaimed Residues.

(a) The final disposition of tailings or wastes at source material milling sites should be such that the need for ongoing active maintenance is not necessary to preserve isolation. As a minimum, annual site inspections shall be conducted by the government agency retaining ultimate custody of the site where tailings or wastes are stored to confirm the integrity of the stabilized tailings or waste systems and to determine the need, if any, for maintenance and monitoring. Results of the inspection shall be reported to the U.S. Nuclear Regulatory Commission within 60 days following each inspection. The U.S. Nuclear Regulatory Commission may require more frequent site inspections, if, on the basis of a site-specific evaluation, such a need appears necessary due to the features of a particular tailings or waste disposal system.

(b) A minimum charge of $405,000 to cover the costs of long-term surveillance shall be paid by each mill operator to the Department prior to the termination of a uranium or thorium mill license. If site surveillance or control requirements at a particular site are determined, on the basis of a site-specific evaluation, to be significantly greater than those specified in paragraph (6)(a), above, additional funding requirements may be specified by the Department. The total charge to cover the cost of long-term surveillance shall be such that, with an assumed 1 percent annual real interest rate, the collected funds will yield interest in an amount sufficient to cover the annual costs of site surveillance. The charge will be assessed quarterly and will be reviewed annually by the Department to recognize or adjust for inflation.

Rulemaking Authority 404.051, 404.061, 404.0612, 404.071, 404.081, 404.111, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (7), (8), (11), 404.061(2), 404.0612, 404.071(1), 404.081(1), 404.111, 404.141 FS. History–New 7-17-85, Formerly 10D-91.312.

- 64E-5.212 Issuance of Specific Licenses.

(1) Upon a determination that an application meets the requirements of Chapter 404, F.S., and these regulations, the Department will issue a specific license authorizing the proposed activity in such form and containing such conditions and limitations as it deems appropriate or necessary.

(2) The Department may incorporate in any license at the time of issuance, or thereafter by appropriate rule, regulation or order, such additional requirements and conditions with respect to the licensee’s receipt, possession, use and transfer of radioactive material subject to this part as it deems appropriate or necessary in order to:

(a) Minimize danger to public health and safety or property;

(b) Require reports and the keeping of records, and to provide for inspections of activities under the license; and

(c) Prevent loss or theft of material subject to this part.

(3) The department shall issue an expiration date authorizing each license to be valid for a period not to exceed 5 years from the last day of the issuance month. The department shall indicate the expiration date on each license. The licensee shall be granted a 90 day extension of the expiration date if written justification is submitted and approved by the department.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (8), 404.081(1), 404.141 FS. History–New 7-17-85, Amended 5-12-93, Formerly 10D-91.313.

- 64E-5.213 Specific Terms and Conditions of License.

(1) Each license issued pursuant to this part shall be subject to all the provisions of the applicable laws, now or hereafter in effect, and to all rules of the Department.

(2) No license issued or granted under this part and no right to possess or utilize radioactive material granted by any license issued pursuant to this part shall be transferred, assigned or in any manner disposed of, either voluntarily or involuntarily, directly or indirectly, through transfer of control to any person unless the Department, after securing a completed specific license application and application fee from the transferee, has issued a proper license in accordance with the provisions of the Act.

(3)(a) Each specific or general licensee shall notify the Department in writing immediately following the filing of a voluntary or involuntary petition for bankruptcy under any chapter of Title 11 (Bankruptcy) of the United States Code (U.S.C.) by or against:

1. The licensee;

2. An entity, as that term is defined in 11 U.S.C. 101(14), controlling the licensee or listing the license or licensee as property of the estate; or

3. An affiliate, as that term is defined in 11 U.S.C. 101(2), of the licensee.

(b) This notification shall indicate the bankruptcy court in which the petition for bankruptcy was filed and the date of the filing of the petition for bankruptcy.

(4)(a) Each person licensed by the Department pursuant to this part shall confine use and possession of the material licensed to the locations and purposes authorized in the license.

(b) Each person specifically licensed by the Department shall maintain a fixed facility located within the state of Florida.

(5) A separate license is required for the following:

(a) Each activity as designated by license category in paragraph 64E-5.204(2)(e), F.A.C.

(b) Facilities for which one or more of the following applies:

1. The facilities are not contiguous;

2. The facilities are not under a single radiation safety program; or

3. The facilities are not under the same management.

4. Temporary job sites lasting more than two years.

(c) Each facility operated by an out-of-state licensee under reciprocity as specified in Rule 64E-5.216, F.A.C., and does not meet the definition of a temporary job site.

(d) Each large irradiator as defined in Rule 64E-5.101, F.A.C.

(6) A separate license is not required for temporary job sites lasting less than two years or for each facility that is authorized under a broad scope license.

(7) A licensee shall notify the department in writing within 30 days after a radiation safety officer permanently discontinues performance of radiation safety officer duties.

(8) A licensee shall apply and receive a license amendment or Department approval:

(a) Before using radioactive material for a method or type or use not permitted by the license;

(b) Before permitting anyone to use radioactive material as an authorized user as authorized by the license;

(c) Before changing a radiation safety officer;

(d) Before ordering or receiving radioactive material in excess of the amount authorized on the license;

(e) Before adding to or changing the areas of use or address or addresses of use identified in the application or on the license; and

(f) Before changing statements, representations, and procedures which are incorporated into the license.

(g) Identifying all sources or devices by manufacturer and model number as registered by the sealed source and device registry or for sources or devices not registered by the sealed source and device registry provide the information in subsection 64E-5.210(14), F.A.C.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.051(1), (4), 404.061(2), (3), 404.141 FS. History–New 7-17-85, Amended 4-4-89, 5-12-93, 8-29-94, Formerly 10D-91.314, Amended 5-18-98, 9-28-06, 2-11-10, 12-26-13.

- 64E-5.214 Expiration and Termination of Licenses and Decommissioning of Sites and Separate Buildings or Outdoor Areas.

(1) Except as provided in Part II, each specific license shall expire at the end of the specified day in the month and year stated therein. Each specific license revoked by the department expires at the end of the day on the date of the department’s final order revoking the license or on the expiration date stated in the final order.

(2)(a) Each licensee shall notify the department in writing within 60 days of the occurrence of any of the following and either begin decommissioning its site or any separate building or outdoor area that contains residual radioactivity so that the building or outdoor area is suitable for release as specified in these rules or send a notice of a decommissioning plan within 12 months as specified in paragraph (4)(c) below and begin decommissioning upon approval of that plan.

1. The license has expired as specified in subsection (1), above.

2. The licensee has ceased principal activities permanently at the entire site or in any separate building or outdoor area.

3. The licensee has conducted no principal activities under the license for 24 months.

4. The licensee has conducted no principal activities for 24 months in any separate building or outdoor area that contains residual radioactivity to the extent that the building or outdoor area is unsuitable for release as specified in these rules.

(b) The notification and request for termination of the license shall include the reports and information specified in subparagraph (4)(a)4. and 5., below.

(3) No less than 30 days before the expiration date specified in the license, the licensee shall either:

(a) Submit an application for license renewal on the same form used for the initial application under Part II, or

(b) Notify the Department, in writing, if the licensee decides not to apply for license renewal.

(4)(a) If a licensee does not submit an application for license renewal under Part II, the licensee shall, on or before the expiration date specified in the license:

1. Terminate the use of radioactive material;

2. Remove residual radioactivity to the extent acceptable to the Department;

3. Properly dispose of the radioactive material;

4. Submit a properly completed DH Form 1059, which is herein incorporated by reference effective 7-17-85; and

5. Submit a radiation survey report to confirm the absence of radioactive materials or to establish the levels of residual radioactivity, unless the licensee demonstrates the absence of residual radioactivity in some other manner. The licensee shall, as appropriate:

a. For gamma radiation, report levels of radiation in units of microroentgens per hour at 10 centimeters and at 1 meter from surfaces.

b. For alpha and beta radiation, report levels of radioactivity in units of transformations per minute or microcuries per 100 square centimeters removable and fixed on surfaces, microcuries per milliliter in water, and picocuries per gram in contaminated solids such as soils or concrete; and

c. Specify the instruments used and certify that each instrument is properly calibrated or tested.

(b)1. If no residual radioactivity attributable to activities conducted under the license is detected, the licensee shall submit a certification that no detectable residual radioactivity was found.

2. Specific licenses including expired licenses will be terminated by written notice to the licensee when the department determines that:

a. Radioactive material has been properly disposed; and

b. A radiation survey has been performed which demonstrates that the premises are suitable for release for unrestricted use or satisfies the requirements specified in Rule 64E-5.221, 64E-5.222, 64E-5.223, or 64E-5.224, F.A.C.; or

c. Other information submitted by the licensee is sufficient to demonstrate that the premises are suitable for release for unrestricted use or satisfies the requirements specified in Rule 64E-5.221, 64E-5.222, 64E-5.223 or 64E-5.224, F.A.C.; or

d. Department has received the following records, if requested:

(I) Disposal records specified in Rule 64E-5.330, paragraphs 64E-5.331(1)(a), (c), (2), (3), or 64E-5.336(2)(d), F.A.C.; and

(II) Records specified in subsection 64E-5.214(6), F.A.C.

(c)1. If detectable levels of residual radioactivity attributable to activities conducted under the license are found or licensee possesses other radioactive materials, the license continues in effect beyond the expiration date, if necessary, with respect to possession of residual radioactivity present or possession of radioactive material, until the department notifies the licensee in writing that the license is terminated. During this time, the licensee is subject to the provisions of subsection (5), below.

2. In addition to the information submitted under (4)(a)4. and 5., above, the licensee shall submit a plan for decommissioning if decommissioning procedures have not been approved previously by the department and could impact the health and safety of workers or the public as follows:

a. More than routine cleanup and maintenance is required;

b. Workers will be in areas with significantly increased surface contamination or radiation levels;

c. Procedures will result in significantly greater airborne concentrations of radioactive materials; or

d. Procedures will result in significantly greater releases of radioactive material to the environment.

3. Procedures which could potentially impact health, safety and the environment may not be performed until the decommissioning plan has been approved.

4. The proposed decommissioning plan must include:

a. A description of the planned decommissioning activities;

b. A description of the methods used to assure protection of workers and the environment against radiation hazards during decommissioning;

c. The time required to complete the decommissioning plan; and

d. A description of the planned final radiation survey.

5. The proposed decommissioning plan will be reviewed by the department and approved or additional information will be requested within 60 days.

6. Upon approval of the decommissioning plan by the department, the licensee shall complete decommissioning in accordance with the approved plan. As a final step in decommissioning, the licensee shall again submit the information required in subparagraph (4)(a)5., above, of this section and shall certify the disposition of accumulated wastes from decommissioning.

7. If the information submitted as specified in subparagraph (4)(a)5. or (4)(c)6. of this section does not adequately demonstrate that the premises are suitable for unrestricted use or does not satisfy the requirements specified in Rule 64E-5.221, 64E-5.222, 64E-5.223, or 64E-5.224, F.A.C., the department will inform the licensee of the appropriate further actions required for termination of the license.

(5) Each licensee who possesses radioactive material under paragraph (4)(c), above, following the expiration date specified in the license shall:

(a) Limit actions involving radioactive material to those related to decontamination, decommissioning, and other activities related to preparation for release for unrestricted use; and

(b) Continue to control entry to restricted areas until they are suitable for release for unrestricted use and the Department notifies the licensee, in writing, that the license is terminated.

(6) Each licensee shall keep records of the decommissioning of the facility in an identified location until the license is terminated by the department. If records of relevant information are kept for other purposes, reference to these records and their location can be used. Records which must be kept include:

(a) Records of spills or other unusual occurrences involving the spread of contamination in and around the facility, equipment, or site. These records can be limited to instances when contamination remains after cleanup procedures or when contaminants have spread to inaccessible areas such as possible seepage into porous materials such as concrete. These records must include any known information on identification of involved nuclides, quantities, forms, and concentrations;

(b) Drawings of structures as originally built, of modifications, and of equipment in restricted areas where radioactive materials are used or stored, and of locations of possible inaccessible contamination such as buried pipes which can be subject to contamination. Drawings and their location can be referenced if not on site. If drawings are not available, the licensee shall substitute appropriate records of available information concerning these areas and locations.

(c) Except for areas containing only radioactive materials having half-lives of less than 65 days or sealed sources that either have not leaked or no contamination remains after any leak, a list contained in a single document and updated every 2 years, of the following:

1. All areas designated and formerly designated restricted areas as defined in Rule 64E-5.101, F.A.C.;

2. All areas outside of restricted areas that require documentation under paragraph 64E-5.214(6)(a), F.A.C.;

3. All areas outside of restricted areas where current and previous wastes have been buried as documented under Rule 64E-5.340, F.A.C.; and

4. All areas outside of restricted areas which contain material such that, if the license expired, the licensee would be required to either decontaminate the area to unrestricted release levels or satisfy the requirements specified in Rule 64E-5.221, 64E-5.222, 64E-5.223, or 64E-5.224, F.A.C.; and

(d) Records of the cost estimate performed for the performance bond required in Rule 64E-5.217, F.A.C., and records of the funding method used.

(7) Confirmatory or closeout surveys will be performed by the department according to the Closeout Inspection and Survey Procedures, November 1991, which are herein incorporated by reference and which are available from the department.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (4), (6), (9), 404.061(2), 404.081(1) FS. History– New 7-17-85, Amended 5-12-93, 8-14-96, Formerly 10D-91.315, Amended 5-18-98, 10-8-00, 12-19-01.

- 64E-5.215 Transfer of Material.

(1) No licensee shall transfer radioactive material except as authorized pursuant to this section.

(2) Except as otherwise provided in his license and subject to the provisions of subsections (3) and (4), below, a licensee may transfer radioactive material:

(a) To the Department after receiving approval from the Department;

(b) To the U.S. Department of Energy;

(c) To any person exempt from these regulations to the extent permitted under such exemption;

(d) To any person authorized to receive such material under terms of a general license or its equivalent, or a specific license or equivalent licensing document, issued by the Department, the U.S. Nuclear Regulatory Commission, an Agreement State, a Licensing State, or to any person otherwise authorized to receive such material by the Federal Government or any agency thereof, the Department, an Agreement State or a Licensing State.

(3) Before transferring radioactive material to a specific licensee of the Department, the U.S. Nuclear Regulatory Commission, an Agreement State, a Licensing State or to a general licensee who is required to register with the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State, prior to receipt of the radioactive material, the licensee transferring the material shall verify that the transferee’s license authorizes the receipt of the type, form and quantity of radioactive material to be transferred.

(4) Any of the following methods for the verification required by subsection (3), above, are applicable:

(a) The transferor may possess and read a current copy of the transferee’s specific or general license.

(b) The transferor may possess a written certification by the transferee that the transferee is authorized by license to receive the type, form and quantity of radioactive material to be transferred, specifying the license number, issuing agency and expiration date.

(c) For emergency shipments, the transferor may accept oral certification by the transferee that the transferee is authorized by license to receive the type, form and quantity of radioactive material to be transferred, specifying the license number, issuing agency, and expiration date; provided, that the oral certification is confirmed in writing within 10 days.

(d) The transferor may obtain other information compiled by a reporting service from official records of the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State regarding the identity of licensees and the scope and expiration dates of the licenses.

(e) When none of the methods of verification described in paragraphs (4)(a) through (d), above, are readily available or when a transferor desires to verify that information received by one of such methods is correct or up-to-date, the transferor may obtain and record confirmation for the Department, the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State that the transferee is licensed to receive the radioactive material.

(5) Shipment and transport of radioactive material shall be in accordance with the provisions of Part XV.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141, 404.20 FS. Law Implemented 404.022, 404.051(1), (2), (4), (11), 404.061(2), 404.081(1), 404.20(1) FS. History–New 7-17-85, Formerly 10D-91.319.

- 64E-5.216 Reciprocal Recognition of Licenses for Byproduct, Source, Naturally Occurring and Accelerator Produced Radioactive Material, and Special Nuclear Material in Quantities Not Sufficient to Form a Critical Mass.

(1) Subject to these regulations, any person who holds a specific license from the NRC, or an Agreement State and issued by the agency having jurisdiction where the licensee maintains an office for directing the licensed activity and at which radiation safety records are normally maintained, will be granted a general license by the Department to conduct the activities authorized in such licensing document within the State of Florida, except for areas of exclusive federal jurisdiction, for a period not in excess of 180 consecutive days provided that:

(a) The out-of-state license document does not limit the performance of the function authorized by such document to specified installations or locations;

(b) The out-of-state licensee notifies the Department in writing at least 3 days prior to engaging in such activity. Such notification shall indicate the location, period and type of proposed possession and use within the State, and shall be accompanied by a copy of the pertinent licensing document. If, for a specific case, the 3-day period would impose an undue hardship on the out-of-state licensee, the licensee may, upon application to the Department, obtain permission to proceed sooner;

(c) The out-of-state licensee complies with these applicable regulations and with all the terms and conditions of the licensing document, except any such terms and conditions that are inconsistent with these applicable regulations; and

(d) The out-of-state licensee shall not transfer or dispose of radioactive material possessed or used under the general license provided in this section except by transfer to a person who is specifically licensed by the Department, by the NRC, an Agreement State or a Licensing State to receive such material.

(e) Any licensee using or storing radioactive material at any location not listed on the license for a period in excess of 180 days in a calendar year, shall notify the department with the information listed in paragraph 64E-5.216(1)(b), F.A.C., prior to exceeding the 180 days.

(2) In addition to the provisions of subsection (1), above, any person who holds a specific license issued by the NRC, an agreement state, or a licensing state authorizing the holder to manufacture, transfer, install or service a device described in paragraph 64E-5.206(4)(a), F.A.C., within areas subject to the jurisdiction of the licensing body may be granted a general license by the Department to install, transfer, demonstrate or service such a device in this State provided that:

(a) Such person shall file a report with the Department within 30 days after the end of each calendar quarter in which any device is transferred to or installed in this State. Each such report shall identify each general licensee to whom such device is transferred by name and address, the type of radioactive material contained in the device;

(b) The device has been manufactured, labeled, installed and serviced in accordance with applicable provisions of the specific license issued to such person by the U.S. Nuclear Regulatory Commission, an Agreement State or a Licensing State;

(c) Such person shall assure that any labels required to be affixed to the device under regulations of the authority which licensed manufacture of the device bear a statement that “Removal of this label is prohibited”; and

(d) The holder of the specific license shall furnish to each general licensee to whom he transfers such device, or on whose premises he installs such device, a copy of the general license contained in subsection 64E-5.206(4), F.A.C., or in equivalent regulations of the agency having jurisdiction over the manufacture and distribution of the device.

(3) The Department may withdraw, limit or qualify its acceptance of any specific license or equivalent licensing document issued by the U.S. Nuclear Regulatory Commission, an Agreement State, or a Licensing State, or any product distributed pursuant to such licensing document, upon determining that such action is necessary in order to prevent undue hazard to public health, safety or property.

Rulemaking Authority 404.051(4), (11), 404.061(2) FS. Law Implemented 404.051(1), (2), (4), (6), (11), 404.061(2), 404.081(1) FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.321, Amended 10-8-00, 2-28-08, 2-11-10, 12-26-13.

- 64E-5.217 Bonding of Persons Licensed Pursuant to Subpart II C.

(1) Any applicant or licensee who is not exempt by the provisions of this subpart shall provide a performance bond.

(a) The bond shall be payable to the State of Florida and shall be in an amount determined by the Department as sufficient to provide for the protection of the environment and the public health and safety in the event of abandonment, insolvency or other inability of the licensee to meet the requirements of the Department. The Department shall use the Bond Risk Factors Calculation Worksheet ‒ March 2014 which is herein incorporated by reference and which can be obtained from the internet at http://www.flrules.org/Gateway/reference.asp?No=Ref-05470 or at http://www.floridahealth.gov/prevention-safety-and-wellness/radiation-control/, to determine the amount of the bond required for each applicant or licensee. The mathematical product of the risk factors will be the amount of the required bond in dollars. In the event that an applicant or licensee feels that the amount of the bond determined by the use of the applicable risk factors is inappropriate, he may submit evidence to the Department in support of a change to the amount. The Department shall determine whether the evidence supports the requested change in the amount.

(b) Licensees must provide the required bond within 90 days after being given notice by the Department of the requirements of a bond and its amount.

(c) The Department may re-evaluate, at any time, the adequacy of an existing bond or guaranty and may require an adjustment by either increasing or decreasing the amount required.

(d) A bond may be issued by a fidelity or surety company authorized to do business in the State of Florida or it may be a cash bond. The bond must initially provide for at least 24 months of coverage from the date of issuance and at no time thereafter shall the period of coverage be less than 12 months, for as long as the license remains in effect.

1. The bond must contain, without limitation, the following conditions:

a. The bond must be open-ended or, if written for a specific term, such as five years, must be renewed automatically unless 90 days or more prior to the renewal date, the issuer notifies the Department’s Bureau of Radiation Control, the beneficiary, and the licensee of its intention not to renew;

b. The bond must provide that the full face amount be paid to the beneficiary automatically prior to the expiration without proof of forfeiture if the licensee fails to provide replacement acceptable to the Department within 30 days after receipt of notification of cancellation; and,

c. The bond must be payable to the State of Florida;

2. The bond must remain in effect until the Department has terminated the license or the requirements of sub-subparagraph 64E-5.217(1)(d)1.b., F.A.C., are met.

(e) The Department may order the bond to be forfeited if it finds any of the following:

1. The facility or site has been abandoned;

2. The licensee is insolvent; or

3. The licensee is unable to perform to the satisfaction of the Department.

(f) Upon determining that a bond shall be forfeited, the Department shall issue a notice to that effect.

(2) The following are exempt from the provisions of this subpart:

(a) Other governmental agencies;

(b) Licensees who are only authorized for possession or use of radioisotopes with a half-life less than or equal to 120 days.

(c) Any licensee whose mathematical product of the risk factors in the Bond Risk Factors Calculation Worksheet ‒ March 2014, is less than or equal to 30,000.

(3) At the time of license application, license renewal and at intervals not to exceed 3 years, the applicant or licensee must submit the bonding determination described in subsection 64E-5.217(1), F.A.C., above, with adjustments as necessary to account for changes in chemical and physical form of radioactive material, radioisotopes authorized, their radiotoxicity and quantity, authorized licensed activities, actual costs and the potential costs of decontamination, treatment, or disposal of radioactive materials and contaminated equipment or facilities.

Rulemaking Authority 404.051, 404.061, 404.111, 404.141 FS. Law Implemented 404.051(1), (4), 404.061(2), 404.111, 404.141 FS. History–New 7-17-85, Amended 4-4-89, 5-12-93, Formerly 10D-91.322, Amended 7-1-15.

- 64E-5.218 Performance of Inspections.

(1) Radioactive material inspections may be announced or unannounced.

(2) Inspection procedures for all license categories will include the following:

(a) At the time of entrance to a facility, the department will inform the licensee management if available the purpose, extent, and approximate length of time required to complete the inspection;

(b) Consultation with workers in accordance with Rule 64E-5.905, F.A.C., may be performed;

(c) The department will review any or all records that are required to be maintained by these regulations or by license conditions;

(d) Radiation surveys will be performed to determine compliance with the regulations and license. The department’s radiation detection and monitoring equipment will be operable and calibrated as required by these regulations;

(e) Upon completion of an inspection, the department will inform the licensee of the preliminary findings of the inspection prior to leaving the facility, if possible. Official notification of the inspection findings will be sent in writing to the licensee.

(3) The department will perform inspections to assure the radioactive materials are used only as specified in these regulations or in the license using instruments calibrated as specified in these regulations.

Rulemaking Authority 404.022, 404.042, 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.042, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.324.

- 64E-5.219 Emergency Planning.

(1) Each application to possess radioactive materials in unsealed form, on foils or plated sources, or sealed in glass in excess of the quantities in Rule 64E-5.220, F.A.C., must contain either:

(a) An evaluation showing that the maximum dose to a person off-site due to a release of radioactive materials would not exceed 1 rem (10 mSv) effective dose equivalent or 5 rem (50 mSv) to the thyroid; or

(b) An emergency plan for responding to a release of radioactive material.

(2) One or more of the following factors can be used to support an evaluation submitted under paragraph (1)(a) of this section:

(a) The radioactive material is physically separated so that only a portion could be involved in an accident.

(b) All or part of the radioactive material is not subject to release during an accident because of the way it is stored or packaged.

(c) The release fraction in the respirable size range would be lower than the release fraction shown in Rule 64E-5.220, F.A.C., due to the chemical or physical form of the material.

(d) The solubility of the radioactive material would reduce the dose received.

(e) Facility design or engineered safety features in the facility would cause the release fraction to be lower than shown in Rule 64E-5.220, F.A.C.

(f) Operating restrictions or procedures would prevent a release fraction as large as that shown in Rule 64E-5.220, F.A.C.

(g) Other factors appropriate for the specific facility.

(3) Each application to possess source material in the form of uranium hexafluoride in excess of 50 kilograms in a single container or 1,000 kilograms total must contain either:

(a) An evaluation showing that the maximum intake of uranium by a member of the public due to a release would not exceed 2 milligrams; or

(b) An emergency plan for responding to the radiological hazards of an accidental release of source material and to any associated chemical hazards.

(4) One or more of the following factors can be used to support an evaluation submitted under paragraph (3)(a) of this section:

(a) All or part of the radioactive material is not subject to release during an accident because of the way it is stored or packaged.

(b) Facility design or engineered safety features in the facility would reduce the amount of the release.

(c) Other factors pertaining to the specific facility.

(5) Each application to possess special nuclear material in the form of uranium hexafluoride in excess of 50 kilograms in a single container or 1,000 kilograms total, or in excess of 2 curies (74 GBg) of plutonium in unsealed form or on foils or plated sources, must contain either:

(a) An evaluation showing that the maximum dose to a member of the public off-site due to a release of radioactive materials would not exceed 1 rem (10mSv) effective dose equivalent; or

(b) An emergency plan for responding to the radiological hazards of an accidental release of special nuclear material and to any associated chemical hazards.

(6) One or more of the following factors can be used to support an evaluation submitted under paragraph (5)(a) of this section:

(a) The radioactive material is physically separated so that only a portion could be involved in an accident.

(b) All or part of the radioactive material is not subject to release during an accident because of the way it is stored or packaged.

(c) In the case of fires or explosions, the release fraction would be lower than 0.001 due to the chemical or physical form of the material.

(d) The solubility of the material released would reduce the dose received.

(e) The facility design or engineered safety features in the facility would cause the release fraction to be lower than 0.001.

(f) Operating restrictions or procedures would prevent a release large enough to cause a member of the public off-site to receive a dose exceeding 1 rem (10mSv) effective dose equivalent.

(g) Other factors pertaining to the specific facility.

(7) An emergency plan responding to a release of radioactive material submitted under paragraph (1)(b), (3)(b) or (5)(b) of this section must include the following information:

(a) A brief description of the licensee’s facility and area near the site.

(b) An identification of each type of radioactive materials accident for which protective actions could be needed.

(c) A classification system for classifying accidents as alerts or site area emergencies.

(d) Identification of the means of detecting each type of accident in a timely manner.

(e) A brief description of the means and equipment for mitigating the consequences of each type of accident, including those provided to protect workers on site, and a description of the program for maintaining the equipment.

(f) A brief description of the methods and equipment to assess releases of radioactive materials.

(g) A brief description of the responsibilities of licensee personnel should an accident occur, including identification of personnel responsible for promptly notifying off-site response organizations and the department and the responsibilities of licensee personnel for developing, maintaining, and updating the plan.

(h) A commitment to and a brief description of the means to promptly notify off-site response organizations and request off-site assistance, including medical assistance for the treatment of contaminated injured on-site workers when appropriate. A control point must be established. The notification and coordination must be planned so that unavailability of some personnel, parts of the facility, or some equipment will not prevent the notification and coordination. The licensee shall also commit to notify the department immediately after notification of the appropriate off-site response organizations and not later than 1 hour after the licensee declares an emergency.

(i) A brief description of the types of information on facility status, radioactive releases, and recommended protective actions, if necessary, to be given to off-site response organizations and the department.

(j) A brief description of the frequency, performance objectives and plans for the training that the licensee will provide workers on how to respond to an emergency, including any special instructions and orientation tours the licensee would offer to fire, police, medical and other emergency personnel. The training shall familiarize personnel with site-specific emergency procedures. Also, the training shall thoroughly prepare site personnel for their responsibilities in the event of accident scenarios postulated as most probable for the specific site, including the use of team training for such scenarios.

(k) A brief description of the means of restoring the facility to a safe condition after an incident.

(l) Provisions for conducting quarterly communications checks with off-site response organizations and biennial on-site exercises to test response to simulated emergencies. Quarterly communications checks with off-site response organizations must include the check and update of all necessary telephone numbers. The licensee shall invite off-site response organizations to participate in the biennial exercises. Participation of off-site response organizations in biennial exercises, although recommended, is not required. Exercises must use accident scenarios postulated as most probable for the specific site and the scenarios shall not be known to most exercise participants. The licensee shall critique each exercise using individuals not having direct implementation responsibility for the plan. Critiques of exercises must evaluate the appropriateness of the plan, emergency procedures, facilities, equipment, training of personnel, and overall effectiveness of the response. Deficiencies found by the critiques must be corrected.

(m) A certification that the applicant has met its responsibilities under the Emergency Planning and Community Right-to-Know Act of 1986, Title III, Pub. L. 99-499, if applicable to the applicant’s activities at the proposed place of use of the radioactive material.

(8) The licensee shall allow the off-site response organizations expected to respond in case of an accident 60 days to comment on the licensee’s emergency plan before submitting it to the department. The licensee shall provide any comments received within the 60 days to the department with the emergency plan.

Rulemaking Authority 404.022, 404.042, 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.042, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.326.

- 64E-5.220 Radioactive Quantities.

(1) The list of quantities of radioactive materials requiring consideration of the need for an emergency plan for responding to a release as required in Rule 64E-5.219, F.A.C., is provided in the Department publication “Emergency Plan Isotopes and Quantities,” July 2014 edition, which is incorporated by reference and can be obtained at http://www.floridahealth.gov/radiation, and at https://www.flrules.org/Gateway/reference.asp?No=Ref-05439.

(2) For combinations of radioactive materials, consideration of the need for an emergency plan is required if the sum of the ratios of the quantity of each radioactive material authorized to the quantity listed for that material in this section exceeds one.

(3) Waste packaged in Type B containers as specified in Rule 64E-5.101, F.A.C., does not require an emergency plan.

Rulemaking Authority 404.022, 404.042, 404.051, 404.061, 404.071, 404.081 FS. Law Implemented 404.022, 404.042, 404.051(1), (4), (6), (9), (10), 404.061(2), (3), 404.071(1), 404.081(1) FS. History–New 5-12-93, Formerly 10D-91.327, Amended 6-3-15.

- 64E-5.221 Radiological Criteria for License Termination.

The criteria in this subpart apply to the decommissioning of facilities licensed under this chapter but do not apply to uranium and thorium recovery facilities as specified in Rule 64E-5.211, F.A.C., or to sites which previously have submitted and received department approval of a license termination plan or decommissioning plan as specified in subsection 64E-5.214(2), F.A.C.

(1) After a site has been decommissioned and the license terminated in accordance with the criteria in this subpart, the department will require additional cleanup only if based on new information or if it determines that the criteria of this subpart were not met and residual activity remaining at the site could result in significant threat to public health and safety.

(2) When calculating total effective dose equivalent to the average member of the critical group, the licensee shall determine the peak annual total effective dose equivalent expected within the first 1,000 years after decommissioning.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (4), (6), (9), 404.061(2), 404.081(1) FS. History–New 12-19-01.

- 64E-5.222 Radiological Criteria for Unrestricted Use.

A site is acceptable for unrestricted use if the total effective dose equivalent to an average member of the critical group from the residual radioactivity that is distinguishable from background radiation does not exceed 25 millirem (0.25 mSv) per year including radioactivity from groundwater sources of drinking water and the residual radioactivity levels are as low as reasonably achievable. Determination of the ALARA levels must take into account any detriments such as deaths from transportation accidents potentially expected to result from decontamination and waste disposal.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (4), (6), (9), 404.061(2), 404.081(1) FS. History–New 12-19-01.

- 64E-5.223 Criteria for License Termination Under Restricted Conditions.

A site is acceptable for license termination under restricted conditions if it meets the criteria below.

(1) The residual levels associated with restricted conditions are ALARA or the licensee can demonstrate that further reductions in residual radioactivity to comply with the provisions of Rule 64E-5.222, F.A.C., would result in an increase in public or environmental harm. Determination of the ALARA levels must take into account any detriments such as traffic accidents potentially expected to result from decontamination and waste disposal.

(2) The licensee has made provisions for legally enforceable institutional controls that provide reasonable assurance that the total effective dose equivalent from residual radioactivity distinguishable from background to the average member of the critical group will not exceed 25 millirem (0.25 mSv) per year.

(3) The licensee has provided sufficient financial assurance to enable an independent third party including a governmental custodian of a site to assume and carry out responsibilities for any necessary control and maintenance of the site. Acceptable financial assurance mechanisms are:

(a) Funds sufficient to pay decommissioning costs placed into an account segregated from the licensee’s assets and outside the licensee’s administrative control before the start of decommissioning operations; or

(b) A bond as specified in Rule 64E-5.217, F.A.C., or

(c) An arrangement deemed acceptable by the governmental entity that is assuming custody and ownership of a site.

(4) The licensee has submitted a decommissioning or license termination plan as specified in subsection 64E-5.214(2), F.A.C., to the department indicating the licensee's intent to decommission in accordance with this part and specifying that the licensee intends to decommission by restricting use of the site. The licensee shall document in the license termination or decommissioning plan how the advice of individuals and institutions in the community who could be affected by the decommissioning has been sought and incorporated, as appropriate, following analysis of that advice.

(a) Licensees proposing to decommission by restricting use of the site shall seek advice from such affected parties regarding the following matters:

1. Whether provisions for institutional controls proposed by the licensee:

a. Will provide reasonable assurance that the total effective dose equivalent from residual radioactivity distinguishable from background to the average member of the critical group will not exceed 25 millirem (0.25 mSv) per year;

b. Will be enforceable; and

c. Will not impose undue burdens on the local community or other affected parties.

2. Whether the licensee has provided sufficient financial assurance to enable an independent third party including a governmental custodian of a site to assume and carry out responsibilities for any necessary control and maintenance of the site.

(b) In seeking advice on the issues identified in paragraph (a), above, the licensee shall provide for:

1. Participation by representatives of a broad cross section of community interests who could be affected by the decommissioning;

2. An opportunity for a comprehensive, collective discussion on the issues by the participants represented; and

3. A publicly available summary of the results of all such discussions including a description of the individual viewpoints of the participants on the issues and the extent of agreement or disagreement among the participants on the issues.

(5) Residual radioactivity at the site has been reduced so that if the institutional controls were no longer in effect there is reasonable assurance that the total effective dose equivalent from residual radioactivity distinguishable from background to the average member of the critical group is as low as reasonably achievable and would not exceed 100 millirem (1 mSv) per year.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (2), (3), (4), (6), (9), 404.061(2), 404.081(1) FS. History–New 12-19-01.

- 64E-5.224 Alternate Criteria for License Termination.

The department will terminate a license using alternate criteria greater than the dose criterion of Rule 64E-5.222, subsection 64E-5.223(2), and sub-subparagraph 64E-5.223(4)(a)1.a., F.A.C., if the licensee:

(1) Provides assurance that public health and safety would continue to be protected and that it is unlikely that the total effective dose equivalent from all combined man-made sources other than medical sources would be more than 100 millirem per year (1 millisievert per year) by submitting an analysis of possible sources of exposure;

(2) Has employed restrictions to the extent practical on site use according to the provisions of Rule 64E-5.223, F.A.C., in minimizing exposures at the site;

(3) Reduces doses to ALARA levels considering any detriments such as traffic accidents potentially expected to result from decontamination and waste disposal; and

(4) Has submitted a decommissioning or license termination plan to the department indicating the licensee’s intent to decommission as specified in subsection 64E-5.214(2), F.A.C., and specifying that the licensee proposes to decommission by use of alternate criteria. The licensee shall document in the license termination or decommissioning plan how the advice of individuals and institutions in the community who could be affected by the decommissioning has been sought and addressed, as appropriate, following analysis of that advice. In seeking such advice, the licensee shall provide for:

(a) Participation by representatives of a broad cross section of community interests who could be affected by the decommissioning;

(b) An opportunity for a comprehensive, collective discussion on the issues by the participants represented; and

(c) A publicly available summary of the results of all such discussions, including a description of the individual viewpoints of the participants on the issues and the extent of agreement and disagreement on the issues.

(5) The use of alternate criteria to terminate a license requires the approval of the department after consideration of any comments provided by the U. S. Environmental Protection Agency and any public comments submitted as specified in Rule 64E-5.225, F.A.C.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (2), (4), (6), (9), 404.061(2), 404.081(1) FS. History–New 12-19-01.

- 64E-5.225 Public Notification and Public Participation.

Upon the receipt of a license termination or decommissioning plan or a proposal for release of a site as specified in Rule 64E-5.223 or 64E-5.224, F.A.C., and the total effective dose equivalent will exceed 50 millirem (0.5 mSv), the department shall:

(1) Notify and solicit comments from:

(a) Local and other state governments in the vicinity of the site and any Indian Nation or other indigenous people that could be affected by the decommissioning; and

(b) The U. S. Environmental Protection Agency if the licensee proposes to release a site as specified in Rule 64E-5.224, F.A.C.

(2) Publish a notice in the Florida Administrative Weekly to solicit comments from affected parties.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (2), (4), (6), (9), 404.061(2), 404.081(1) FS. History–New 12-19-01.

- 64E-5.226 Minimizing Contamination.

After the effective date of this rule, applicants for licenses other than renewals shall describe in the application how facility design and procedures for operation will minimize contamination of the facility and the environment to the extent practical, facilitate eventual decommissioning, and minimize the generation of radioactive waste to the extent practical.

Rulemaking Authority 404.051(4), (6), (9), 404.061(2), 404.081 FS. Law Implemented 404.051(1), (4), (6), (9), 404.061(2), 404.081(1) FS. History– New 12-19-01.

- 64E-5.301 Standards for Protection Against Radiation.

(1) The rules in this part control the receipt, possession, use, disposal, and transfer of sources of radiation by any licensee or registrant so the total dose to an individual, including doses resulting from all sources of radiation other than background radiation, does not exceed the standards for protection against radiation prescribed in this part. However, nothing in this part shall be construed to limit actions necessary to protect health and safety.

(2) Except as specifically provided in other parts of these rules, this part applies to persons licensed or registered by the department to receive, possess, use, or transfer sources of radiation. The limits in this part do not apply to doses from background radiation, to exposure of patients to radiation for medical diagnosis or therapy, to exposure from individuals administered radioactive material and released as specified in Rule 64E-5.622, F.A.C., or to voluntary participation in medical research programs.

Rulemaking Authority 404.051(1) FS. Law Implemented 404.022, 404.051(1), (4), 404.181(1)(b) FS. History–New 1-1-94, Amended 5-15-96, Formerly 10D-91.431, Amended 10-8-00.

- 64E-5.302 Implementation.

(1) Any existing license or registration condition that is more restrictive than Part III remains in force until there is an amendment or renewal of the license or registration.

(2) If a license or registration condition exempts a licensee or registrant from a provision of the part in effect on or before the effective date of this rule, it also exempts the licensee or registrant from the corresponding provisions of this part.

(3) If a license or registration condition cites provisions of this part in effect prior to the effective date of this rule which do not correspond to any provisions of this part, the license or registration condition remains in force until there is an amendment or renewal of the license or registration that modifies or removes this condition.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.433.

- 64E-5.303 Radiation Protection Programs.

(1) Each licensee or registrant shall develop, document, and implement a radiation protection program sufficient to ensure compliance with the provisions of this part. See Rule 64E-5.335, F.A.C., for recordkeeping requirements relating to these programs.

(2) The licensee or registrant shall use to the extent practical procedures and engineering controls based upon sound radiation protection principles to achieve occupational doses and public doses that are as low as reasonably achievable.

(3) The licensee or registrant shall, at intervals not to exceed 12 months, review the radiation protection program content and implementation.

(4) Dental and podiatry registrants are exempt from subsections (1) and (3), above.

(5) To implement the ALARA requirements of subsection 64E-5.303(2), F.A.C., and notwithstanding the requirements of Rule 64E-5.312, F.A.C., of this part, licensees shall establish constraints on air emissions of radioactive material, excluding radon 222 and its daughters, to the environment so that individual members of the public who are likely to receive the highest doses are not expected to receive a total effective dose equivalent in excess of 10 millirems (0.10 mSv) per year from these emissions. If a licensee subject to this requirement exceeds this dose constraint, the licensee shall report the occurrence as specified in Rule 64E-5.345, F.A.C., and promptly take corrective action to ensure against recurrence.

Rulemaking Authority 404.051(4), 404.081(1) FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Amended 11-20-94, Formerly 10D-91.434, Amended 10-8-00.

- 64E-5.304 Occupational Dose Limits for Adults.

(1) The licensee or registrant shall control the occupational dose to individual adults, except for planned special exposures as specified in Rule 64E-5.309, F.A.C., to the following dose limits:

(a) An annual limit, which is the more limiting of:

1. The total effective dose equivalent equal to 5 rem (0.05 sievert); or

2. The sum of the deep dose equivalent and the committed dose equivalent to any individual organ or tissue other than the lens of the eye equal to 50 rem (0.5 sievert).

(b) The annual limits to the lens of the eye, to the skin of the whole body, and to the skin of the extremities which are:

1. A lens dose equivalent of 15 rem (0.15 sievert), and

2. A shallow dose equivalent of 50 rem (0.5 sievert) to the skin of the whole body or to skin of any extremity.

(2) Doses received in excess of the annual limits, including doses received during accidents, emergencies, and planned special exposures, shall be subtracted from the limits for planned special exposures that the individual could receive during the current year and during the individual’s lifetime as specified in paragraphs 64E-5.309(5)(a) and (b), F.A.C.

(3) When the external exposure is determined by measurement with an external personal monitoring device, the deep-dose equivalent must be used in place of the effective dose equivalent, unless the effective dose equivalent is determined by a dosimetry method approved by the Department. The assigned deep dose equivalent must be for the part of the body receiving the highest exposure. The assigned shallow dose equivalent must be the dose averaged over the contiguous 10 square centimeters of skin receiving the highest exposure. The deep dose equivalent, lens dose equivalent and shallow dose equivalent may be assessed from surveys or other radiation measurements to demonstrate compliance with the occupational dose limits if the individual monitoring device was not in the region of highest potential exposure or the results of individual monitoring are unavailable.

(4) Derived air concentration (DAC) and annual limit on intake (ALI) values are presented in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, Table I, and can be used to determine the individual’s dose and to demonstrate compliance with the occupational dose limits. See Rule 64E-5.339, F.A.C.

(5) In addition to the annual dose limits, the licensee shall limit the soluble uranium intake by an individual to 10 milligrams in a week in consideration of chemical toxicity. See footnote 3 of State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012 (see 64E-5.101, F.A.C.).

(6) The licensee or registrant shall reduce the dose that an individual can be allowed to receive in the current year by the amount of occupational dose received while employed by any other person. See subsection 64E-5.308(5), F.A.C.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Formerly 10D-91.435, Amended 10-8-00, 9-28-06, 12-26-13.

- 64E-5.305 Compliance with Requirements for Summation of External and Internal Doses.

(1) If the licensee is required to monitor as specified in both subsections 64E-5.315(1) and (2), F.A.C., the licensee shall demonstrate compliance with the dose limits by summing external and internal doses. If the licensee or registrant is required to monitor only as specified in subsection 64E-5.315(1), F.A.C., or only as specified in subsection 64E-5.315(2), F.A.C., then summation is not required to demonstrate compliance with the dose limits. The licensee can demonstrate compliance with the requirements for summation of external and internal doses as specified in subsections 64E-5.305(2), (3) and (4), F.A.C. The dose equivalents for the lens of the eye, the skin, and the extremities are not included in the summation, but are subject to separate limits.

(2) Intake by Inhalation. If the only intake of radionuclides is by inhalation, the total effective dose equivalent limit is not exceeded if the sum of the deep dose equivalent divided by the total effective dose equivalent limit and one of the following does not exceed unity:

(a) The sum of the fractions of the inhalation ALI for each radionuclide;

(b) The total number of derived air concentration-hours (DAC-hours) for all radionuclides divided by 2,000; or

(c) The sum of the calculated committed effective dose equivalents to all significantly irradiated organs or tissues (T) calculated from bioassay data using appropriate biological models and expressed as a fraction of the annual limit. For purposes of this requirement, an organ or tissue is considered significantly irradiated if, for that organ or tissue, the product of the weighting factors, WT, and the committed dose equivalent, HT, 50, per unit intake is greater than 10 percent of the maximum weighted value of H50, or WT HT,50, per unit intake for any organ or tissue.

(3) Intake by Oral Ingestion. If the occupationally exposed individual also receives an intake of radionuclides by oral ingestion greater than 10 percent of the applicable oral ALI, the licensee shall account for this intake and include it in demonstrating compliance with the limits.

(4) Intake through Wounds or Absorption through Skin. The licensee shall evaluate and to the extent practical account for intakes through wounds or skin absorption. The intake through intact skin has been included in the calculation of DAC for hydrogen 3 and does not need to be evaluated or accounted for as specified in this subsection.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.436.

- 64E-5.306 Determination of External Dose from Airborne Radioactive Material.

(1) Licensees shall include the contribution to the deep dose equivalent, eye dose equivalent, and shallow dose equivalent from external exposure to the radioactive cloud when determining the dose from airborne radioactive material. See State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) footnotes 1 and 2.

(2) Airborne radioactivity measurements and DAC values shall not be used as the primary means to assess the deep dose equivalent when the airborne radioactive material includes radionuclides other than noble gases or if the cloud of airborne radioactive material is not relatively uniform. The determination of the deep dose equivalent to an individual shall be based upon measurements using instruments or individual monitoring devices.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Formerly 10D-91.437, Amended 12-26-13.

- 64E-5.307 Determination of Internal Exposure.

(1) To assess dose used to determine compliance with occupational dose equivalent limits when required as specified in Rule 64E-5.315, F.A.C., the licensee shall take suitable and timely measurements of:

(a) Concentrations of radioactive materials in air in work areas;

(b) Quantities of radionuclides in the body;

(c) Quantities of radionuclides excreted from the body; or

(d) Combinations of these measurements.

(2) Unless respiratory protective equipment is used as specified in Rule 64E-5.319, F.A.C., or the assessment of intake is based on bioassays, the licensee shall assume that an individual inhales radioactive material at the airborne concentration in which the individual is present.

(3) When specific information on the physical and biochemical properties of the radionuclides taken into the body or the behavior of the material in an individual is known, the licensee is permitted to:

(a) Use that information to calculate the committed effective dose equivalent, and, if used, the licensee shall document that information in the individual’s record;

(b) Upon prior approval of the department, adjust the DAC or ALI values to reflect the actual physical and chemical characteristics of airborne radioactive material, for example, aerosol size distribution or density; and

(c) Separately assess the contribution of fractional intakes of Class D, W, or Y compounds of a given radionuclide to the committed effective dose equivalent. See State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012.

(4) If the licensee chooses to assess intakes of Class Y material using the measurements given in paragraph 64E-5.307(1)(b) or (c), F.A.C., the licensee can delay the recording and reporting of the assessments for periods up to 7 months, unless otherwise required by Rule 64E-5.344 or 64E-5.345, F.A.C. This delay permits the licensee to make additional measurements basic to the assessments.

(5) If the identity and concentration of each radionuclide in a mixture are known, the fraction of the DAC applicable to the mixture for use in calculating DAC-hours shall be either:

(a) The sum of the ratios of the concentration to the appropriate DAC value, that is D, W, or Y, from State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, for each radionuclide in the mixture; or

(b) The ratio of the total concentration for all radionuclides in the mixture to the most restrictive DAC value for any radionuclide in the mixture.

(6) If the identity of each radionuclide in a mixture is known, but the concentration of one or more of the radionuclides in the mixture is not known, the DAC for the mixture shall be the most restrictive DAC of any radionuclide in the mixture.

(7) When a mixture of radionuclides in air exists, a licensee is permitted to disregard certain radionuclides in the mixture if:

(a) The licensee uses the total activity of the mixture in demonstrating compliance with the dose limits in Rule 64E-5.304, F.A.C., and in complying with the monitoring requirements in subsection 64E-5.315(2), F.A.C.;

(b) The concentration of any radionuclide disregarded is less than 10 percent of its DAC; and

(c) The sum of these percentages for all of the radionuclides disregarded in the mixture does not exceed 30 percent.

(8) When determining the committed effective dose equivalent, the following information can be considered:

(a) To calculate the committed effective dose equivalent, the licensee can assume that the inhalation of one ALI or an exposure of 2,000 DAC-hours results in a committed effective dose equivalent of 5 rem (0.05 sievert) for radionuclides that have their ALIs or DACs based on the committed effective dose equivalent.

(b) For an ALI and the associated DAC determined by the nonstochastic organ dose limit of 50 rem (0.5 sievert), the intake of radionuclides that would result in a committed effective dose equivalent of 5 rem (0.05 sievert), that is, the stochastic ALI, as listed in parentheses in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table I. The licensee can use the stochastic ALI to determine committed effective dose equivalent as a simplifying assumption. However, if the licensee uses the stochastic ALI, the licensee shall also demonstrate that the limit in subparagraph 64E-5.304(1)(a)2., F.A.C., is met.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Formerly 10D-91.438, Amended 12-26-13.

- 64E-5.308 Determination of Prior Occupational Dose.

(1) For each individual who is likely to receive in a year an occupational dose requiring monitoring as specified in Rule 64E-5.315, F.A.C., the licensee or registrant shall:

(a) Determine the occupational radiation dose received during the current year; and

(b) Attempt to obtain the records of lifetime cumulative occupational radiation dose.

(2) Prior to permitting an individual to participate in a planned special exposure, the licensee or registrant shall determine:

(a) The internal and external doses from all previous planned special exposures;

(b) All doses in excess of the limits received during the lifetime of the individual, including doses received during accidents and emergencies; and

(c) All lifetime cumulative occupational radiation doses.

(3) In complying with the requirements of subsection 64E-5.308(1), F.A.C., a licensee or registrant can:

(a) Accept a written signed statement from the individual or from the individual’s most recent employer for work involving radiation exposure that discloses the nature and the amount of any occupational dose that the individual received during the current year as a record of the occupational dose that the individual received during the current year;

(b) Accept an up-to-date DOH Form 1623 July 1993, which is herein incorporated by reference and which is available from the department, or an equivalent signed by the individual and countersigned by an appropriate official of the most recent employer for work involving radiation exposure or the individual’s current employer if the individual is not employed by the licensee or registrant as the record of lifetime cumulative radiation dose; and

(c) Obtain reports of the individual’s dose equivalent from the most recent employer for work involving radiation exposure or the individual’s current employer if the individual is not employed by the licensee or registrant by telephone, telegram, facsimile, or letter. The licensee or registrant shall request a written verification of the dose data if the authenticity of the transmitted report cannot be established.

(4) The licensee or registrant shall record the exposure history, as required by subsection 64E-5.308(1), F.A.C., on DOH Form 1623 July 1993 or other clear and legible record of all the information required on that form. The form or record shall show each period in which the individual received occupational exposure to radiation or radioactive material and shall be signed by the individual who received the exposure. For each period for which the licensee or registrant obtains reports, the licensee or registrant shall use the dose shown in the report in preparing DOH Form 1623 July 1993 or equivalent. For any period in which the licensee or registrant does not obtain a report, the licensee or registrant shall place a notation on DOH Form 1623 July 1993 or equivalent indicating the periods of time for which data are not available.

(5) Licensees are not required to reevaluate the separate external dose equivalents and internal committed dose equivalents or intakes of radionuclides assessed as specified in the rules in this part in effect before the effective date of this rule. Further, occupational exposure histories obtained and recorded on DOH Form 1623 or equivalent before the effective date of this rule would not have included effective dose equivalents, but can be used in the absence of specific information on the intake of radionuclides by the individual.

(6) If the licensee or registrant is unable to obtain a complete record of an individual’s current and previously accumulated occupational dose, the licensee or registrant shall assume:

(a) That the allowable dose limit for the individual is reduced by 1.25 rem (12.5 millisievert) for each quarter for which records were unavailable and the individual was engaged in activities that could have resulted in occupational radiation exposure to establish administrative controls as specified in subsection 64E-5.304(6), F.A.C., for the current year; and

(b) That the individual is not available for planned special exposures.

(7) The licensee or registrant shall retain the records on DOH Form 1623 July 1993 or equivalent until the department terminates each pertinent license or registration requiring this record. The licensee or registrant shall retain records used in preparing DOH Form 1623 July 1993 or equivalent for 3 years after the record is made.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.439.

- 64E-5.309 Planned Special Exposures.

A licensee or registrant can authorize an adult worker to receive doses in addition to and accounted for separately from the doses received under the limits specified in Rule 64E-5.304, F.A.C., if each of the following conditions is satisfied:

(1) The licensee or registrant authorizes a planned special exposure only in an exceptional situation when alternatives that might avoid the dose are unavailable or impractical.

(2) The licensee or registrant and employer if the employer is not the licensee or registrant specifically authorizes the planned special exposure in writing before the exposure occurs.

(3) Before a planned special exposure, the licensee or registrant ensures that each individual involved is:

(a) Informed of the purpose of the planned operation;

(b) Informed of the estimated doses and associated potential risks and specific radiation levels or other conditions that might be involved in performing the task; and

(c) Instructed in the measures to be taken to keep the dose ALARA considering other risks that might be present.

(4) The licensee or registrant ascertains prior doses as required by subsection 64E-5.308(2), F.A.C., during the lifetime of each individual involved prior to permitting an individual to participate in a planned special exposure.

(5) As specified in subsection 64E-5.304(2), F.A.C., the licensee or registrant shall not authorize a planned special exposure that would cause an individual to receive a dose from all planned special exposures and all doses in excess of:

(a) The numerical values of any of the dose limits in subsection 64E-5.304(1), F.A.C., in any year; and

(b) Five times the annual dose limits in subsection 64E-5.304(1), F.A.C., during the individual’s lifetime.

(6) The licensee or registrant maintains records of the conduct of a planned special exposure as specified in Rule 64E-5.338, F.A.C., and submits a written report to the department as specified in Rule 64E-5.346, F.A.C.

(7) The licensee or registrant records the best estimate of the dose resulting from the planned special exposure in the individual’s record and informs the individual in writing of the dose within 30 days after the date of the planned special exposure. The dose from planned special exposures shall not be considered in controlling future occupational dose of the individual as specified in subsection 64E-5.304(1), F.A.C., but shall be included in evaluations required by subsections 64E-5.309(4) and (5), F.A.C.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.440, Amended 10-8-00.

- 64E-5.310 Occupation Dose Limits for Minors.

The annual occupational dose limits for minors are 10 percent of the annual occupational dose limits specified for adult workers in Rule 64E-5.304, F.A.C.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.441.

- 64E-5.311 Dose to an Embryo or Fetus.

(1) The licensee or registrant shall ensure that the dose to an embryo or fetus during the entire pregnancy from occupational exposure of a declared pregnant woman does not exceed 0.5 rem (5 mSv). See Rule 64E-5.339, F.A.C., for recordkeeping requirements.

(2) The licensee or registrant shall make efforts to avoid substantial variation above a uniform monthly exposure rate to a declared pregnant woman so as to satisfy the limit in subsection 64E-5.311(1), F.A.C. It is recommended that no more than 0.05 rem (0.5 mSv) be received by the embryo or fetus in any one month.

(3) The dose to an embryo or fetus shall be taken as the sum of:

(a) The deep dose equivalent to the declared pregnant woman; and

(b) The dose to the embryo or fetus from radionuclides in the embryo or fetus and radionuclides in the declared pregnant woman.

(4) If by the time the woman declares pregnancy to the licensee or registrant the dose to the embryo or fetus has exceeded 0.5 rem (5 mSv) or is within 0.05 rem (0.5 mSv) of this dose, the licensee or registrant shall be considered in compliance with subsection 64E-5.311(1), F.A.C., if the additional dose to the embryo or fetus does not exceed 0.05 rem (0.50 mSv) during the remainder of the pregnancy.

(5) Each individual who has declared pregnancy shall wear a radiation monitor at waist level at all times at work. This monitor shall be used to estimate the fetal deep-dose equivalent. When the declared pregnant worker wears protective clothing, this monitor shall be worn under the protective clothing. The fetal dose is to be kept as low as reasonably achievable, but shall not exceed 10 percent of the standards specified in Rule 64E-5.304, F.A.C. Each declared pregnant worker whose duties require protective clothing shall also wear a radiation monitor outside the protective clothing to estimate dose to the worker, and the standard occupational limits will apply.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.442, Amended 10-8-00.

- 64E-5.312 Dose Limits for Individual Members of the Public.

(1) Each licensee or registrant shall conduct operations so that:

(a) Except as specified in paragraph 64E-5.312(1)(b), F.A.C., the total effective dose equivalent to individual members of the public from the licensed or registered operation does not exceed 0.1 rem (1 millisievert) in a year, exclusive of the dose contribution from background radiation, from any medical administration the individual has received, from exposure to individuals administered radioactive materials and released as specified in Rule 64E-5.622, F.A.C., from voluntary participation in medical research programs and from the licensee’s disposal of radioactive material into sanitary sewerage as specified in Rule 64E-5.330, F.A.C.;

(b) In facilities in operation before January 1, 1994, the total effective dose equivalent to individual members of the public from infrequent exposure to radiation from diagnostic and therapeutic radiation machines does not exceed 0.5 rem (5 millisievert) in a year; and

(c) The dose in any unrestricted area from external sources, exclusive of the dose contribution from patients administered radioactive material and released as specified in Rule 64E-5.622, F.A.C., does not exceed 0.002 rem (0.02 millisievert) in any one hour.

(2) If the licensee or registrant permits members of the public to have access to restricted areas, the limits for members of the public continue to apply to those individuals.

(3) A licensee, registrant, or an applicant for a license or registration can apply for prior department authorization to operate up to an annual dose limit for an individual member of the public of 0.5 rem (5 millisievert). This application shall include the following information:

(a) Demonstration of the need for and the expected duration of operations in excess of the limit in subsection 64E-5.304(1), F.A.C.;

(b) The licensee’s or registrant’s program to assess and control dose within the 0.5 rem (5 millisievert) annual limit; and

(c) The procedures to be followed to maintain the dose ALARA.

(4) In addition to the requirements of this part, a licensee subject to the provisions of the U.S. Environmental Protection Agency’s generally applicable environmental radiation standards in 40 C.F.R. 190 shall comply with those standards.

(5) A licensee or applicant for a license may permit visitors to an individual who cannot be released under Rule 64E-5.622, F.A.C., to receive a radiation dose greater than 0.1 rem (1 millisievert) provided the following are satisfied:

(a) The radiation dose received does not exceed 0.5 rem (5 millisievert);

(b) The authorized user, as defined in Rule 64E-5.6011, F.A.C., has determined before the visit that it is appropriate.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Amended 5-15-96, Formerly 10D-91.443, Amended 10-8-00, 2-11-10.

- 64E-5.313 Compliance with Dose Limits for Individual Members of the Public.

(1) The licensee or registrant shall make or cause to be made surveys of radiation levels in unrestricted areas and radioactive materials in effluents released to unrestricted areas to demonstrate compliance with the dose limits for individual members of the public in Rule 64E-5.312, F.A.C.

(2) A licensee or registrant shall show compliance with the annual dose limit in Rule 64E-5.312, F.A.C., by:

(a) Demonstrating by measurement or calculation that the total effective dose equivalent to the individual who is likely to receive the highest dose from the licensed or registered operation does not exceed the annual dose limit; or

(b) Demonstrating that:

1. The annual average concentrations of radioactive material released in gaseous and liquid effluents at the boundary of the unrestricted area do not exceed the values specified in State of Florida Bureau of Radiation Control ALIs, DACs and Effluent Concentrations, June 2012, Table II; and

2. The dose from external sources would not exceed 0.002 rem (0.02 millisievert) in an hour and 0.05 rem (0.5 millisievert) in a year if an individual were continually present in an unrestricted area.

(3) Upon approval from the Department, the licensee can adjust the effluent concentration values in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) for members of the public to take into account the actual physical and chemical characteristics of the effluents, such as aerosol size distribution, solubility, density, radioactive decay equilibrium, and chemical form.

(4) Dental and podiatry registrants are exempt from subsections (1), (2), and (3), above.

(5) Each licensee or registrant shall maintain records sufficient to demonstrate compliance with the dose limit for individual members of the public until the department terminates each pertinent license or registration requiring the record.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Amended 11-20-94, 5-15-96, Formerly 10D-91.444, Amended 12-26-13.

- 64E-5.314 General.

(1) Each licensee or registrant shall make or cause to be made surveys that:

(a) Are necessary for the licensee or registrant to comply with this part; and

(b) Are necessary under the circumstances to evaluate:

1. The magnitude and extent of radiation levels;

2. Concentrations or quantities of radioactive material; and

3. The potential radiological hazards.

(2) The licensee or registrant shall ensure that instruments and equipment used for quantitative radiation measurements such as dose rate and effluent monitoring are calibrated annually for the radiation measured.

(3) All personnel dosimeters except for direct and indirect reading pocket ionization chambers and those dosimeters used to measure the dose to any extremity that require processing to determine the radiation dose and that are used by licensees and registrants to comply with Rule 64E-5.304, F.A.C., with other applicable provisions of these rules, or with conditions specified in a license or registration shall be processed and evaluated by a dosimetry processor:

(a) Holding current personnel dosimetry accreditation from the National Voluntary Laboratory Accreditation Program (NVLAP) of the National Institute of Standards and Technology; and

(b) Approved in this accreditation process for the type of radiation or radiations included in the NVLAP program that most closely approximates the type of radiation or radiations for which the individual wearing the dosimeter is monitored.

(4) The licensee or registrant shall ensure that adequate precautions are taken to prevent a deceptive exposure of an individual monitoring device.

(5) Dental and podiatry registrants are exempt from subsections (1) and (2), above.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Amended 11-20-94, Formerly 10D-91.445, Amended 10-8-00.

- 64E-5.315 Conditions Requiring Individual Monitoring of External and Internal Occupational Dose.

Each licensee or registrant shall monitor exposures from sources of radiation at levels sufficient to demonstrate compliance with the occupational dose limits of this part. As a minimum:

(1) Each licensee or registrant shall monitor occupational exposure to radiation and shall supply and require the use of individual monitoring devices by:

(a) Adults likely to receive in 1 year from sources external to the body a dose in excess of 10 percent of the limits in subsection 64E-5.304(1), F.A.C.;

(b) Minors likely to receive in 1 year from radiation sources external to the body a deep dose equivalent in excess of 0.1 rem (1 mSv), a lens dose equivalent in excess of 0.15 rem (1.5 mSv) or a shallow dose equivalent to the skin or to the extremities in excess of 0.5 rem (5 mSv);

(c) Declared pregnant women likely to receive during the entire pregnancy from radiation sources external to the body a deep dose equivalent in excess of 0.1 rem (1 mSv); and

(d) Individuals entering a high or very high radiation area.

(2) Each licensee shall monitor to determine compliance with Rule 64E-5.307, F.A.C., the occupational intake of radioactive material by and assess the committed effective dose equivalent to:

(a) Adults likely to receive in 1 year an intake in excess of 10 percent of the applicable ALI in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations June 2012, (see 64E-5.101, F.A.C.) Table I, Columns 1 and 2; and

(b) Minors likely to receive in 1 year a committed effective dose equivalent in excess of 0.1 rem (1.0 millisievert); and

(c) Declared pregnant women likely to receive during the entire pregnancy a committed effective dose equivalent in excess of 0.1 rem (1 mSv).

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Formerly 10D-91.446, Amended 10-8-00, 12-26-13.

- 64E-5.316 Control of Access to High Radiation Areas.

(1) The licensee or registrant shall ensure that each entrance or access point to a high radiation area has one or more of the following features:

(a) A control device that upon entry into the area causes the level of radiation to be reduced below that level at which an individual might receive a deep dose equivalent of 0.1 rem (1 millisievert) in 1 hour at 30 centimeters from the source of radiation from any surface that the radiation penetrates;

(b) A control device that energizes a conspicuous visible or audible signal so that the individual entering the high radiation area and the supervisor of the activity are made aware of the entry; or

(c) Entryways that are locked except during periods when access to the areas is required with positive control over each individual entry.

(2) The licensee or registrant can substitute continuous direct or electronic surveillance that is capable of preventing unauthorized entry in place of the controls required by subsection 64E-5.316(1), F.A.C., for a high radiation area.

(3) The licensee or registrant can apply to the department for approval of alternative methods for controlling access to high radiation areas.

(4) The licensee or registrant shall establish the controls required by subsections 64E-5.316(1) and (3), F.A.C., in a way that does not prevent individuals from leaving a high radiation area.

(5) The licensee is not required to control each entrance or access point to a room or other area that is a high radiation area solely because of the presence of radioactive materials prepared for transport and packaged and labeled as specified in the regulations of the U.S. Department of Transportation if:

(a) The packages do not remain in the area longer than 3 days; and

(b) The dose rate at 1 meter from the external surface of any package does not exceed 0.01 rem (0.1 millisievert) per hour.

(6) The licensee is not required to control entrance or access to rooms or other areas in hospitals solely because of the presence of patients containing radioactive material if there are personnel in attendance who are taking the necessary precautions to prevent the exposure of individuals to radiation or radioactive material in excess of the established limits in this part and to operate within the ALARA provisions of the licensee’s radiation protection program.

(7) The registrant is not required to control entrance or access to rooms or other areas containing sources of radiation capable of producing a high radiation area as described in Rule 64E-5.316, F.A.C., if the registrant has met all the specific requirements for access and control specified in other applicable parts of these rules, such as Part IV for industrial radiographic operations, Part V for x-rays in the healing arts, and Part VIII for particle accelerators.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.447.

- 64E-5.317 Control of Access to Very High Radiation Areas.

(1) In addition to the requirements in Rule 64E-5.316, F.A.C., the licensee or registrant shall institute measures to ensure that an individual is not able to gain unauthorized or inadvertent access to areas in which radiation levels could be encountered at 500 rad (5 gray) or more in 1 hour at 1 meter from a source of radiation or any surface through which the radiation penetrates. This requirement does not apply to rooms or areas in which diagnostic x-ray systems are the only source of radiation or to non-self-shielded irradiators.

(2) The registrant is not required to control entrance or access to rooms or other areas containing sources of radiation capable of producing a very high radiation area as described in subsection 64E-5.317(1), F.A.C., if the registrant has met all the specific requirements for access and control specified in other applicable parts of these rules, such as Part IV for industrial radiographic operations, Part V for x-rays in the healing arts, and Part VIII for particle accelerators.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.448.

- 64E-5.318 Use of Process or Other Engineering Controls.

The licensee shall use to the extent practical process or other engineering controls such as containment, decontamination, or ventilation to control the concentrations of radioactive material in air.

(1) When it is not practical to apply process or other engineering controls to control the concentrations of radioactive material in the air to values below those that define an airborne radioactivity area, the licensee shall, consistent with maintaining the total effective dose equivalent ALARA, increase monitoring and limit intakes by one or more of the following means:

(a) Control of access;

(b) Limitation of exposure time;

(c) Use of respiratory protection equipment; or

(d) Other controls.

(2) If the licensee performs an ALARA analysis to determine whether or not to use respirators, the licensee can consider safety factors other than radiological factors. The licensee also should consider the impact of respirator use on workers’ industrial health and safety.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Amended 5-15-96, Formerly 10D-91.450, Amended 9-28-06.

- 64E-5.319 Use of Individual Respiratory Protection Equipment.

(1) If the licensee uses respiratory protection equipment to limit intakes as specified in Rule 64E-5.318, F.A.C.:

(a) The licensee shall use only respiratory protection equipment that is tested and certified by the National Institute for Occupational Safety and Health except as provided in paragraph 64E-5.319(1)(b), F.A.C.

(b) If the licensee wishes to use equipment that has not been tested or certified by the National Institute for Occupational Safety and Health or for which there is no schedule for testing or certification, the licensee shall submit an application for authorized use of that equipment, including evidence that the material and performance characteristics of the equipment are capable of providing the proposed degree of protection under anticipated conditions of use demonstrated by testing or on the basis of reliable test information.

(c) The licensee shall implement and maintain a respiratory protection program that includes:

1. Air sampling sufficient to identify the potential hazard, permit proper equipment selection, and estimate doses;

2. Surveys and bioassays as needed to evaluate actual intakes;

3. Testing of respirators for operability including user seal checks for face sealing devices and functional checks for other devices immediately before each use;

4. Written procedures regarding supervision and training of respirator users; monitoring, including air sampling and bioassays; fit testing; respirator selection; breathing air quality; inventory and control; storage, issuance, maintenance, repair, testing, and quality assurance of respiratory protection equipment; limitations on periods of respirator use and relief from respirator use; and recordkeeping;

5. Determination by a physician before initial fitting of face sealing respirators, before the first field use of non-face sealing respirators, and every 12 months thereafter or periodically at a frequency determined by a physician that the individual user is medically fit to use respiratory protection equipment; and

6. Fit testing before the first field use of tight fitting face-sealing respirators and periodically thereafter at a frequency not to exceed 1 year with fit factor ≥ 10 times the APF for negative pressure devices and a fit factor ≥ 500 for any positive pressure, continuous flow, and pressure-demand devices. Fit testing shall be performed with the facepiece operating in the negative pressure mode.

(d) The licensee shall advise each respirator user that the user can leave the area at any time for relief from respirator use in the event of equipment malfunction, physical or psychological distress, procedural or communication failure, significant deterioration of operating conditions, or any other conditions that might require such relief.

(e) The licensee also shall consider limitations appropriate to the type and mode of use. When selecting respiratory devices, the licensee shall provide for vision correction, adequate communication, low temperature work environments, and the concurrent use of other safety or radiological protection equipment. The licensee shall use equipment in such a way as not to interfere with the proper operation of the respirator.

(f) Standby rescue persons are required whenever one-piece atmosphere-supplying suits or any combination of supplied-air respiratory protection device and personnel protective equipment are used from which an unaided individual would have difficulty extricating himself or herself. Standby persons shall be equipped with respiratory protection devices or other apparatus appropriate for the potential hazards. The standby rescue persons shall observe or otherwise maintain continuous communication with the workers through visual, voice, signal line, telephone, radio, or other suitable means and be available immediately to assist them in case of a failure of the air supply or for any other reason that requires relief from distress. A sufficient number of standby rescue persons shall be available immediately to assist all users of this type of equipment and to provide effective emergency rescue if needed.

(g) Atmosphere-supplying respirators shall be supplied with respirable air of grade D quality or better as defined by the Compressed Gas Association in publication G-7.1, “Commodity Specification for Air,” 1997, which is herein incorporated by reference and, available from the Compressed Gas Association, Inc., and included in the regulations of the Occupational Safety and Health Administration. Grade D quality air criteria include:

1. Oxygen content (v/v) of 19.5—23.5%;

2. Condensed hydrocarbon content of 5 milligrams per cubic meter of air or less;

3. Carbon monoxide content of 10 ppm or less;

4. Carbon dioxide content of 1,000 ppm or less; and

5. Lack of noticeable odor.

(h) The licensee shall ensure that no objects, materials, or substances such as facial hair, or any conditions that interfere with the face-to-facepiece seal or valve function and that are under the control of the respirator wearer are between the skin of the wearer’s face and the sealing surface of a tight-fitting respirator facepiece.

(2) When estimating the dose to individuals from airborne radioactive materials, the concentration of radioactive material in the air that is inhaled when respirators are worn is assumed initially to be the ambient concentration in air without respiratory protection divided by the assigned protection factor. If the dose later is found to be greater than the estimated dose, the corrected value shall be used. If the dose later is found to be less than the estimated dose, the corrected value can be used.

(a) Licensees shall take actions to limit doses to individuals from intakes of airborne radioactive materials to maintain total effective dose equivalent ALARA, which could include using process or other engineering controls and limiting the use of respiratory protection equipment.

(b) The licensee shall obtain authorization from the Department before using assigned protection factors in excess of those specified in State of Florida, Bureau of Radiation Control Protection Factors for Respirators, May 2006. The Department can authorize a licensee to use higher protection factors on receipt of an application that:

1. Describes the situation for which a need exists for higher protection factors; and

2. Demonstrates that the respiratory protection equipment provides these higher protection factors under the proposed conditions of use.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.452, Amended 5-18-98, 9-28-06.

- 64E-5.320 Security of Stored Sources of Radiation.

The licensee shall secure from unauthorized removal or access licensed sources of radiation that are stored in restricted or unrestricted areas.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.453.

- 64E-5.321 Control of Sources of Radiation Not in Storage.

(1) The licensee shall control and maintain constant surveillance of licensed radioactive material that is in an unrestricted area and that is not in storage or in a patient.

(2) The registrant shall maintain control of radiation machines that are in a restricted or unrestricted area and that are not in storage.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.454.

- 64E-5.322 Caution Signs.

(1) Standard Radiation Symbol. Unless otherwise authorized by the department, the symbol prescribed in this section shall use the colors magenta or purple or black on yellow background. The symbol prescribed is the three-bladed design as follows:

RADIATION SYMBOL

(a) Cross-hatched area is to be magenta or purple or black, and

(b) The background is to be yellow.

(2) Exception to Color Requirements for Standard Radiation Symbol. In spite of the requirements of subsection 64E-5.322(1), F.A.C., licensees or registrants are authorized to label sources, source holders, or device components containing sources of radiation that are subjected to high temperatures, with conspicuously etched or stamped radiation caution symbols and without a color requirement.

(3) Additional Information on Signs and Labels. In addition to contents of signs and labels prescribed in this part, the licensee or registrant shall provide on or near the required signs and labels additional information to make individuals aware of potential radiation exposures and to minimize the exposures.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.455.

- 64E-5.323 Posting Requirements.

(1) Posting of Radiation Areas. The licensee or registrant shall post each radiation area with a conspicuous sign or signs bearing the radiation symbol and the words “CAUTION, RADIATION AREA.”

(2) Posting of High Radiation Areas. The licensee or registrant shall post each high radiation area with a conspicuous sign or signs bearing the radiation symbol and the words “CAUTION, HIGH RADIATION AREA” or “DANGER, HIGH RADIATION AREA.”

(3) Posting of Very High Radiation Areas. The licensee or registrant shall post each very high radiation area with a conspicuous sign or signs bearing the radiation symbol and words “GRAVE DANGER, VERY HIGH RADIATION AREA.”

(4) Posting of Airborne Radioactivity Areas. The licensee shall post each airborne radioactivity area with a conspicuous sign or signs bearing the radiation symbol and the words “CAUTION, AIRBORNE RADIOACTIVITY AREA” or “DANGER, AIRBORNE RADIOACTIVITY AREA.”

(5) Posting of Areas or Rooms in which Licensed Material is Used or Stored. The licensee shall post each area or room in which there is used or stored an amount of licensed material exceeding 10 times the quantity of such material specified in State of Florida Bureau of Radiation Control Radioactive Material Requiring Labeling, May 2000, which is herein incorporated by reference and which is available from the department, with a conspicuous sign or signs bearing the radiation symbol and the words “CAUTION, RADIOACTIVE MATERIAL(S)” or “DANGER, RADIOACTIVE MATERIAL(S).”

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.456, Amended 10-8-00.

- 64E-5.324 Exceptions to Posting Requirements.

(1) A licensee or registrant is not required to post caution signs in areas or rooms containing sources of radiation for periods of less than 8 hours if each of the following conditions is met:

(a) The sources of radiation are constantly attended during these periods by an individual who takes the precautions necessary to prevent the exposure of individuals to sources of radiation in excess of the limits established in this part; and

(b) The area or room is subject to the licensee’s or registrant’s control.

(2) Rooms or other areas in hospitals that are occupied by patients are not required to be posted with caution signs as specified in Rule 64E-5.323, F.A.C., if the patient could be released from confinement as specified in Rule 64E-5.622, F.A.C.

(3) A room or area is not required to be posted with a caution sign because of the presence of a sealed source provided the radiation level at 30 centimeters from the surface of the sealed source container or housing does not exceed 0.005 rem (0.05 millisievert) per hour.

(4) A room or area is not required to be posted with a caution sign because of the presence of radiation machines used solely for diagnosis in the healing arts.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.457.

- 64E-5.325 Labeling Containers and Radiation Machines.

(1) The licensee shall ensure that each container of licensed material bears a durable, clearly visible label bearing the radiation symbol and the words “CAUTION, RADIOACTIVE MATERIAL” or “DANGER, RADIOACTIVE MATERIAL.” The label shall also provide information, such as the radionuclides present, an estimate of the quantity of radioactivity, the date for which the activity is estimated, radiation levels, kinds of materials, and mass enrichment, to permit individuals handling or using the containers or working in the vicinity of the containers to take precautions to avoid or minimize exposures.

(2) Each licensee prior to removal or disposal of empty uncontaminated containers to unrestricted areas shall remove or deface the radioactive material label or otherwise clearly indicate that the container no longer contains radioactive materials.

(3) Each registrant shall ensure that each radiation machine is labeled in a conspicuous manner which cautions individuals that radiation is produced when it is energized.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.458.

- 64E-5.326 Exemptions to Labeling Requirements.

A licensee is not required to label:

(1) Containers holding licensed material in quantities less than the quantities listed in State of Florida Bureau of Radiation Control Radioactive Material Requiring Labeling, May 2000;

(2) Containers holding licensed material in concentrations less than those specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table III;

(3) Containers attended by an individual who takes the precautions necessary to prevent the exposure of individuals in excess of the limits established by this part;

(4) Containers when they are in transport and packaged and labeled as specified by the rules of the U.S. Department of Transportation;

(5) Containers that are accessible only to individuals authorized to handle or use them or to work in the vicinity of the containers if the contents are identified to these individuals by a readily available written record. Examples of containers of this type are containers in locations such as water-filled canals, storage vaults, or hot cells. The record shall be retained as long as the containers are in use for the purpose indicated on the record; or

(6) Installed manufacturing or process equipment, such as piping and tanks.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Formerly 10D-91.459, Amended 10-8-00, 12-26-13.

- 64E-5.327 Procedures for Receiving and Opening Packages.

(1) Each licensee who expects to receive a package containing quantities of radioactive material in excess of A1 or A2 quantities as defined in Part XV shall make arrangements to receive:

(a) The package when the carrier offers it for delivery; or

(b) The notification of the arrival of the package at the carrier’s terminal and to take possession of the package expeditiously.

(2) Each licensee shall:

(a) Monitor the external surfaces of a package for radioactive contamination that are labeled with a Radioactive White I, Yellow II, or Yellow III label as specified in U.S. Department of Transportation regulations in 49 C.F.R. 172.403 and 172.436-.440, unless the package contains only radioactive material in the form of gas or in special form as defined in Part XV;

(b) Monitor the external surfaces of a package for radiation levels that are labeled with a Radioactive White I, Yellow II, or Yellow III label as specified in U.S. Department of Transportation regulations in 49 C.F.R. 172.403 and 172.436-.440, unless the package contains quantities of radioactive material that are less than or equal to the A1 or A2 quantities as defined in Part XV; and

(c) Monitor all packages known to contain radioactive material for radioactive contamination and radiation levels if there is evidence of degradation of package integrity, such as packages that are crushed, wet, or damaged.

(3) The licensee or registrant shall perform the monitoring required by subsection 64E-5.327(2), F.A.C., as soon as practicable after receipt of the package but not later than 3 hours after the package is received at the licensee’s facility if it is received during the licensee’s normal working hours, or not later than 3 hours from the beginning of the next working day if it is received after working hours.

(4) The licensee shall immediately notify the final delivery carrier and the department by telephone and telegram, mailgram, or facsimile when:

(a) Removable radioactive surface contamination exceeds the limits of subsection 64E-5.1505(8), F.A.C.; or

(b) External radiation levels exceed the limits of subsection 64E-5.1505(9), F.A.C.

(5) Each licensee shall:

(a) Establish, maintain, and retain written procedures for safely opening packages in which radioactive material is received; and

(b) Ensure that the procedures are followed and that consideration is given to special instructions for the type of package being opened.

(6) Licensees transferring special form sources in vehicles owned or operated by the licensee to and from a work site are exempt from the contamination monitoring requirements of paragraph 64E-5.327(2)(a), F.A.C., but are not exempt from the monitoring requirement in paragraph 64E-5.327(2)(b), F.A.C., for measuring radiation levels that ensures that the source is still properly lodged in its shield.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.460.

- 64E-5.328 General Requirements.

(1) Unless otherwise exempted, a licensee shall transfer waste for disposal, discharge, or decay licensed material only:

(a) By transfer to an authorized recipient as specified in Rule 64E-5.332, F.A.C., or in Part II of these regulations or to the U.S. Department of Energy;

(b) By decay in storage;

(c) By release in effluents within the limits in Rule 64E-5.312, F.A.C.; or

(d) As authorized in this subpart.

(2) A person shall be specifically licensed to receive waste containing licensed material from other persons for:

(a) Treatment prior to disposal;

(b) Treatment by incineration;

(c) Decay in storage;

(d) Disposal at a licensed land disposal facility; or

(e) Storage until transferred to a storage or disposal facility authorized to receive the waste.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.461.

- 64E-5.329 Method of Obtaining Approval of Proposed Disposal Procedures.

(1) A person can apply to the department for approval of proposed procedures to dispose of radioactive material in a manner not otherwise authorized in this part. Each application shall include a description of the radioactive material, including the quantities and kinds of radioactive material and levels of radioactivity involved, and the proposed manner and conditions of disposal. The application where appropriate should also include an analysis and evaluation of pertinent information of the nature of the environment, including topographical, geological, meteorological, and hydrological characteristics; usage of ground and surface waters in the general area; the nature and location of other potentially affected facilities; and procedures to be observed to minimize the risk of unexpected or hazardous exposure.

(2) The department will not approve any application for a licensee to receive radioactive material from other persons for disposal on land not owned by a state or the federal government.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.462.

- 64E-5.330 Discharge by Release into Sanitary Sewerage.

(1) A licensee can discharge licensed material into sanitary sewerage if each of the following conditions is satisfied:

(a) The material is readily soluble or is readily dispersible biological material in water;

(b) The quantity of licensed radioactive material that the licensee releases into the sewer in 1 month divided by the average monthly volume of water released into the sewer by the licensee does not exceed the concentration listed in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, Table III;

(c) If more than one radionuclide is released, the following conditions must also be satisfied;

1. The licensee shall determine the fraction of the limit in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, Table III represented by discharges into sanitary sewerage by dividing the actual monthly average concentration of each radionuclide released by the licensee into the sewer by the concentration of that radionuclide listed in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table III; and

2. The sum of the fractions for each radionuclide required by subparagraph 64E-5.330(1)(c)1., F.A.C., does not exceed unity; and

(d) The total quantity of licensed radioactive material that the licensee releases into the sanitary sewerage in a year does not exceed 5 curies (185 gigabecquerels) of hydrogen 3, 1 curie (37 gigabecquerels) of carbon 14, and 1 curie (37 gigabecquerels) of all other radioactive materials combined.

(2) Excreta from individuals undergoing medical diagnosis or therapy with radioactive material are not subject to the limitations contained in subsection 64E-5.330(1), F.A.C.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 1-1-94, Formerly 10D-91.463, Amended 12-26-13.

SUBPART J WASTE MANAGEMENT

- 64E-5.331 Disposal of Specific Wastes.

(1) A licensee can dispose of the following licensed material without regard to its radioactivity:

(a) 0.05 microcurie (1.85 kBq) or less of hydrogen 3 or carbon 14 per gram of medium used for liquid scintillation counting;

(b) 0.05 microcurie (1.85 kBq) or less of hydrogen 3 or carbon 14 per gram of animal tissue, averaged over the weight of the entire animal.

(c) Any radioactive material which is not a sealed source with a physical half-life of less than 120 days if all of the following are met:

1. Radioactive material to be disposed is held for decay in storage a minimum of 10 half-lives;

2. The radioactive material is monitored at the container surface before disposal as ordinary trash and its radioactivity cannot be distinguished from the background radiation level in a low background radiation area with an appropriate radiation survey instrument set on its most sensitive scale and with no interposed shielding;

3. All radiation labels are removed or obliterated, unless specifically authorized in writing or license condition by the department;

4. Each generator column is separated and monitored individually with all radiation shielding removed to ensure that its contents have decayed to background levels before disposal; and

5. The licensee shall retain a record of each disposal for 3 years. The record shall include:

a. The date of the disposal;

b. The date on which the radioactive material was placed in storage;

c. The radionuclides disposed;

d. The model and serial number of the radiation survey instrument used;

e. The background dose rate;

f. The radiation dose rate measured at the surface of each container; and

g. The name of the individual who performed the disposal.

(d) Licensed material as defined in paragraphs 64E-5.101(21)(c) and (d), F.A.C., may be disposed of at a licensed low-level radioactive waste disposal facility, even though it is not defined as low-level radioactive waste provided the requirements of Rule 64E-5.332, F.A.C., are satisfied or at a disposal facility authorized to dispose of such material in accordance with any Federal or State solid or hazardous waste law, including the Solid Waste Disposal Act, as authorized under the Energy Policy Act of 2005.

(2) A licensee shall not dispose of tissue as specified in subsection 64E-5.331(1), F.A.C., in a manner that would permit its use either as food for humans or as animal feed.

(3) The licensee shall maintain records as specified in Rule 64E-5.340, F.A.C.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.465, Amended 2-11-10, 12-26-13.

- 64E-5.332 Transfer for Disposal and Manifests.

(1) The requirements of this section, Requirements for Transfers of Low-Level Radioactive Waste Intended for Disposal at Licensed Land Disposal Facilities and Manifest, July 1997, hereafter referred to as “Requirements for Low-Level Radioactive Waste Disposal,” which is herein incorporated by reference and which is available from the department, and Part XV are designed to control transfers of low-level radioactive waste by any waste generator, waste collector, or waste processor licensee, as defined in Requirements for Low-Level Radioactive Waste Disposal, who ships low-level waste directly or indirectly through a waste collector or waste processor to a licensed low-level waste land disposal facility as defined in Requirements for Low-Level Radioactive Waste Disposal, establish a manifest tracking system, and supplement existing requirements concerning transfers and recordkeeping for those wastes. Requirements for Low-Level Radioactive Waste Disposal incoporates NRC Form 540 (3-95), Uniform Low-Level Radioactive Waste Manifest – Shipping Paper; NRC Form 541 (11-96), Uniform Low-Level Radioactive Waste Manifest – Container and Waste Description; and NRC Form 542 (3-95), Uniform Low-Level Radioactive Waste Manifest – Manifest Index and Regional Compact Tabulation.

(2) Prior to March 1, 1998, each shipment of radioactive waste designated for disposal at a licensed low-level radioactive waste disposal facility shall be accompanied by a shipment manifest as specified in subsection 64E-5.333(12), F.A.C. Beginning March 1, 1998, any licensee shipping radioactive waste intended for ultimate disposal at a licensed land disposal facility shall document the information required on forms specified in Requirements for Low-Level Radioactive Waste Disposal and transfer this recorded information to the intended consignee as specified in Requirements for Low-Level Radioactive Waste Disposal.

(3) Prior to March 1, 1998, each shipment manifest shall include a certification by the waste generator as specified in subsection 64E-5.333(12), F.A.C. Beginning March 1, 1998, each shipment manifest shall include a certification by the waste generator as specified in Requirements for Low-Level Radioactive Waste Disposal.

(4) Prior to March 1, 1998, each person involved in the transfer of waste for disposal, including the waste generator, waste collector, waste processor and disposal facility operator, shall comply with the requirements specified in subsection 64E-5.333(12), F.A.C. Beginning March 1, 1998, each person participating in the transfer of waste for disposal, including the waste generator, waste collector, waste processor and disposal facility operator, shall comply with the requirements specified in Requirements for Low-Level Radioactive Waste Disposal.

Rulemaking Authority 404.051, 404.081, 404.20 FS. Law Implemented 404.051(1), (4), 404.081, 404.20 FS. History–New 1-1-94, Formerly 10D-91.466, Amended 5-18-98.

- 64E-5.333 Classification and Characteristics of Low-Level Radioactive Waste for Near-Surface Land Disposal, Labeling and Manifest Requirements.

(1) Physical Half-life Considerations.

(a) Consideration by the licensee must be given to the concentration of long-lived radionuclides, and their shorter-lived precursors, whose potential hazard will persist long after such precautions as institutional controls, improved waste form and deeper disposal have ceased to be effective. These precautions delay the time when long-lived radionuclides could cause exposures. In addition, the magnitude of the potential dose is limited by the concentration and availability of the radionuclide at the time of exposure.

(b) Consideration by the licensee must be given to the concentration of shorter-lived radionuclides for which requirements on institutional controls, waste form and disposal methods are effective.

(2) Classes of Low Level Radioactive Waste.

(a) Class A waste is waste that is usually segregated from other waste classes at the disposal site. The physical form and characteristics of Class A waste must meet the minimum requirements set forth in paragraph (9)(a), below. If Class A waste also meets the stability requirements set forth in paragraph (9)(b), below, it is not necessary to segregate the waste for disposal.

(b) Class B waste is waste that must meet more rigorous requirements on waste form to ensure stability after disposal. The physical form and characteristics of Class B waste must meet both the minimum and stability requirements set forth in subsection (9), below.

(c) Class C waste is waste that not only must meet more rigorous requirements on waste form to ensure stability but also requires additional measures at the disposal facility to protect against inadvertent intrusion. The physical form and characteristics of Class C waste must meet both the minimum and stability requirements set forth in subsection (9), below.

(3) Classification of Low Level Radioactive Waste Determined by Long-Lived Radionuclides Present. If the low-level radioactive waste contains only radionuclides listed in Table 1, classification shall be determined as follows:

(a) If the concentration does not exceed 0.1 times the value in Table 1, the waste is Class A.

(b) If the concentration exceeds 0.1 times the value in Table 1, the waste is Class C.

(c) If the concentration exceeds the value in Table 1, the waste is not generally acceptable for near-surface disposal.

(d) For wastes containing mixtures of radionuclides listed in Table 1, the total concentration shall be determined by the sum of fractions rule described in (7), below.

Table 1

Concentration

Radionuclide	(Curies per cubic meter)

C-14	8

C-14 in activated metal	80

Ni-59 in activated metal	220

Nb-94 in activated metal	0.2

Tc-99	3

I-129	0.08

Radionuclide	(Nanocuries per gram)

Alpha emitting transuranic radionuclides	100

with half-life greater than 5 years

Pu-241	3,100

Cm-242	20,000

Ra-226	100

(e) To convert nanocuries to becquerels (Bq), multiply by 37. To convert curies to gigabecquerels (GBq), multiply by 37.

(4) Classification of Low Level Radioactive Waste Determined by Short-Lived Radionuclides Present. If the low level radioactive waste does not contain any of the radionuclides listed in Table 1, classification shall be determined based on the concentrations shown in Table 2.

(a) If the radioactive waste does not contain any radionuclides listed in either Table 1 or 2, the waste is Class A.

(b) If the concentration does not exceed the value in Table 2, Column 1, the waste is Class A.

(c) If the concentration exceeds the value in Table 2, Column 1, but does not exceed the value in Column 2, the waste is Class B.

(d) If the concentration exceeds the value in Table 2, Column 2, but does not exceed the value in Column 3, the waste is Class C.

(e) If the concentration exceeds the value in Table 2, Column 3, the waste is not generally acceptable for near-surface disposal.

(f) For wastes containing mixtures of the radionuclides listed in Table 2, the total concentration shall be determined by the sum of fractions rule described in subsection (7), below.

Table 2

Concentration Curies per cubic meter

Radionuclide	Column 1 	Column 2	Column 3

Total of all radionuclides

with less than 5 year

half-life 	700	See paragraph (4)(g), below

H-3	40 	See paragraph (4)(g), below

Co-60	700 	See paragraph (4)(g), below

Ni-63	3.5 	70	700

Ni-63 in activated metal	35 	700	7,000

Sr-90	0.04	150	7,000

Cs-137	1 	44	4,600

(g) There are not limits established for the radionuclides specified in Table 2 for Class B or C wastes. Practical considerations such as the effects of external radiation and internal heat generation on transportation, handling and disposal will limit the concentrations for such wastes. Such wastes shall be Class B unless the concentrations of other radionuclides in Table 2 determine the waste to be Class C, independent of these radionuclides.

(5) Classification of Low Level Radioactive Waste Determined by Both Long- and Short-Lived Radionuclides Present. If the radioactive waste contains a mixture of radionuclides, some of which are listed in Table 1, and some of which are listed in Table 2, classification shall be determined as follows:

(a) If the concentration of a radionuclide listed in Table 1 is less than 0.1 times the value listed in Table 1, the class shall be that determined by the concentration of radionuclides listed in Table 2.

(b) If the concentration of a radionuclide listed in Table 1 exceeds 0.1 times the value listed in Table 1, the waste shall be Class C, provided the concentration of radionuclides listed in Table 2 does not exceed the value shown in Table 2, Column 3.

(6) Classification of Low Level Radioactive Wastes with Radionuclides Other Than Those Listed in Tables 1 and 2. If the radioactive waste does not contain any radionuclides listed in either Table 1 or 2, it is Class A.

(7) The Sum of the Fractions Rule for Mixtures of Radionuclides. For determining classification for waste that contains a mixture of radionuclides, it is necessary to determine the sum of fractions by dividing each radionuclide’s concentration by the appropriate limit and adding the resulting values. The appropriate limits must all be taken from the same column of the same table. The sum of the fractions for the column must be less than 1.0 if the waste class is to be determined by that column.

Example: A waste contains Sr-90 in a concentration of 50 Ci per m3 (1.85 TBq per m3) and Cs-137 in a concentration of 22 Ci per m3 (814 GBq per m3). Since the concentrations both exceed the values in Table 2, Column 1, they must be compared to Column 2 values. For Sr-90 fraction 50/150 = 0.33; for Cs-137 fraction, 22/44 = 0.5; the sum of the fractions = 0.83. Since the sum is less than 1.0, the waste is Class B.

(8) Determination of Concentrations in Low Level Radioactive Wastes. The concentration of a radionuclide may be determined by indirect methods such as use of scaling factors which relate the inferred concentration of one radionuclide to another that is measured, or radionuclide material accountability, if there is reasonable assurance that the indirect methods can be correlated with actual measurements. The concentration of a radionuclide may be averaged over the volume of the waste, or weight of the waste, if the units are expressed as nanocuries per gram.

(9) Low Level Radioactive Waste Characteristics.

(a) The following are minimum requirements for all three classes of radioactive waste specified in subsection (2), above, and are intended to facilitate handling and provide protection of health and safety of personnel at the disposal site.

1. Radioactive wastes should be packaged in conformance with the conditions of the license issued to the site operator to which the waste will be shipped.

2. Radioactive wastes shall not be packaged for disposal in cardboard or fiberboard boxes.

3. Liquid radioactive waste must be solidified or packaged in sufficient absorbent material to absorb twice the volume of the liquid.

4. Solid radioactive waste containing liquid shall contain as little freestanding and noncorrosive liquid as is reasonably achievable, but in no case shall the liquid be corrosive or exceed 1 percent of the volume.

5. Radioactive waste shall not be readily capable of detonation or of explosive decomposition or reaction at normal pressures and temperatures, or of explosive reaction with water.

6. Radioactive waste shall not contain, or be capable of generating, quantities of toxic gases, vapors or fumes harmful to persons transporting, handling or disposing of the waste. This does not apply to radioactive gaseous waste packaged in accordance with subparagraph (9)(a)8., below.

7. Pyrophoric materials contained in radioactive wastes shall be treated, prepared and packaged to be nonflammable.

8. Radioactive wastes in gaseous form shall be packaged at a gauge pressure that does not exceed 1.5 atmospheres at 20 degrees Celsius. Total activity shall not exceed 100 curies (3.7 TBq) per container.

9. Radioactive wastes containing hazardous, biological, pathogenic or infectious material shall be treated to reduce to the maximum extent practicable the potential hazard from the nonradiological materials.

(b) Radioactive waste shall have structural stability. A structurally stable waste form will generally maintain its physical dimensions and its form, under the expected disposal conditions such as weight of overburden and compaction equipment, the presence of moisture, and microbial activity, and internal factors such as radiation effects and chemical changes. Structural stability can be provided by the waste form itself, processing the waste to a stable form, or placing the waste in a disposal container or structure that provides stability after disposal.

(c) Notwithstanding the provisions in subparagraphs (9)(a)3. and 4., above, liquid radioactive wastes, or radioactive wastes containing liquid, shall be converted into a form that contains as little freestanding and noncorrosive liquid as reasonably achievable, but in no case shall the liquid be corrosive or exceed 1 percent of the volume of the waste when the waste is in a disposal container designed to ensure stability, or 0.5 percent of the volume of the waste for waste processed to a stable form.

(d) Void spaces within the radioactive waste and between the waste and its package shall be reduced to the extent practicable.

(10) Package Labeling of Low Level Radioactive Waste. Each package of radioactive waste shall be clearly labeled to identify whether it is Class A, Class B or Class C waste, in accordance with subsections (1) through (8), above.

(11) Reserved.

(12) Transfer for Disposal and Manifests of Low Level Radioactive Waste.

(a) Each shipment of radioactive waste to a licensed land disposal facility shall be accompanied by a shipment manifest that contains the name, address and telephone number of the person generating the waste. The manifest shall also include the name, address and telephone number or the name and U.S. Environmental Protection Agency hazardous waste identification number of the person transporting the waste to the land disposal facility. The manifest shall also indicate as completely as practicable:

1. A physical description of the waste;

2. The waste volume;

3. Radionuclide identity and quantity;

4. The total radioactivity; and

5. The principal chemical form.

The solidification agent, is used, shall be specified. Wastes containing more than 0.1 percent chelating agents by weight shall be identified and the weight percentage of the chelating agent estimated. Wastes classified as Class A, Class B or Class C in subsections (1) through (8), above, shall be clearly identified as such in the manifest. The total quantity of the radionuclides H-3, C-14, Tc-99 and I-129 shall be shown.

(b) The manifest required in paragraph (12)(a), above, may be shipping papers used to meet U.S. Department of Transportation or U.S. Environmental Protection Agency regulations or requirements of the receiver, provided all the required information is included.

(c) Each manifest shall include a certification by the radioactive waste generator that the transported materials are properly classified, described, packaged, marked, and labeled and are in proper condition for transportation according to the applicable regulations of the U.S. Department of Transportation and the Department. An authorized representative of the waste generator shall sign and date the manifest.

(d) Any licensee who transfers radioactive waste to a land disposal facility or a licensed waste collector shall comply with the following requirements. Any licensee who transfers waste to a licensed waste processor who treats or repackages waste shall comply with the requirements of subparagraphs (12)(d)4. through 8., below. A licensee shall:

1. Prepare all radioactive wastes so that the waste is classified according to subsections (1) through (8), above, and meets the waste characteristics requirements in subsection (9), above;

2. Label each package of waste to identify whether it is Class A waste, Class B waste or Class C waste, in accordance with subsections (1) through (8), above;

3. Conduct a quality control program to assure compliance with subsections (1) through (9), above; the program must include management evaluation of audits;

4. Prepare shipping manifests to meet the requirements of paragraphs (12)(a) and (c), above;

5. Forward a copy of the manifest to the intended recipient at the time of shipment; or, deliver to a collector at the time the radioactive waste is collected, obtaining acknowledgment of receipt in the form of a signed copy of the manifest from the collector;

6. Include one copy of the manifest with the shipment;

7. Retain a copy of the manifest with documentation of acknowledgment of receipt as the record of transfer of licensed material as required by these regulations; and

8. For any low level radioactive waste shipments or any part of a shipment for which acknowledgment of receipt has not been received within the times set forth in this subsection, conduct an investigation in accordance with paragraph (12)(g), below.

(e) Any waste collector licensee who handles only radioactive wastes that have been prepackaged shall:

1. Acknowledge receipt of the radioactive waste from the generator within 1 week of receipt by returning a signed copy of the manifest;

2. Prepare a new manifest to reflect consolidated shipments; the new manifest shall serve as a listing or index for the detailed generator manifests. Copies of the generator manifests shall be a part of the new manifest. The waste collector may prepare a new manifest without attaching the generator manifests, provided the new manifest contains for each package the information specified in paragraph (12)(a), above. The collector licensee shall certify that nothing has been done to the waste which would invalidate the generator’s certification;

3. Forward a copy of the new manifest to the land disposal facility operator at the time of shipment;

4. Include the new manifest with the radioactive waste shipment to the disposal site;

5. Retain a copy of the manifest with documentation of acknowledgment of receipt as the record of transfer of licensed material as required by these regulations, and retain information from generator manifests until disposition is authorized by the department; and

6. For any low level radioactive waste shipments or any part of a shipment for which acknowledgment of receipt is not received within the times set forth in this subsection, conduct an investigation in accordance with paragraph (12)(g), below.

(f) Any licensed waste processor who treats or repackages low level radioactive wastes shall:

1. Acknowledge receipt of the radioactive waste from the generator within 1 week of receipt by returning a signed copy of the manifest or equivalent documentation;

2. Prepare a new manifest that meets the requirements of paragraphs (12)(a) and (c), above. Preparation of the new manifest reflects that the processor is responsible for the waste;

3. Prepare all low level radioactive wastes so that the waste is classified according to subsections (1) through (8), above, and meets the waste characteristics requirement in subsection (9), above;

4. Label each package of waste to identify whether it is Class A waste, Class B waste or Class C waste, as specified in subsections (1) through (8) and (10), above;

5. Conduct a quality control program to assure compliance with subsections (1) through (9), above. The program shall include management evaluation of audits;

6. Forward a copy of the new manifest to the disposal site operator or waste collector at the time of shipment, or deliver to a collector at the time the radioactive waste is collected, obtaining acknowledgment of receipt in the form of a signed copy of the manifest by the collector;

7. Include the new manifest with the shipment;

8. Retain copies of original manifests and new manifests with documentation of acknowledgment of receipt as the record or transfer of licensed material required by these regulations; and

9. For any low level waste shipment or part of a shipment for which acknowledgment if not received within the times set forth in this section, conduct an investigation in accordance with paragraph (12)(g), below.

(g) The land disposal facility operator shall:

1. Acknowledge receipt of the waste within 1 week of receipt by returning a signed copy of the manifest or equivalent documentation to the shipper. The shipper to be notified is the licensee who last possessed the waste and transferred the waste to the operator. The returned copy of the manifest or equivalent documentation shall indicate any discrepancies between materials listed on the manifest and materials received;

2. Maintain copies of all completed manifests or equivalent documentation until the department authorizes their disposition; and

3. Notify the shipper and the department when any shipment or portion of a shipment has not arrived within 60 days after the advance manifest was received.

(h) Any low level radioactive waste shipment or part of a shipment for which acknowledgment is not received within the times set forth in this section must:

1. Be investigated by the shipper if the shipper has not received notification of receipt within 20 days after transfer; and

2. Be traced and reported. The investigation shall include tracing the shipment and filing a report with the Department. Each licensee who conducts a trace investigation shall file a written report with the department within 2 weeks of completion of the investigation.

(13) Beginning March 1, 1998, all licensees shall comply with Requirements for Low-Level Radioactive Waste Disposal. Prior to March 1, 1998, a low-level waste disposal facility operator or its regulatory authority can require the shipper to use requirements specified in subsection (12), above.

Rulemaking Authority 404.051, 404.061, 404.20 FS. Law Implemented 404.051(1), (4), 404.061(2), 404.20(1) FS. History–New 1-1-94, Amended 5-15-96, Formerly 10D-91.468, Amended 5-18-98.

- 64E-5.334 General Provisions.

(1) Each licensee or registrant shall use the SI unit becquerel, gray, sievert and coulomb per kilogram, or the special units curie, rad, rem and roentgen, including multiples and subdivisions, and shall clearly indicate the units of all quantities on records required by this part. The information on shipping manifests, specified in subsection 64E-5.332(2), F.A.C., shall be recorded in SI units or in SI units and special units curie, rad, rem and roentgen.

(2) The licensee or registrant shall make a clear distinction among the quantities entered on the records required by this part, such as total effective dose equivalent, shallow dose equivalent, lens dose equivalent, deep dose equivalent, or committed effective dose equivalent.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.469, Amended 5-18-98, 10-8-00.

- 64E-5.335 Records of Radiation Protection Programs.

(1) Each licensee or registrant shall maintain records of the radiation protection program, including:

(a) The provisions of the program; and

(b) Audits and other reviews of program content and implementation.

(2) The licensee or registrant shall retain the records required by paragraph 64E-5.335(1)(a), F.A.C., until the department terminates each pertinent license or registration requiring the record. The licensee or registrant shall retain the records required by paragraph 64E-5.335(1)(b), F.A.C., for 3 years after the record is made.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.470.

- 64E-5.336 Records of Surveys.

(1) Each licensee or registrant shall maintain records showing the results of surveys and calibrations required by Rule 64E-5.314 and subsection 64E-5.327(2), F.A.C. The licensee or registrant shall retain these records for 3 years after the record is made.

(2) The licensee or registrant shall retain each of the following records until the department terminates each pertinent license or registration requiring the record:

(a) Records of the results of surveys to determine the dose from external sources of radiation used in the assessment of individual dose equivalents in the absence of or in combination with individual monitoring data;

(b) Records of the results of measurements and calculations used to determine individual intakes of radioactive material and used in the assessment of internal dose;

(c) Records showing the results of air sampling, surveys, and bioassays specified in subparagraphs 64E-5.319(1)(c)1. and 2., F.A.C.; and

(d) Records of the results of measurements and calculations used to evaluate the release of radioactive effluents to the environment.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.471.

- 64E-5.337 Records of Tests for Leakage or Contamination of Sealed Sources.

Records of tests for leakage or contamination of sealed sources required by these regulations shall be kept in units of becquerel or microcurie and maintained for inspection by the department for 3 years after the records are made.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10E-91.472.

- 64E-5.338 Records of Planned Special Exposures.

(1) For each planned special exposure, the licensee or registrant shall maintain records that describe:

(a) The exceptional circumstances requiring the use of a planned special exposure;

(b) The name of the management official who authorized the planned special exposure and a copy of the signed authorization;

(c) What actions were necessary;

(d) Why the actions were necessary;

(e) What precautions were taken to assure that doses were maintained ALARA;

(f) What individual and collective doses were expected to result; and

(g) The doses actually received in the planned special exposure.

(2) The licensee or registrant shall retain the records until the department terminates each pertinent license or registration requiring these records.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.474.

- 64E-5.339 Records of Individual Monitoring Results.

(1) Each licensee or registrant shall maintain records of doses received by all individuals for whom monitoring is required as specified in Rule 64E-5.315, F.A.C., and records of doses received during planned special exposures, accidents, and emergency conditions. Assessments of dose equivalent and records made using units in effect before the effective date of the rule need not be changed. These records shall include when applicable:

(a) The deep dose equivalent to the whole body, lens dose equivalent, shallow dose equivalent to the skin, and shallow dose equivalent to the extremities;

(b) The estimated intake of radionuclides as specified in Rule 64E-5.305, F.A.C.;

(c) The committed effective dose equivalent assigned to the intake of radionuclides;

(d) The specific information used to calculate the committed effective dose equivalent as specified in subsection 64E-5.307(3), F.A.C.;

(e) The total effective dose equivalent when required by Rule 64E-5.305, F.A.C.; and

(f) The total of the deep dose equivalent and the committed dose to the organ receiving the highest total dose.

(2) Recordkeeping Frequency. The licensee or registrant shall make entries of the records specified in subsection 64E-5.339(1), F.A.C., annually.

(3) Recordkeeping Format. The licensee or registrant shall maintain the records specified in subsection 64E-5.339(1), F.A.C., on DOH Form 1622 July 1993, which is herein incorporated by reference and which is available from the department, according to the instructions for DOH Form 1622 July 1993, or in clear and legible records containing all the information required by DOH Form 1622 July 1993.

(4) The licensee or registrant shall maintain the records of dose to an embryo or fetus with the records of dose to the declared pregnant woman. The declaration of pregnancy including the estimated date of conception shall also be kept on file but can be maintained separately from the dose records.

(5) The licensee or registrant shall retain each required form or record until the department terminates each pertinent license or registration requiring the record.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.475, Amended 10-8-00.

- 64E-5.340 Records of Waste Disposal or Transfers.

(1) Each licensee shall maintain records of the disposal of licensed materials.

(2) The licensee shall retain the records required by subsection 64E-5.340(1), F.A.C., until the department terminates each pertinent license requiring the record except as provided in subparagraph 64E-5.331(1)(c)5. and subsection 64E-5.624(2), F.A.C.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.477.

- 64E-5.341 Records of Testing Entry Control Devices for Very High Radiation Areas.

(1) Each licensee or registrant shall maintain records of tests specified in paragraph 64E-5.317(2)(a), F.A.C., on entry control devices for very high radiation areas. The records must include the date, time, and results of each such test of function.

(2) The licensee or registrant shall retain the records required by subsection 64E-5.341(1), F.A.C., for 3 years after the record is made.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.478.

- 64E-5.342 Form of Records.

Each record required by this part shall be legible throughout the specified retention period. The record shall be the original or a reproduced copy or a microform if the copy or microform is authenticated by authorized personnel and the microform is capable of producing a clear copy throughout the required retention period. The record can also be stored in electronic media capable of producing legible, accurate, and complete records during the required retention period. Records such as letters, drawings, and specifications shall include all pertinent information such as stamps, initials, and signatures. The licensee or registrant shall maintain adequate safeguards against tampering with and loss of records.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.479.

- 64E-5.343 Reports of Stolen, Lost, or Missing Licensed or Registered Sources of Radiation.

(1) Telephone Reports. Each licensee or registrant shall report to the department by telephone the following:

(a) Stolen, lost or missing licensed radioactive material in an aggregate quantity equal to or greater than 1,000 times the quantity specified in State of Florida Bureau of Radiation Control Radioactive Material Requiring Labeling, May 2000, immediately after its occurrence becomes known to the licensee if it appears to the licensee that an exposure could result to individuals in unrestricted areas; or

(b) Lost, stolen, or missing licensed radioactive material in an aggregate quantity greater than 10 times the quantity specified in State of Florida Bureau of Radiation Control Radioactive Material Requiring Labeling, May 2000, that is still missing within 30 days after its occurrence becomes known.

(c) A stolen, lost, or missing radiation machine immediately after its occurrence becomes known.

(2) Written Reports. Each licensee or registrant required to make a report as specified in subsection 64E-5.343(1), F.A.C., shall make a written report to the department setting forth the following information within 30 days after making the telephone report:

(a) A description of the licensed or registered source of radiation involved, including, for radioactive material, the kind, quantity, and chemical and physical form; and, for radiation machines, the manufacturer, model and serial number, type and maximum energy of radiation emitted;

(b) A description of the circumstances under which the loss or theft occurred;

(c) A statement of disposition or probable disposition of the licensed or registered source of radiation involved;

(d) Exposures of individuals to radiation, circumstances under which the exposures occurred, and the possible total effective dose equivalent to persons in unrestricted areas;

(e) Actions that have been or will be taken to recover the source of radiation; and

(f) Procedures or measures that have been or will be adopted to ensure against a recurrence of the loss or theft of licensed or registered sources of radiation.

(3) Subsequent to filing the written report, the licensee or registrant shall also report additional substantive information on the loss or theft within 30 days after the licensee or registrant learns of such information.

(4) The licensee or registrant shall prepare any report filed with the department as specified in Rule 64E-5.343, F.A.C., so that names of individuals who have received exposure to radiation are stated in a separate and detachable portion of the report.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.480, Amended 10-8-00.

SUBPART REPORTS

- 64E-5.344 Notification of Incidents.

(1) Immediate Notification. Regardless of other requirements for notification, each licensee or registrant shall immediately report each event involving a source of radiation possessed by the licensee or registrant that might have caused or threatens to cause any of the following conditions:

(a) An individual to receive:

1. A total effective dose equivalent of 25 rem (0.25 sievert) or more;

2. A lens dose equivalent of 75 rem (0.75 sievert) or more; or

3. A shallow dose equivalent to the skin or extremities or a total organ dose equivalent of 250 rad (2.5 gray) or more; or

(b) The release of radioactive material inside or outside of a restricted area so that if an individual had been present for 24 hours, the individual could have received an intake five times the occupational ALI. This provision does not apply to locations where personnel are not normally stationed during routine operations, such as hot-cells or process enclosures.

(2) Twenty-Four Hour Notification. Each licensee or registrant shall report to the department within 24 hours of discovery of the event each event involving loss of control of a licensed or registered source of radiation possessed by the licensee or registrant that might have caused or threatens to cause any of the following conditions:

(a) An individual to receive in a period of 24 hours:

1. A total effective dose equivalent exceeding 5 rem (0.05 sievert);

2. A lens dose equivalent exceeding 15 rem (0.15 sievert); or

3. A shallow dose equivalent to the skin or extremities or a total organ dose equivalent exceeding 50 rem (0.5 sievert); or

(b) The release of radioactive material inside or outside of a restricted area so that if an individual had been present for 24 hours, the individual could have received an intake in excess of one occupational ALI. This provision does not apply to locations where personnel are not normally stationed during routine operations such as hot-cells or process enclosures.

(3) The licensee or registrant shall prepare each report filed with the department as specified Rule 64E-5.344, F.A.C., so that names of individuals who have received exposure to sources of radiation are stated in a separate and detachable portion of the report.

(4) Licensees or registrants shall make the reports required by subsections 64E-5.344(1) and (2), F.A.C., to the department by telephone, telegram, mailgram, or facsimile to the department.

(5) The provisions of Rule 64E-5.344, F.A.C., do not apply to doses that result from planned special exposures if such doses are within the limits for planned special exposures and are reported as specified in Rule 64E-5.346, F.A.C.

(6) Immediate notification. In addition to the other reporting requirements in these regulations, each licensee shall notify the department as soon as possible but not later than 4 hours after the discovery of an event, such as a fire, explosion, or toxic gas release, that prevents immediate protective actions necessary to avoid exposure to radiation or radioactive materials that could exceed regulatory limits or to avoid releases of licensed material that could exceed regulatory limits.

(7) Twenty-four hour report. Each licensee shall notify the Department within 24 hours after the discovery of any of the following events involving licensed material:

(a) An unplanned contamination event that:

1. Requires access to the contaminated area, by workers or the public, to be restricted for more than 24 hours by imposing additional radiological controls or by prohibiting entry into the area;

2. Involves a quantity of material greater than five times the lowest annual limit on intake of materials as specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012 (see 64E-5.101, F.A.C.); and

3. Has access to the area restricted for a reason other than to allow isotopes with a half-life of less than 24 hours to decay prior to decontamination.

(b) An event in which equipment is disabled or fails to function as designed when:

1. The equipment is required by regulation or license condition to prevent releases exceeding regulatory limits, to prevent exposure to radiation and radioactive materials exceeding regulatory limits, or to mitigate the consequences of an accident;

2. The equipment is required to be available and operable when it is disabled or fails to function; and

3. No redundant equipment is available and operable to perform the required safety function.

(c) An event that requires unplanned medical treatment at a medical facility of an individual with spreadable radioactive contamination on the individual’s clothing or body;

(d) An unplanned fire or explosion damaging any licensed material or any device, container, or equipment containing licensed materials when:

1. The quantity of material involved is five times the lowest annual limit on intake for material specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012; and

2. The damage affects the integrity of the licensed material or its container.

(e) Dose to an embryo/fetus that is greater than 50 mSv (5 rem) dose equivalent that is a result of an administration of radioactive material or radiation from radioactive material to a pregnant individual unless the dose to the embryo/fetus was specifically approved, in advance, by the authorized user as defined in Rule 64E-5.6011, F.A.C.

(f) Dose to a nursing child that is a result of an administration of radioactive material to a breast-feeding individual that meets one of the following:

1. Greater than 50 mSv (5 rem) total effective dose equivalent; or

2. Has resulted in unintended permanent functional damage to an organ or a physiological system of the child, as determined by a physician.

(8) Preparation and submission of reports. Reports made by licensees in response to the requirements of this section must be made as follows:

(a) Licensees shall make reports required by subsections 64E-5.344(6) and (7), F.A.C., by telephone to the department. If the information is available at the time of notification, the information provided in these reports must include:

1. The caller’s name and call back telephone number;

2. A description of the event, including date and time;

3. The exact location of the event;

4. The isotopes, quantities, and chemical and physical forms of the licensed material involved; and

5. Any personnel radiation exposure data available.

(b) Written report. Each licensee who makes a report required by subsections 64E-5.344(6) and (7), F.A.C., shall submit a written follow-up report within 30 days of the initial report. Written reports prepared as required by other regulations may be submitted to fulfill this requirement if the reports contain all of the necessary information. The reports must include the following:

1. A description of the event, including the probable cause and the manufacturer and model number of any equipment that failed or malfunctioned;

2. The exact location of the event;

3. The isotopes, quantities, and chemical and physical form of the licensed material involved;

4. Date and time of the event;

5. Corrective actions taken or planned and the results of any evaluations or assessments; and

6. The extent of exposure of individuals to radiation or to radioactive materials without identification of the individuals by name.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Amended 5-15-96, Formerly 10D-91.481, Amended 10-8-00, 2-11-10, 12-26-13.

- 64E-5.345 Reports of Exposures, Radiation Levels, Concentrations of Radioactive Material Exceeding the Constraints or Limits, Medical Events and Dose to an Embryo/Fetus or a Nursing Child.

(1) Reportable Events. In addition to the notification required by Rule 64E-5.344, F.A.C., each licensee or registrant shall submit a written report within 30 days after learning of any of the following occurrences:

(a) Incidents for which notification is required by Rule 64E-5.344, F.A.C.; or

(b) Doses in excess of any of the following:

1. The occupational dose limits for adults in Rule 64E-5.304, F.A.C.;

2. The occupational dose limits for a minor in Rule 64E-5.310, F.A.C.;

3. The limits for an embryo or fetus of a declared pregnant woman in Rule 64E-5.311, F.A.C.;

4. The limits for an individual member of the public in Rule 64E-5.312, F.A.C.;

5. Any applicable limit in the license or registration;

6. The ALARA constraints for air emissions specified in subsection 64E-5.303(5), F.A.C.; or

(c) Levels of radiation or concentrations of radioactive material in:

1. A restricted area in excess of applicable limits in the license or registration; or

2. An unrestricted area in excess of 10 times the applicable limit set forth in this part or in the license or registration, whether or not involving exposure of any individual in excess of the limits in Rule 64E-5.312, F.A.C.; or

(d) For licensees subject to the provisions of U.S. Environmental Protection Agency’s generally applicable environmental radiation standards in 40 CFR 190, levels of radiation or releases of radioactive material in excess of those standards, or of license conditions related to those standards.

(2) Contents of Reports.

(a) Each report required by subsection 64E-5.345(1), F.A.C., shall describe the extent of exposure of individuals to radiation and radioactive material, including as appropriate:

1. Estimates of each individual’s dose;

2. The levels of radiation and concentrations of radioactive material involved;

3. The cause of the elevated exposures, dose rates, or concentrations; and

4. Corrective steps taken or planned to ensure against a recurrence, including the schedule for achieving conformance with applicable limits, ALARA constraints, generally applicable environmental standards, and associated license or registration conditions.

(b) Each report filed as specified in subsection 64E-5.345(1), F.A.C., shall include for each occupationally overexposed individual: the name, social security account number, and date of birth. With respect to the limit for the embryo or fetus in Rule 64E-5.311, F.A.C., the identifiers should be those of the declared pregnant woman. The report shall be prepared so that this information is stated in a separate and detachable portion of the report.

(3) All licensees or registrants who make reports as specified in subsection 64E-5.345(1), F.A.C., shall submit the report in writing to the department.

(4) Reports of Medical Events.

(a) The licensee or registrant shall notify the department by telephone no later than the next calendar day after the discovery of the medical event. The licensee or registrant shall also notify the referring physician of the affected individual and the individual or a responsible relative or guardian, unless the referring physician personally informs the licensee either that he will inform the individual or believes, based on medical judgment, that telling the individual or the individual’s responsible relative or guardian would be harmful to either. These notifications shall be made within 24 hours after the licensee or registrant discovers the medical event. If the referring physician, individual or the individual’s responsible relative or guardian cannot be reached within 24 hours, the licensee or registrant shall notify them as soon as practicable. The licensee is not required to notify the individual or the individual’s responsible relative or guardian without first consulting the referring physician; however, the licensee or registrant shall not delay medical care for the individual because of this. If a verbal notification is made, the licensee shall inform the individual, or appropriate responsible relative or guardian, that a written description of the event can be obtained from the licensee upon request. The licensee shall provide such a written description if requested.

(b) Written Report. Within 15 days after the medical event report to the department, the licensee or registrant shall report in writing to the department and to the referring physician and furnish a copy of the report to the individual or the individual’s responsible relative or guardian if either was previously notified by the licensee or registrant as specified in (4)(a), above, or a brief description of both event and consequences as they affect the individual or the individual’s responsible relative or guardian if a statement is included that the report submitted to the department can be obtained from the licensee or registrant. The written report shall include the licensee’s or registrant’s name; the prescribing physician’s name; the referring physician’s name; a brief description of the event; why the event occurred; the effect on the individual; the action taken to prevent recurrence; whether the licensee or registrant informed the individual or the individual’s responsible relative or guardian and what information was provided to the individual or individual’s responsible relative or guardian, and if not, a written medical justification. The report shall not include the individual’s name or other information that could lead to identification of the individual.

(5) Records of medical event. Each licensee or registrant shall retain a record of each medical event for 20 years. The record shall contain the names of all individuals involved in the event, including the prescribing physician, the allied health personnel, the individual, and the individual’s referring physician, the individual’s identification number if one has been assigned, a brief description of the event, why it occurred, the effect on the individual, what improvements are needed to prevent recurrence, and the actions taken, if any, to prevent recurrence.

(6) Rights and Duties of Licensees or Registrants. Aside from the notification requirement, nothing in this section shall affect any rights or duties of licensees, registrants or physicians in relation to each other, the individual, or responsible relatives or guardians.

(7) Reports of a dose to an embryo/fetus or a nursing child.

(a) The licensee shall provide notification of the event to the referring physician and also notify the pregnant individual or mother, both hereafter referred to as the mother, no later than 24 hours after discovery of an event that would require reporting under paragraph 64E-5.344(7)(e) or (f), F.A.C., unless the referring physician personally informs the licensee either that he or she will inform the mother or that, based on medical judgment, telling the mother would be harmful. The licensee is not required to notify the mother without first consulting with the referring physician. If the referring physician or mother cannot be reached within 24 hours, the licensee shall make the appropriate notifications as soon as possible thereafter. The licensee may not delay any appropriate medical care for the embryo/fetus or for the nursing child, including any necessary remedial care as a result of the event, because of any delay in notification. To meet the requirements of this subsection, the notification may be made to the mother or child's responsible relative or guardian instead of the mother. If a verbal notification is made, the licensee shall inform the mother, or the mother or child’s responsible relative or guardian, that a written description of the event can be obtained from the licensee upon request. The licensee shall provide such a written description if requested.

(b) Written Report.

1. Within 15 days after the discovery of an event that would require reporting under paragraph 64E-5.344(7)(e) or (f), F.A.C., the licensee or registrant shall report in writing as described below, to the department and to the referring physician.

2. Within 15 days after the discovery of an event that would require reporting under paragraph 64E-5.344(7)(e) or (f), F.A.C., the licensee or registrant shall also furnish a copy of the report or a brief description of both the event and the consequences of the event as they affect the embryo/fetus or nursing child, to the mother, or the mother or child’s responsible relative or guardian if either was previously notified by the licensee or registrant as specified in paragraph (7)(a), above. If a brief description of both the event and consequences of the event is provided in lieu of the report, such description shall include a statement that the report submitted to the department can be obtained from the licensee or registrant.

3. The written report shall include the licensee's or registrant's name, the prescribing physician’s name, the referring physician's name, a brief description of the event, why the event occurred, the effect on the embryo/fetus or nursing child, the action taken to prevent recurrence, whether the licensee or registrant informed the pregnant individual or mother or the mother’s or child’s responsible relative or guardian and what information was provided to the individual or individual’s responsible relative or guardian, and if not, a written medical justification. The report shall not include the individual’s or child’s name or other information that could lead to identification of the individual or child.

(8) Records of reports of dose to an embryo/fetus or a nursing child. Each licensee or registrant shall retain a record of each report of dose to an embryo/fetus or a nursing child for 20 years. The record shall contain the names of all individuals involved in the event, including the prescribing physician, the allied health personnel, mother or the nursing child’s name, and the mother or nursing child’s referring physician, the social security number of the mother, the nursing child’s social security number or identification number if either has been assigned, a brief description of the event, why it occurred, the effect on the individual, what improvements are needed to prevent recurrence, and the actions taken, if any, to prevent recurrence.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.482, Amended 10-8-00, 2-11-10.

- 64E-5.346 Reports of Planned Special Exposures.

The licensee or registrant shall submit a written report to the department within 30 days following any planned special exposure as specified in Rule 64E-5.309, F.A.C., informing the department that a planned special exposure was conducted and indicating the date the planned special exposure occurred and the information required by Rule 64E-5.338, F.A.C.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.483.

- 64E-5.347 Notifications and Reports to Individuals.

(1) Requirements for notification and reports to individuals of exposure to radiation or radioactive material are specified in Part IX of these regulations.

(2) When a licensee or registrant is required by Rules 64E-5.345, 64E-5.346 or 64E-5.347, F.A.C., to report to the department any occupational exposure of an individual or an identified member of the public to radiation or radioactive material, the licensee or registrant shall also provide a copy of the report submitted to the department to the individual. Such notice shall be transmitted no later than the transmittal to the department, and shall comply with the provisions of Part IX.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.484, Amended 5-18-98.

- 64E-5.348 Reports of Leaking or Contaminated Sealed Sources.

The licensee shall immediately notify the department if the test for leakage or contamination required by these regulations indicates a sealed source is leaking or contaminated. A written report of a leaking or contaminated source shall be submitted to the department within 5 days. The report shall include the equipment involved, the test results and the corrective action taken.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.485.

- 64E-5.349 Vacating Premises.

Each specific licensee or registrant shall notify the department in writing of the intent to vacate no less than 30 days before vacating or relinquishing possession or control of premises which might have been contaminated with radioactive material as a result of his activities. The licensee shall decommission the premises for subsequent use as an unrestricted area.

Rulemaking Authority 404.051, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.081 FS. History–New 1-1-94, Formerly 10D-91.486.

- 64E-5.350 Reports of Transactions Involving Nationally Tracked Sources.

Each licensee who manufactures, transfers, receives, disassembles, or disposes of a nationally tracked source shall complete and submit to the NRC a National Source Tracking Transaction Report as specified in subsections (1) through (5) of this section for each type of transaction.

(1) Each licensee who manufactures a nationally tracked source shall complete and submit a National Source Tracking Transaction Report. The report must include the following information:

(a) The name, address, and license number of the reporting licensee;

(b) The name of the individual preparing the report;

(c) The manufacturer, model, and serial number of the source;

(d) The radioactive material in the source;

(e) The initial source strength in becquerels (curies) at the time of manufacture; and

(f) The manufacture date of the source.

(2) Each licensee that transfers a nationally tracked source to another person shall complete and submit a National Source Tracking Transaction Report. The report must include the following information:

(a) The name, address, and license number of the reporting licensee;

(b) The name of the individual preparing the report;

(c) The name and license number of the recipient facility and the shipping address;

(d) The manufacturer, model, and serial number of the source or, if not available, other information to uniquely identify the source;

(e) The radioactive material in the source;

(f) The initial or current source strength in becquerels (curies);

(g) The date for which the source strength is reported;

(h) The shipping date;

(i) The estimated arrival date; and

(j) For nationally tracked sources transferred as waste under a Uniform Low-Level Radioactive Waste Manifest, the waste manifest number and the container identification of the container with the nationally tracked source.

(3) Each licensee that receives a nationally tracked source shall complete and submit a National Source Tracking Transaction Report. The report must include the following information:

(a) The name, address, and license number of the reporting licensee;

(b) The name of the individual preparing the report;

(c) The name, address, and license number of the person that provided the source;

(d) The manufacturer, model, and serial number of the source or, if not available, other information to uniquely identify the source;

(e) The radioactive material in the source;

(f) The initial or current source strength in becquerels (curies);

(g) The date for which the source strength is reported;

(h) The date of receipt; and

(i) For material received under a Uniform Low-Level Radioactive Waste Manifest, the waste manifest number and the container identification with the nationally tracked source.

(4) Each licensee that disassembles a nationally tracked source shall complete and submit a National Source Tracking Transaction Report. The report must include the following information:

(a) The name, address, and license number of the reporting licensee;

(b) The name of the individual preparing the report;

(c) The manufacturer, model, and serial number of the source or, if not available, other information to uniquely identify the source;

(d) The radioactive material in the source;

(e) The initial or current source strength in becquerels (curies);

(f) The date for which the source strength is reported; and

(g) The disassemble date of the source.

(5) Each licensee who disposes of a nationally tracked source shall complete and submit a National Source Tracking Transaction Report. The report must include the following information:

(a) The name, address, and license number of the reporting licensee;

(b) The name of the individual preparing the report;

(c) The waste manifest number;

(d) The container identification with the nationally tracked source.

(e) The date of disposal; and

(f) The method of disposal.

(6) The National Source Tracking Transaction Report discussed in subsections (1) through (5) of this section must be submitted to the NRC by the close of the next business day after the transaction. A single report may be submitted for multiple sources and transactions. The reports must be submitted to the National Source Tracking System by using:

(a) The on-line National Source Tracking System;

(b) Electronically using a computer-readable format;

(c) By facsimile;

(d) By mail to the address on the NRC Form 748 National Source Tracking Transaction Report Form; or

(e) By telephone with followup by facsimile or mail.

(7)(a) Each licensee shall correct any error in previously filed reports or file a new report for any missed transaction within 5 business days of the discovery of the error or missed transaction. Such errors may be detected by a variety of methods such as administrative reviews or by physical inventories required by regulation.

(b) In addition, every year each licensee shall reconcile the inventory of nationally tracked sources possessed by the licensee against that licensee's data in the National Source Tracking System. The reconciliation must be conducted during the month of January in each year. The reconciliation process must include resolving any discrepancies between the National Source Tracking System and the actual inventory by filing the reports identified by subsections (1) through (5) of this section.  In order to reconcile each transaction, the licensee shall file a report for missed transactions or file a corrected report for previously submitted reports containing inaccuracies. By January 31 of each year, each licensee must submit to the National Source Tracking System confirmation that the data in the National Source Tracking System is correct.

(8) Each licensee that possesses Category 1 nationally tracked sources shall report its initial inventory of Category 1 nationally tracked sources to the National Source Tracking System by January 31, 2009 or as specified in 10 C.F.R. section 20.2207(h), 1-1-13 edition, which is herein incorporated by reference and is available at http://www.flrules.org/Gateway/reference.asp?No=Ref-03456 or http://www.gpo.gov/fdsys/pkg/CFR-2013-title10-vol1/pdf/CFR-2013-title10-vol1-sec20-2207.pdf. Each licensee that possesses Category 2 nationally tracked sources shall report its initial inventory of Category 2 nationally tracked sources to the National Source Tracking System by January 31, 2009 or as specified in 10 C.F.R. section 20.2207(h), 1-1-13 edition. The information may be submitted by using any of the methods identified by paragraphs (6)(a) through (6)(e) of this section. The initial inventory report must include the following information:

(a) The name, address, and license number of the reporting licensee;

(b) The name of the individual preparing the report;

(c) The manufacturer, model, and serial number of each nationally tracked source or, if not available, other information to uniquely identify the source;

(d) The radioactive material in the sealed source;

(e) The initial or current source strength in becquerels (curies); and

(f) The date for which the source strength is reported.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051, 404.081 FS. History–New 2-28-08, Amended 12-26-13.

- 64E-5.351 Nationally Tracked Source Thresholds.

The nationally tracked source thresholds are listed in table 1 below with the Terabecquerel (TBq) values as the regulatory standard. The curie (Ci) values specified are obtained by converting from the TBq value. The curie values are provided for practical usefulness only and are rounded after conversion.

Table 1

| Radioactive Material | Category 1 (TBq) | Category 1 (Ci) | Category 2 (TBq) | Category 2 (Ci) |
| --- | --- | --- | --- | --- |
| Actinium-227 | 20 | 540 | 0 .2 | 5 .4 |
| Americium-241 | 60 | 1,600 | 0 .6 | 16 |
| Americium-241/Be | 60 | 1,600 | 0 .6 | 16 |
| Californium-252 | 20 | 540 | 0 .2 | 5 .4 |
| Cobalt-60 | 30 | 810 | 0 .3 | 8 .1 |
| Curium-244 | 50 | 1,400 | 0 .5 | 14 |
| Cesium-137 | 100 | 2,700 | 1 | 27 |
| Gadolinium-153 | 1,000 | 27,000 | 10 | 270 |
| Iridium-192 | 80 | 2,200 | 0 .8 | 22 |
| Plutonium-238 | 60 | 1,600 | 0 .6 | 16 |
| Plutonium-239/Be | 60 | 1,600 | 0 .6 | 16 |
| Polonium-210 | 60 | 1,600 | 0 .6 | 16 |
| Promethium-147 | 40,000 | 1,100,000 | 400 | 11,000 |
| Radium-226 | 40 | 1,100 | 0 .4 | 11 |
| Selenium-75 | 200 | 5,400 | 2 | 54 |
| Strontium-90 | 1,000 | 27,000 | 10 | 270 |
| Thorium-228 | 20 | 540 | 0 .2 | 5 .4 |
| Thorium-229 | 20 | 540 | 0 .2 | 5 .4 |
| Thulium-170 | 20,000 | 540,000 | 200 | 5,400 |
| Ytterbium-169 | 300 | 8,100 | 3 | 81 |

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051, 404.081 FS. History–New 2-28-08, Amended 12-26-13.

- 64E-5.423 Definitions.

As used in this part, the following definitions apply:

(1) “Associated equipment” means equipment, such as guide tubes, control tubes, control cables, removable source stops, J-tubes, and collimators, used in conjunction with a radiographic exposure device that drives, guides, or comes in contact with the sealed source.

(2) “Certifying entity” means:

(a) For radiographic operations using radioactive materials, an independent certifying organization that meets the requirements of Appendix A of 10 CFR Part 34, which is herein incorporated by reference and which is available from the department, or an agreement state that meets the requirements of Appendix A, Parts II and III of 10 CFR Part 34.

(b) For radiographic operations using radiation machines, any agreement state or organization approved by the Conference of Radiation Control Directors, Inc.

(3) “Collimator” means a radiation shield that is placed on the end of the guide tube or directly onto a radiographic exposure device to restrict the size of the radiation beam when the sealed source is cranked into position to make a radiographic exposure.

(4) “Control cable” means the cable that is connected to the source assembly and used to drive the source from and return it to the shielded position. It also is known as a drive cable.

(5) “Control drive mechanism” means a device that enables the source assembly to be moved to and from the shielded position. It also is known as a crank assembly.

(6) “Control tube” means a protective sheath for guiding the control cable. The control tube connects the control drive mechanism to the radiographic exposure device.

(7) “Exposure head” means a device that locates the sealed source in the selected position. It also is known as a source stop.

(8) “Guide tube” means a flexible or rigid tube for guiding the source assembly and the attached control cable from the radiographic exposure device to the exposure head and includes the connections to attach to the radiographic exposure device and to the exposure head. It also is known as a projection sheath or source tube.

(9) “Industrial cabinet x-ray system” means a cabinet x-ray system used to perform industrial radiography excluding baggage x-ray systems.

(10) “Lay-barge radiography” means industrial radiography performed on any water vessel used for laying pipe.

(11) “Platform radiography” means industrial radiography performed on an offshore platform or other structure over a body of water.

(12) “Radiographer certification” means a written document received from a certifying entity stating that an individual has met radiation safety training, testing, and experience criteria satisfactorily.

(13) “Radiographic operations” means all activities including surveys that involve the use or transport of radiation machines, radiographic exposure devices, source changers, or industrial cabinet x-ray systems to conduct industrial radiography.

(14) “Radiographic personnel” means radiographers and radiographer’s assistants.

(15) “Reference survey” means a survey made with a radiation survey instrument within 6 inches (15 cm) of the surface of a radiographic exposure device or source changer at a location established by the licensee. The reference survey is used to verify that the sealed source is located properly in the shielded position and to establish a radiation level for reference before, during, and after radiographic operations.

(16) “S-tube” means a tube through which the radioactive source travels inside a radiographic exposure device.

(17) “Source assembly” means a set of assembled parts consisting of a sealed source and a connector that attaches the source to the control cable. The source assembly sometimes includes a stop ball used to secure the source in the shielded position. It also is known as a pigtail.

(18) “Special training session” means training not conducted during production radiography.

(19) “Transport container” means a package that is designed to provide radiation safety and security when sealed sources are transported and that meets all applicable requirements of the U.S. Department of Transportation (USDOT).

(20) “Underwater radiography” means industrial radiography performed when the radiation machine, radiographic exposure device, or related equipment are beneath the surface of the water.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.031, 404.051(1), (4), (6) FS. History–New 9-11-01.

- 64E-5.424 Requirements for Industrial Radiography Equipment Using Sealed Sources.

(1) Equipment used in radiographic operations shall meet the criteria specified below.

(a) Each radiographic exposure device, source assembly or sealed source, and all associated equipment shall meet the requirements specified in American National Standards Institute (ANSI) N432-1980 “Radiological Safety for the Design and Construction of Apparatus for Gamma Radiography,’’ published as National Bureau of Standards Handbook 136, January 1981, which is herein incorporated by reference and which is available from the department. Engineering analyses that demonstrate that the radiography equipment components are equivalent are an acceptable alternative to actual testing of the component.

(b) Equipment used in radiographic operations is not required to comply with section 8.9.2(c) of the Endurance Test in ANSI N432-1980 if the prototype equipment has been tested using a torque value representative of the torque that an individual using the radiography equipment realistically can exert on the lever or crankshaft of the drive mechanism.

(2) In addition to the requirements specified in subsection 64E-5.424(1), F.A.C., radiographic exposure devices, source changers, source assemblies, and sealed sources must meet the requirements specified below.

(a) Each radiographic exposure device shall have a durable, legible, clearly visible label attached that specifies:

1. The chemical symbol and mass number of the radionuclide in the radiographic exposure device;

2. The activity of the sealed source and the date on which this activity was last measured;

3. The manufacturer’s name and the model and serial number of the sealed source; and

4. The name, address, and telephone number of the licensee.

(b) Each radiographic exposure device, source changer, storage container, and transport container shall have a durable, legible, clearly visible marking or label attached that includes the standard radiation symbol as specified in Rule 64E-5.322, F.A.C., in conventional colors of magenta, purple, or black on a yellow background, has a minimum diameter of 25 millimeters, and has the following wording:

CAUTION (or DANGER)

RADIOACTIVE MATERIAL – DO NOT HANDLE

NOTIFY CIVIL AUTHORITIES (or NAME OF COMPANY)

(c) Modification of radiographic exposure devices, source changers, source assemblies, and associated equipment is prohibited unless the design of any replacement component, including source holder, source assembly, controls, or guide tubes will not compromise design safety features.

(3) Radiographic exposure devices, source assemblies, and associated equipment that allow the source to be moved out of the radiographic exposure device for radiographic operations or to source changers must meet the requirements specified below.

(a) The coupling between the source assembly and the control cable shall be designed so that the source assembly will not become disconnected if cranked outside the guide tube. The coupling shall be designed so that it cannot be disconnected unintentionally under normal and reasonably foreseeable abnormal conditions.

(b) The radiographic exposure device shall secure the source assembly automatically when it is cranked back into the fully shielded position within the device. This securing system shall be able to be released only by a deliberate operation on the exposure device.

(c) The outlet fittings, lock box, and drive cable fittings on each radiographic exposure device shall be equipped with safety plugs or covers that are installed during storage and transportation to protect the source assembly from water, mud, sand, or other foreign matter.

(d)1. Each sealed source or source assembly shall have attached to it or engraved on it a durable, legible, visible label with the words: “DANGER – RADIOACTIVE.”

2. The label cannot interfere with the safe operation of the radiographic exposure device, source changer, or associated equipment.

(e) The guide tube shall be able to withstand a crushing test that approximates closely the crushing forces that are likely to be encountered during use and be able to withstand a kinking resistance test that approximates closely the kinking forces that are likely to be encountered during use.

(f) Guide tubes shall be used when moving the source out of the device.

(g) An exposure head or similar device designed to prevent the source assembly from passing out of the end of the guide tube shall be attached to the outermost end of the guide tube during radiographic operations.

(h) The guide tube exposure head connection shall be able to withstand the tensile test for control units specified in ANSI N432-1980.

(i) Source changers shall have a system to ensure that the source will not be withdrawn from the changer accidentally when connecting or disconnecting the drive cable to or from a source assembly.

(4) The maximum exposure rate limits for storage containers and source changers are 200 millirem (2 mSv) per hour at any exterior surface and 10 millirem (0.1 mSv) per hour at 1 meter from any exterior surface with the sealed source in the shielded position.

(5) Each radiographic exposure device, source changer, and storage container shall have a lock or outer locked container designed to prevent unauthorized or accidental removal of the sealed source from its shielded position.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), (6) FS. History–New 9-11-01.

- 64E-5.425 Locking of Sources of Radiation, Storage Precautions, and Surveillance.

(1) Each radiation machine, radiographic exposure device, source changer, and storage container shall be kept locked with the key removed from any keyed lock except when under the direct supervision of radiographic personnel or as specified in subsection (6), below.

(2) Each radiation machine, radiographic exposure device, source changer, and storage container shall be locked and the key removed from any keyed lock before being moved or transported and before being stored at a given location, except at permanent radiographic installations as specified in Rule 64E-5.431, F.A.C. Keys to radiation machines, radiographic exposure devices, source changers, storage containers, transport containers, and transport vehicles shall be maintained in the possession of the radiographer or radiographer’s assistant responsible for the equipment in a manner that prevents access to sources of radiation by unauthorized personnel.

(3) Locked radiographic exposure devices, source changers, storage containers, and radiation machines shall be secured physically except when under the direct surveillance of radiographic personnel or as specified in subsection (6), below, to prevent tampering or removal by unauthorized personnel. The licensee shall store licensed material in a manner that minimizes danger from explosion or fire.

(4) Each sealed source shall be secured in its shielded position by locking the radiographic exposure device or source changer each time the sealed source is returned to the shielded position.

(5) Transport containers containing licensed material shall be locked and secured in the transporting vehicle to prevent accidental loss, tampering, or unauthorized removal of the licensed material from the vehicle.

(6) During each radiographic operation, the radiographer or radiographer’s assistant shall maintain continuous direct visual surveillance of the operation to protect against unauthorized entry into a high radiation area, except at permanent radiographic installations where all entryways are locked and the requirements of Rule 64E-5.431, F.A.C., are met.

(7) During each radiographic operation using an industrial cabinet x-ray system, direct surveillance of the operation shall be maintained to protect against unauthorized entry into a high radiation area.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), (6) FS. History–New 9-11-01.

- 64E-5.426 Radiation Survey Instruments.

(1) The licensee or registrant shall maintain enough calibrated and operable radiation survey instruments to make physical radiation surveys as required by the rules contained in this part and Chapter 64E-5, Part III, F.A.C. Such instrumentation shall be able to measure a range from 2 millirem (0.02 mSv) per hour through 1 rem (0.01 Sv) per hour.

(2) Radiation survey instruments used to establish dose rates shall be calibrated:

(a) At intervals not to exceed 6 months and after each instrument servicing other than battery replacement;

(b) At energies and geometries appropriate for use;

(c) To demonstrate accuracy within 20% of the true radiation level at each point checked;

(d) For linear scale instruments, at two points located approximately 1/3 and 2/3 of full-scale on each scale; for logarithmic scale instruments, at midrange of each decade and at two points at least one decade apart; and for digital instruments, at three points between 2 millirem (0.02 mSv) per hour and 1 rem (0.01 Sv) per hour; and

(e) By a person licensed by the department, another agreement state, licensing state or the NRC.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 9-11-01.

- 64E-5.427 Leak Testing, Repairing, Tagging, Opening, Modifying, and Replacing Sealed Sources and Devices.

(1) The replacement, leak testing, leak test sample analysis, repair, tagging, opening, or any other modification of any sealed source shall be performed only by persons authorized specifically to do so by the department, another agreement state, licensing state, or the NRC.

(2) Each sealed source shall be tested for radioactive contamination leakage at intervals not to exceed 6 months. In the absence of a certificate from a transferor indicating that a test has been made within the 6 months before the transfer, the sealed source shall not be used until tested. Sealed sources that are listed in a department license for storage only do not require leak testing during storage but shall be tested before use or transfer to another person if the interval of storage exceeds 6 months.

(3) Each exposure device using depleted uranium (DU) shielding and an S-tube configuration shall be tested for DU contamination at intervals not to exceed 12 months. DU shielded devices do not have to be tested for DU contamination while in storage and not in use. However, the DU devices shall be tested for DU contamination before use or transfer if the interval of storage exceeds 12 months. Licensees must comply with the DU leak testing requirements of this section within 6 months after the effective date of this rule.

(4) Leak testing as specified in subsections 64E-5.427(2) and (3), F.A.C., shall be capable of detecting the presence of 0.005 microcurie (185 Bq) of removable contamination on the test sample. The wipe sample shall be taken from the nearest accessible point to the sealed source where contamination could accumulate.

(5) If any test conducted pursuant to this section reveals the presence of 0.005 microcurie (185 Bq) or more of removable radioactive material, the licensee immediately shall withdraw the equipment from use and cause it to be decontaminated and repaired or disposed of in accordance with Rule 64E-5.1303, F.A.C., and the applicable sections of rules contained in Parts III and XV of Chapter 64E-5, F.A.C. If DU leak testing reveals the presence of 0.005 microcurie (185 Bq) or more of removable DU contamination, the exposure device shall be removed from use until an evaluation of the wear on the S-tube has been made. If the evaluation reveals that the S-tube is worn through, the device shall not be used. The licensee shall file a report with the department describing the equipment involved, the test results, and the corrective action taken within 5 days after obtaining results of the test.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), (6) FS. History–New 9-11-01, Amended 9-28-06.

- 64E-5.428 Quarterly Inventory.

Each licensee or registrant shall conduct a quarterly physical inventory to account for all sources of radiation received or possessed during the quarter. The inventory shall cover all sources of radiation requiring licensure or registration by the department, including sealed sources, radiation machines, radiographic exposure devices, and source changers containing DU.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1) FS. History–New 9-11-01.

- 64E-5.429 Source Movement Logs, Daily Survey Reports, and Individual Dosimeter Logs.

(1) Each time a radiation source is removed from storage, the licensee or registrant shall complete and maintain source movement logs for each radiation source with the following information, as applicable:

(a) The locations where used, the names of the jobs or clients, and the dates of use including the dates removed and returned to storage;

(b) The manufacturer’s name, model, and serial number of the radiographic exposure device, source changer, or radiation machine used;

(c) The sealed source manufacturer’s name, model, and serial number, activity in curies (becquerels) on the date of receipt and each date of use, and the due date of the next leak test;

(d) The results of the reference survey of the radiographic exposure device or source changer performed upon removal and return to storage; and

(e) The identity and signature or initials of the radiographer to whom the radiation source has been assigned.

(2) Before performing industrial radiography, leak tests, source exchanges, or quarterly inspection and maintenance of radiographic equipment, the licensee or registrant shall prepare and maintain a daily survey report for each radiation source with the information described below as it becomes available:

(a) The location where used, the name of the job or client, and the date of use;

(b) The manufacturer’s name, model, and serial number of the radiographic exposure device, source changer, or radiation machine used;

(c) The sealed source manufacturer’s name, model, and serial number and activity in curies (becquerels) for the date of use;

(d) The names and titles of the radiographic personnel working with the radiation source;

(e) The serial number of the personnel monitoring badge, pocket dosimeter, and alarm ratemeter used by each of the radiography crew members;

(f) The manufacturer’s name, model, serial number, and date of calibration or calibration due date for each survey meter used;

(g) The results of the reference survey performed when the radiographic exposure device or source changer is removed from or returned to storage;

(h) Evidence of performance of the equipment checks described in subsection 64E-5.430(1), F.A.C.;

(i) The results of the survey of the posted perimeter in mR/hr (mSv/hr) and feet (meters);

(j) The total exposure time; and

(k) The start, end, and total pocket dosimeter readings for all radiographic personnel.

(3) Radiographic personnel shall maintain an individual log of their daily dosimeter totals. Each individual shall record the doses measured by his or her dosimeter at the end of each day of radiographic operations and total the recorded doses at the end of each week and at the end of each month. Copies of the individual dosimeter logs shall be provided to the radiation safety officer (RSO) or the RSO’s designee no later than 7 days after each month. The RSO or the RSO’s designee shall review the logs within 7 days of receipt and shall date and sign or initial the logs at the time of the review. Each log shall include the following information:

(a) The name of the individual;

(b) The dates of the monitoring periods;

(c) The daily, weekly, and monthly individual radiation dose totals as measured by the dosimeter; and

(d) The date the log was reviewed by the RSO or the RSO’s designee and the signature or initials of the RSO or the RSO’s designee.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1) FS. History–New 9-11-01, Amended 9-28-06.

- 64E-5.430 Inspection and Maintenance.

(1) Each licensee or registrant shall perform visual and operability checks on survey instruments, radiation machines, radiographic exposure devices, associated equipment, transport containers, storage containers, and source changers before use on each day the equipment is to be used to ensure the equipment is in good working condition, the sources are shielded adequately, and required labeling is present. All appropriate parts shall be maintained in accordance with the manufacturer’s specifications. Each radiation survey instrument shall be visually inspected, have its batteries checked, and have its operability checked with a radiation source at the beginning of each day of use and at the beginning of each work shift. If equipment problems are found, the equipment shall be removed from service until repaired.

(2) Each licensee or registrant shall have written procedures and perform equipment inspection and maintenance as described below. (Pursuant to Section 120.54(6), F.S., subsection 64E-5.430(2), F.A.C., is substantively identical to 10 CFR 34.31(b) published on 01/01/2007.)

(a) Inspection and maintenance of survey instruments, radiation machines, radiographic exposure devices, associated equipment, source changers, storage containers, and transport containers shall be performed quarterly to assure proper functioning of components important to safety. All appropriate parts shall be maintained in accordance with the manufacturer’s specifications. Verification of compliance with radiation limits specified in subsection 64E-5.424(4), F.A.C., shall be included in each quarterly inspection. If equipment problems are found, the equipment shall be labeled as defective and removed from service until repaired. Replacement components shall meet manufacturer’s specifications.

(b) Inspection and maintenance of Type B packages used to transport radioactive materials shall be performed quarterly in accordance with each package’s certificate of compliance or other approval.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1) FS. History–New 9-11-01, Amended 2-28-08.

- 64E-5.431 Permanent Radiographic Installations.

(1) Each entrance used for personnel access to a high radiation area in a permanent radiographic installation shall have either:

(a) An entrance control that reduces the radiation level to below the level at which an individual might receive a deep dose equivalent of 0.1 rem (1 millisievert) in 1 hour at 30 centimeters from the source of radiation from any surface the radiation penetrates, or

(b) Conspicuous visible and audible signals to warn of the presence of radiation. The visible signal shall be actuated by radiation. The audible signal shall be actuated when an attempt is made to enter the installation while the source is exposed or the radiation machine is activated.

(2) The alarm system shall be tested for proper operation with a radiation source each day before radiographic operations. The test shall include a check of both the visible and audible signals. Entrance control devices that reduce the radiation level upon entry shall be tested monthly. If an entrance control device or an alarm is operating improperly, it shall be labeled immediately as defective and repaired within 7 days. The installation can continue to be used by an unaccompanied radiographer during this 7-day period if the continuous surveillance requirements of subsection 64E-5.425(6), F.A.C., are implemented and an alarming ratemeter is used.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4) FS. History–New 9-11-01.

- 64E-5.432 Radiation Protection Program.

The radiation protection program specified in Rule 64E-5.303, F.A.C., for registrants performing radiography and license applications, renewals, and requests for amendments for licensees performing radiography shall include the components specified below and the location of all records required.

(1) A description of the overall organizational structure pertaining to the licensee’s or registrant’s radiation protection program, including specific delegation of authority and responsibility, the name of the RSO, and the minimum qualifications of the RSO and the RSO’s designees.

(2) A radiation safety training program for radiographic personnel that meets the requirements of Rule 64E-5.434, F.A.C., and includes the components described below.

(a) Initial, periodic, and on-the-job training.

(b) Written and practical examinations to determine knowledge, understanding of, and ability to comply with department and applicable USDOT rules, licensee or registrant requirements, operating and emergency procedures, and use of radiographic and related equipment.

(3) Procedures to verify the certification of radiographers and to ensure that the certification remains valid.

(4) A written policy to maintain radiation doses as low as reasonably achievable as specified in Rule 64E-5.303, F.A.C. The policy shall include:

(a) A commitment by management to keep radiation doses as low as reasonably achievable and a description of the participation of management, the RSO, and radiographic personnel in the implementation of the policy;

(b) Investigation within 30 days by the RSO of any exposure level that exceeds established monthly and quarterly levels and implementation of corrective actions to halt unnecessary exposures and prevent recurrence; and

(c) An audit of the program to evaluate its effectiveness in minimizing exposures in conjunction with the annual review of the radiation protection program specified in subsection 64E-5.303(3), F.A.C. A summary of the results of each audit, including a description of corrective actions taken, shall be prepared by the RSO and approved by the licensee or registrant.

(5) An auditing program for internal inspections of the job performance of all radiographic personnel at intervals not to exceed 6 months as described in Rule 64E-5.434, F.A.C.

(6) Written operating and emergency procedures as described in Rule 64E-5.436, F.A.C.

(7) Leak testing procedures, including a description of:

(a) The method of taking wipes and preparing samples for analysis using only radiographers or radiographer’s assistants working under the personal supervision of a radiographer or persons specifically licensed by the department, another agreement state, licensing state, or the NRC to perform such services; and

(b) The method of performing leak test sample analyses, including instrumentation to be used and experience of the individuals who will perform the analyses or a commitment to use vendors specifically licensed to perform such analyses by the department, another agreement state, licensing state, or the NRC.

(8) Procedures for the semiannual calibration of survey instruments and the annual calibration of alarm ratemeters, including a description of the calibration instrumentation and the experience of the person who will perform the calibrations or a commitment to use persons specifically licensed to perform such calibrations by the department, another agreement state, licensing state, or the NRC. All survey instrument calibrations shall be performed in accordance with subsection 64E-5.426(2), F.A.C.

(9) Procedures for quarterly inspection and maintenance of survey instruments, radiation machines, radiographic exposure devices, associated equipment, source changers, storage containers, and transport containers to assure proper function of components important to safety, performed in accordance with Rule 64E-5.430, F.A.C.

(10) Procedures for annual calibration of pocket or electronic dosimeters, including a description of the calibration instrumentation and the experience of the person who will perform the calibrations or a commitment to use persons specifically licensed to perform such calibrations by the department, another agreement state, licensing state, or the NRC.

(11) Procedures for lay-barge, offshore platform and underwater radiography if conducting such activities.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (4), (6), (9), (10), 404.061(2), 404.081(1), 404.141 FS. History– New 9-11-01.

- 64E-5.433 Radiation Safety Officer.

(1) The licensee or registrant shall appoint an RSO and delegate the authority needed to fulfill the duties of the position. Except as specified in subsection 64E-5.433(2), F.A.C., below, the minimum qualifications, training, and experience for the RSO shall be:

(a) One year of documented industrial radiography experience as a radiographer; and

(b) Sixteen hours of formal instruction in the establishment and maintenance of a radiation protection program, including training to perform internal audits and mitigation of radiological incidents. Individuals identified as an RSO on an industrial radiography license or registration before the effective date of this rule are not required to comply with the training requirements of this paragraph.

(2) Equivalent alternative radiation and safety training and experience in radiographic operations and formal training in the establishment and maintenance of a radiation protection program can substitute for the requirements specified in paragraphs 64E-5.433(1)(a), (b), F.A.C., above.

(3) In addition to other duties specified in this part, the RSO shall:

(a) Ensure compliance with all components of the licensee’s or registrant’s radiation protection program as specified in Rule 64E-5.432, F.A.C., the terms and conditions of the license, and this rule;

(b) Investigate incidents and direct corrective actions, including halting operations when necessary;

(c) Serve as the licensee’s or registrant’s contact with the department; and

(d) Ensure that radiation safety activities are performed using approved procedures and requirements in Chapter 64E-5, F.A.C., in the daily operation of the licensee’s program.

Rulemaking Authority 404.051, 404.061, FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2) FS. History–New 9-11-01.

- 64E-5.434 Training, Testing, Certification, and Audits.

(1) The licensee or registrant shall not permit any individual to act as a radiographer’s assistant until such individual:

(a) Receives a copy of the licensee’s or registrant’s operating and emergency procedures;

(b) Completes 8 hours of training, including instruction in the licensee’s or registrant’s operating and emergency procedures and supervised instruction in use of the licensee’s or registrant’s radiographic equipment, related handling tools, radiation survey instruments, and personnel monitoring devices during a special training session; and

(c) Successfully completes a closed-book, written examination on the licensee’s or registrant’s operating and emergency procedures and a practical examination that is not conducted during production radiography to demonstrate competence in the use of the licensee’s or registrant’s radiographic equipment, related handling tools, radiation survey instruments, and personnel monitoring devices.

(2) Licensees and registrants can allow individuals who have completed the training and testing specified in paragraphs 64E-5.434(2)(a)-(d), F.A.C., below, to perform industrial radiography for 12 months after the effective date of these rules. The licensee or registrant shall not permit any individual to act as a radiographer until such individual:

(a) Receives copies of rules contained in Chapter 64E-5, Parts I-IV, IX and XV, F.A.C., applicable USDOT regulations, the appropriate license or certificate of registration, and the licensee’s or registrant’s operating and emergency procedures;

(b)1. For radioactive material radiographic operations, completes 320 hours of on-the-job training in industrial radiography, excluding hours as specified in subparagraph 64E-5.434(2)(b)2., F.A.C., below, as a radiographer’s assistant using radioactive material; or

2. For machine produced radiographic operations, completes 200 hours of on-the-job training using radiation machines;

(c) Receives 40 hours of formal instruction in the subjects outlined in subsection 64E-5.434(6), F.A.C., and supervised instruction during a special training session in the inspection and use of the licensee’s or registrant’s radiographic equipment, related handling tools, radiation survey instruments, and personnel monitoring devices;

(d) Successfully completes a closed-book, written examination on the subjects outlined in subsection 64E-5.434(6), F.A.C., and a practical examination to demonstrate competence in the use of the licensee’s or registrant’s radiographic and safety equipment; and

(e) Is certified by a certifying entity.

(3) Radiographers who work for an out-of-state radioactive materials license under reciprocal recognition are authorized to conduct radiographic operations within the state if they have a valid certification from a certifying entity for the activities being conducted before entering the state.

(4) Any individual who has completed all requirements specified in subsection 64E-5.434(2), F.A.C., above, and begins work for a different Florida licensee or registrant shall complete 4 hours of additional training and testing before conducting radiographic operations. The training shall consist of instructions in the licensee’s or registrant’s operating and emergency procedures and supervised instruction during a special training session in the use of the licensee’s or registrant’s radiographic and safety equipment. The testing shall consist of successful completion of the written and practical examinations described in paragraph 64E-5.434(1)(c), F.A.C. The RSO shall document how the prior radiation training and experience was verified.

(5) Personnel using industrial cabinet x-ray systems for industrial radiography shall complete 16 hours of training and testing as described below:

(a) Ten hours of training and testing as described in subsection 64E-5.434(6), F.A.C.; and

(b) Two hours of instruction in the registrant’s operating and emergency procedures pertaining to industrial radiography using industrial cabinet x-ray systems, 2 hours of supervised instruction during a special training session in the use of the registrant’s industrial cabinet x-ray system, related handling tools, radiation survey instruments, and personnel monitoring devices, and 2 hours of testing, which shall consist of a written examination covering operating and emergency procedures and equipment use and a practical examination to demonstrate competence in the use of the registrant’s industrial cabinet x-ray system and related equipment.

(6) The subjects to be covered during the instruction of radiographers shall include:

(a) Fundamentals of radiation safety, including characteristics of radiation, units of radiation dose, quantities of radioactivity, hazards of radiation exposure, radiation protection standards, radiation levels from sources of radiation, and methods of minimizing radiation dose.

(b) Radiation detection instruments, including:

1. Use, operation, calibration, and limitations of radiation survey instruments;

2. Survey techniques; and

3. Use of personnel monitoring equipment.

(c) Equipment to be used, including, as applicable:

1. Operation and control of radiation machines, radiographic exposure equipment, remote handling equipment, source changers, storage containers, and transport containers, including pictures or models of source assemblies;

2. Storage, control, and disposal of licensed material; and

3. Inspection and maintenance of equipment.

(d) The applicable requirements of these rules and NRC and USDOT regulations.

(e) The licensee’s or registrant’s operating and emergency procedures.

(f) Case histories of industrial radiography accidents.

(7) Each licensee or registrant shall provide 8 hours of refresher annual radiation safety training to all radiographic personnel, which can be conducted in multiple sessions.

(8) The RSO or the RSO’s designee shall audit the job performance of each radiographer and radiographer’s assistant to ensure that the department’s regulations, license requirements, and the licensee’s or registrant’s operating and emergency procedures are followed. The audits shall include observation of the performance of each radiographer or radiographer’s assistant during an actual radiographic operation at intervals not to exceed 6 months. Radiographers or radiographer’s assistants who have not participated in a radiographic operation for more than 6 months since the last audit shall demonstrate knowledge of the licensee’s or registrant’s operating and emergency procedures and safe use of radiographic and related equipment by a practical examination before participating in a radiographic operation. Audits of the RSO are not required.

(9) Individuals conducting internal radiation safety training or audits shall meet the minimum qualifications specified in subsection 64E-5.433(1), F.A.C., for the RSO.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2) FS. History–New 9-11-01, Amended 9-28-06.

- 64E-5.435 Conducting Industrial Radiographic Operations.

(1) With the exception of industrial cabinet x-ray systems, the radiographer shall be accompanied by at least one other radiographer or radiographer’s assistant whenever radiography is performed at a location other than a permanent radiographic installation. The additional qualified individual shall observe the radiographic operations and be capable of providing immediate assistance to prevent unauthorized entry. Radiography is prohibited if only one qualified individual is present. Radiography performed in an industrial cabinet x-ray system by a single individual meeting the training and testing requirements specified in subsection 64E-5.434(5), F.A.C., is permitted.

(2) The radiographer’s assistant shall be under the personal supervision of a radiographer when using a radiation machine, radiographic exposure device, source changer, or related source handling tools or conducting radiation surveys to determine that the sealed source has returned to the shielded position or that the radiation machine is off after an exposure.

(3) All radiographic operations conducted at a licensee’s or registrant’s permanent facility shall be conducted in a permanent radiographic installation or an industrial cabinet x-ray system or using equipment, facilities, and procedures that are adequate to protect public health, safety, and property and included in the radiation protection program specified in Rule 64E-5.432, F.A.C.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), (11), 404.081(1) FS. History–New 9-11-01.

- 64E-5.436 Operating and Emergency Procedures.

The licensee’s or registrant’s procedures shall include instructions in the following:

(1) Handling and use of sources of radiation to be used so that exposures are maintained as low as reasonably achievable and no individual is likely to be exposed to radiation doses in excess of the limits established in rules contained in Part III of Chapter 64E-5, F.A.C.;

(2) Methods and occasions to conduct radiation surveys;

(3) Methods to control access to radiographic areas;

(4) Methods and occasions to lock and secure sources of radiation;

(5) Personnel monitoring and the use of personnel monitoring equipment, including steps to be taken immediately by radiography personnel when a pocket dosimeter is found off-scale, an alarm ratemeter alarms unexpectedly, or a personnel monitoring badge is damaged or lost;

(6) Transportation of licensed material to field locations and preparation of packages for shipment by common or contract carriers, including packaging, marking, labeling, shipping papers, emergency response information, blocking and bracing, security, surveys, and vehicle placarding in accordance with applicable requirements of the USDOT;

(7) Leak testing, quarterly inventories, and equipment inspection, maintenance and operability checks, and disposal of licensed material;

(8) Source exchanges for licensees who perform source exchanges;

(9) Calibration of survey instruments, dosimeters, and alarm ratemeters for licensees who perform calibrations;

(10) Emergency response, including response to loss, damage, or theft of sources of radiation, unauthorized entries into restricted areas, notifications, exposure minimization, and source recovery;

(11) Identifying and reporting equipment defects and noncompliance issues; and

(12) Maintenance of records.

Rulemaking Authority 404.051, 404.20 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.081, 404.20(1) FS. History–New 9-11-01.

- 64E-5.437 Personnel Monitoring.

(1) The licensee or registrant shall not permit any individual to act as a radiographer or a radiographer’s assistant unless the individual wears on the trunk of his or her body at all times during radiographic operations:

(a) A NVLAP-approved personnel monitoring badge such as a film badge, thermoluminescent dosimeter (TLD) or optically stimulated luminescent device (OSLD);

(b) A direct reading pocket dosimeter, which can be either an ion chamber or electronic personal dosimeter; and

(c) An alarming ratemeter. Alarm ratemeters are not required for radiography performed in an approved permanent radiographic installation meeting the requirements of Rule 64E-5.431, F.A.C.

(2) Each personnel monitoring badge shall be assigned to and worn by only one individual and shall be exchanged monthly. After exchange each badge shall be processed as soon as possible. If a report is received from the badge processor that indicates an individual has received a radiation exposure in excess of 5 rem (0.05 Sv), the licensee or registrant shall notify the department within 24 hours as specified in subsection 64E-5.344(2), F.A.C. If a personnel monitoring badge is lost or damaged, the worker shall cease work immediately until a replacement badge is provided and the exposure is calculated by the RSO or the RSO’s designee for the time period from issuance to loss or damage of the badge. The results of the calculated exposure and the time period for which the personnel monitoring badge was lost or damaged shall be provided to the processor to adjust the individual’s occupational exposure record.

(3) Pocket dosimeters shall have a range from 0 to 200 millirem (2 mSv) and shall be recharged at the start of each shift and when 75% of the full scale of the dosimeter is exceeded. Initial, final, and total pocket dosimeter readings shall be recorded at the start and end of each shift.

(4) If an individual’s pocket dosimeter is found to be off-scale or if an individual’s electronic personal dosimeter reads more than 200 millirem (2 mSv) and the possibility of radiation exposure cannot be ruled out as the cause, the individual’s personnel monitoring badge shall be sent for processing within 24 hours. In addition, the individual shall not resume radiographic operations until a determination of the individual’s radiation exposure has been made by the RSO or the RSO’s designee. The results of this determination shall be reported in writing to the department within 30 days of the determination.

(5) Each alarming ratemeter shall:

(a) Have a function test without being exposed to radiation to ensure that the audible alarm is functioning properly before use at the start of each work shift;

(b) Give an alarm at a preset dose rate of no more than 500 millirem (0.5 mSv) per hour; and

(c) Require special means to change the preset alarm function.

(6) Pocket dosimeters and alarm ratemeters shall be calibrated annually for correct response to radiation by a person licensed by the department, another agreement state, licensing state, or the NRC. Acceptable dosimeters shall read within 20% of the true radiation exposure. Ion chamber dosimeters also shall be checked for response to drift by setting the dosimeter at zero and storing it in a low background area for at least 24 hours and for electrical leakage, which shall be no more than 1% of full scale for each 24 hours. Acceptable ratemeters shall alarm within 20% of the true radiation dose rate.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1), (2) FS. History–New 9-11-01.

- 64E-5.438 Radiation Surveys.

(1) No radiographic operations shall be conducted unless at least one calibrated and operable radiation survey instrument meeting the requirements of Rule 64E-5.426, F.A.C., is available for each radiographic exposure device and radiation machine in use at each site where radiographic exposures are made. All radiation surveys shall be performed with a calibrated and operable radiation survey instrument meeting the requirements of Rule 64E-5.426, F.A.C.

(2) The surveys described below shall be performed by the licensee or registrant where applicable.

(a) A reference survey of each radiographic exposure device or source changer immediately following removal from a storage area, including removal from storage following transportation.

(b) An area survey during the first radiographic exposure to verify that the posting requirements specified in subsection 64E-5.439(1), F.A.C., have been met and that unrestricted areas do not have radiation levels in excess of the limits specified in paragraph 64E-5.312(1)(c), F.A.C.

(c) A survey of the radiographic exposure device and the length of the guide tube after each exposure when approaching the device or guide tube, concluding with a reference survey of the radiographic exposure device at the location established by the licensee after each radiographic exposure. The surveys shall be performed before exchanging film, repositioning the exposure head, or dismantling equipment.

(d) A reference survey of the radiographic exposure device and source changer before and after source exchanges.

(e) A reference survey of the radiographic exposure device, source changer, or storage container after returning the sealed source to a storage area.

(f) A survey after each radiographic exposure using radiation machines to verify that the machine is off.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1) FS. History–New 9-11-01.

- 64E-5.439 Posting.

In addition to the posting requirements specified in Rule 64E-5.901, F.A.C., the licensee or registrant shall comply with the requirements described below.

(1) Radiation areas and high radiation areas created by radiographic operations shall be posted conspicuously as specified in subsections 64E-5.323(1) and (2), F.A.C. Areas or rooms in which licensed material is used or stored shall be posted as specified in subsection 64E-5.323(5), F.A.C. The exceptions to posting specified in subsection 64E-5.324(1), F.A.C., do not apply to industrial radiography.

(2) Source movement logs specified in Rule 64E-5.429, F.A.C., that document the current location of each source of radiation and source movements for the previous 30 days shall be posted conspicuously adjacent to the area where the source of radiation is stored.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1), 404.20 FS. History–New 9-11-01.

- 64E-5.440 Records.

(1) Each licensee or registrant shall maintain the following records for 3 years after the event at the location specified in Rule 64E-5.432, F.A.C., for inspection by the Department:

(a) Survey instrument, dosimeter, and alarm ratemeter calibrations specified in Rule 64E-5.426 and subsections 64E-5.437(5)-(6), F.A.C.;

(b) Leak test results specified in Rule 64E-5.427, F.A.C., which shall contain the manufacturer’s name, model, and serial number of each sealed source or device tested, including the device the source was stored in, the identity of each radionuclide, the estimated activity of each sealed source, the measured activity of each test sample expressed in microcuries (becquerels), the date of the test, and the signature or initials of the RSO or the RSO’s designee;

(c) Quarterly inventories specified in Rule 64E-5.428, F.A.C., which shall include the name of the person conducting the inventory, the radionuclide, number of curies (becquerels) or mass in each device, location of each sealed source, device, and machine, the manufacturer, model, and serial number of each sealed source, device, and machine, the date of the inventory, and the signature or initials of the RSO or the RSO’s designee;

(d) Source movement logs and daily survey reports specified in Rule 64E-5.429, F.A.C.

(e) Quarterly equipment inspection and maintenance specified in subsection 64E-5.430(2), F.A.C., including the date of the inspection, the name of inspector, the equipment involved, any problems found, and what repair or maintenance was done;

(f) Operation tests on permanent radiographic installation entrance controls and audible and visual alarms specified in Rule 64E-5.431, F.A.C.;

(g) Records of internal audits specified in subsection 64E-5.434(8), F.A.C., including lists of audit items checked and any violations observed;

(h) Records showing receipts and transfers of sealed sources and devices using DU for shielding, including the date, the name of the individual making the record, radionuclide, number of curies (becquerels) or mass, manufacturer, model, and serial number of each sealed source and device, as appropriate;

(i) Records of annual ALARA audits specified in paragraph 64E-5.432(4)(c), F.A.C.

(2) Each licensee or registrant shall maintain the following records until the Department terminates the license or registration requiring the record:

(a) Individual dosimeter logs specified in Rule 64E-5.429, F.A.C.;

(b) Initial and refresher radiation safety training specified in Rule 64E-5.434, F.A.C., including lists of the topics discussed, dates the training was conducted, names of the instructors and attendees, and written and practical examinations;

(c) Verification of previous radiography experience;

(d) Radiographer certification documents specified in paragraph 64E-5.434(2)(d), F.A.C., and verification of certification status;

(e) Records of personnel exposure investigations specified in paragraph 64E-5.432(4)(b), F.A.C., including the names of the individuals involved, the exposures received, the dates the exposures were received, a description of the cause of the exposures, the corrective actions taken, and the signature of the RSO;

(f) Records of estimates of exposures as a result of off-scale dosimeters or lost or damaged personnel monitoring badges, including records of surveys used to determine an individual’s exposure and reports submitted to the department as specified in subsection 64E-5.437(3), F.A.C.;

(g) Personnel monitoring badge records from the accredited NVLAP processor as specified in subsection 64E-5.437(2), F.A.C.; and

(h) Operating and emergency procedures. Licensees shall retain superseded material for 3 years after making changes to operating or emergency procedures.

(3) Each licensee or registrant conducting industrial radiography at a temporary job site shall have the following records available at that site for inspection by the department:

(a) Appropriate license or registration;

(b) Certification by a certifying entity;

(c) Operating and emergency procedures;

(d) Rules contained in Chapter 64E-5, Parts I-IV, IX, and XV, F.A.C.;

(e) Calibration records for the survey instruments, pocket dosimeters, and alarm ratemeters used at the site or calibration tags or labels that are affixed to the devices;

(f) Records of the latest leak test results for the specific devices in use at the site or leak test tags or labels that are affixed to the devices; and

(g) Source movement logs and daily survey reports for the period of operation at the site.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1), 404.20 FS. History–New 9-11-01, Amended 9-28-06, 2-28-08.

- 64E-5.441 Reporting Requirements.

(1) In addition to the reporting requirements specified in rules contained in Chapter 64E-5, Parts III and IX, F.A.C., and other sections of this part, each licensee shall provide a written report to the department within 30 days of the occurrence of any of the incidents involving radiographic equipment described below. Such reports shall be mailed to the Bureau of Radiation Control, Radioactive Materials Section, Bin C21, 4052 Bald Cypress Way, Tallahassee, Florida 32399-1741 for incidents involving radioactive materials or to the Bureau of Radiation Control, Radiation Machine Section, 705 Wells Road, Suite 300, Orange Park, Florida 32073 for incidents involving radiation machines.

(a) Unintentional disconnection of the source assembly from the control cable.

(b) Inability to retract and secure the source assembly to the fully shielded position.

(c) Failure of any component critical to safe operation of the device to perform its intended function properly.

(2) The licensee shall include the information described below in each report submitted as specified in this section.

(a) A description of the equipment problem.

(b) Cause of each incident if known.

(c) Manufacturer name and model number of the equipment involved in the incident.

(d) Place, time, and date of the incident.

(e) Actions taken to establish normal operations.

(f) Corrective actions taken or planned to prevent recurrence.

(g) Qualifications of the personnel involved in the incident.

(3) Reports of overexposures submitted as specified in rules contained in Part III of Chapter 64E-5, F.A.C., that involve failure of safety components of radiography equipment also must include the information specified in subsection 64E-5.441(2), F.A.C.

(4) Any licensee conducting radiographic operations or storing radioactive material at any location not listed on the license for a period in excess of 180 days in a calendar year, shall notify the Department prior to exceeding the 180 days. (Pursuant to Section 120.54(6), F.S., subsection 64E-5.441(4), F.A.C., is substantively identical to 10 CFR 34.101(c) published on 01/01/2007.)

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1) FS. History–New 9-11-01, Amended 9-28-06, 2-28-08.

- 64E-5.501 Definitions.

As used in this part, the following definitions apply:

(1) “Accessible surface” means the external surface of any enclosure or housing.

(2) “Added filtration” means any filtration which is in addition to the inherent filtration.

(3) “Aluminum equivalent” means the thickness of type 1100 aluminum alloy affording the same degree of radiation attenuation, under specified conditions, as the material in question. The nominal chemical composition of type 1100 aluminum alloy is 99.00 percent minimum aluminum, 0.12 percent copper.

(4) “Assembler” means any person engaged in the business of assembling, replacing or installing one or more components into an x-ray system or subsystem. The term includes the owner of an x-ray system or his employee or agent who assembles components into an x-ray system that is subsequently used to provide professional or commercial services.

(5) “Attenuation block” means a block or stack, having dimensions 20 centimeters by 20 centimeters by 3.8 centimeters, of type 1100 aluminum alloy or other materials having equivalent attenuation.

(6) “Automatic exposure control” means a device which automatically controls one or more technique factors in order to obtain, at a preselected location, a required quantity of radiation. See also “Phototimer”.

(7) “Barrier”. See “Protective barrier”.

(8) “Beam axis” means a line from the source through the centers of the x-ray field.

(9) “Beam-limiting device” means a device which provides a means to restrict the dimensions of the x-ray field, except for beam-blocking or beam-shaping devices used in radiation therapy.

(10) “Beam monitoring system” means a system designed to detect and measure the radiation present in the useful beam.

(11) “Cephalometric device” means a device intended for the radiographic visualization and measurement of the dimensions of the human head.

(12) “Certified components” means components of x-ray systems which are subject to regulations promulgated under Public Law 90-602, the Radiation Control for Health and Safety Act of 1968.

(13) “Certified system” means any x-ray system which has one or more certified components.

(14) “Changeable filters” means any filter which can be readily removed from the useful beam through any electronic, mechanical or physical process.

(15) “Contact therapy system” means an x-ray system used for therapy with the x-ray tube port placed in contact with or within 5 centimeters of the surface being treated.

(16) “Control panel” means that part of the x-ray control upon which are mounted the switches, knobs, pushbuttons and other hardware necessary for manually setting the technique factors and operating modes.

(17) “Cooling curve” means the graphical relationship between the heat units stored and cooling time.

(18) “Dead-man switch” means a switch so constructed that a circuit closing contact can be maintained only by continuous pressure on the switch by the operator.

(19) “Diagnostic source assembly” means the tube housing assembly with a beam-limiting device attached.

(20) “Diagnostic x-ray system” means an x-ray system designed for irradiation of any part of the human body as an aid to diagnosis through visualization of anatomical parts.

(21) “Direct scattered radiation” means that radiation which has been deviated in direction only by materials irradiated by the useful beam. See also “Scattered radiation”.

(22) “Entrance exposure rate” means the roentgens (C per kg) per unit time at the point where the center of the useful beam enters the patient.

(23) “Equipment”. See “X-ray equipment”.

(24) “Field emission equipment” means equipment which uses an x-ray tube in which electron emission from the cathode is due solely to the action of an electric field.

(25) “Filter” means material placed in the useful beam to preferentially absorb selected radiation.

(26) “Fluoroscopic imaging assembly” means a subsystem in which x-ray photons produce a fluoroscopic image. It includes the image receptors such as the image intensifier and spot film device, equipment housings, electrical interlocks, if any, the primary protective barrier, and structural material providing linkage between the image receptor and the diagnostic source assembly.

(27) “General purpose radiographic x-ray system” means any diagnostic radiographic x-ray system, except computed tomography systems, which, by design, is not limited to radiographic examination of a specific anatomical region.

(28) “Gonad shield” means a primary protective barrier for the testes or ovaries.

(29) “Half-value layer (HVL)” means the thickness of specified material which attenuates the beam of radiation to an extent such that the exposure rate is reduced to one-half of its original value. In this definition, the contribution of all scattered radiation, other than any which might be present initially in the beam concerned, is deemed to be excluded.

(30) “Healing arts self-referral” means the testing of human beings using x-ray machines for the detection or evaluation of health indications when such tests are not specifically and individually ordered by a licensed practitioner of the healing arts legally authorized to prescribe such x-ray tests for the purposes of diagnosis or medical treatment.

(31) “Heat unit” means a unit of energy equal to the product of the peak kilovoltage, milliamperes, and seconds, or kVp × mA × seconds.

(32) “Image intensifier” means a device, when installed in its housing, that instantaneously converts an x-ray pattern into a corresponding light image of higher energy density.

(33) “Image receptor” means any device, such as a fluorescent screen or radiographic film, which transforms incident x-ray photons either into a visible image or into another form which can be made into a visible image by further electronic or chemical transformations.

(34) “Image receptor support” means that part of a mammographic system designed to support the image receptor in a horizontal plane during the mammographic examination.

(35) “Inherent filtration” means the filtration of the useful beam provided by the permanently installed components of the x-ray tube housing assembly.

(36) “Irradiation” means the exposure of matter to ionizing radiation.

(37) “Kilovolts peak (kVp)”. See “Peak tube potential”.

(38) “kV” means kilovolts.

(39) “kWs” means kilowatt second. It is equivalent to kV × mA × seconds × 10-3.

(40) “Lead equivalent” means the thickness of lead affording the same attenuation, under specified conditions, as the material in question.

(41) “Leakage radiation” means radiation emanating from the diagnostic or therapeutic source assembly, except for the useful beam and radiation produced when the exposure switch or timer is not activated.

(42) “Leakage technique factors” means the technique factors associated with the diagnostic or therapeutic source assembly which are used in measuring leakage radiation. They are defined as follows:

(a) For diagnostic source assemblies intended for capacitor energy storage equipment, the maximum rated peak tube potential and the maximum rated number of exposures in an hour for operation at the maximum rated peak tube potential with the quantity of charge per exposure being 10 milliampere seconds (10 mAs) or the minimum obtainable from the unit, whichever is larger.

(b) For diagnostic source assemblies intended for field emission equipment rated for pulsed operation, the maximum rated peak tube potential and the maximum rated number of x-ray pulses in an hour for operation at the maximum rated peak tube potential.

(c) For all other diagnostic or therapeutic source assemblies, the maximum rated peak tube potential and the maximum rated continuous tube current for the maximum rated peak tube potential.

(43) “Light field” means that area of the intersection of the light beam from the beam-limiting device and one of the sets of planes parallel to and including the plane of the image receptor, whose perimeter is the locus of points at which the illumination is one-fourth of the maximum in the intersection.

(44) “Line-voltage regulation” means the difference between the no-load and the load line potentials expressed as a percent of the load line potential. It is calculated using the following equation: Percent line-voltage regulation = 100 (Vn - V1)/V1 where Vn = no-load line potential and V1 = load line potential.

(45) “mAs” means milliampere second.

(46) “Maximum line current” means the root-mean-square current in the supply line of an x-ray machine operating at its maximum rating.

(47) “Mobile x-ray equipment”. See “X-ray equipment”.

(48) “Peak tube potential” means the maximum value of the potential difference across the x-ray tube during an exposure.

(49) “Phantom” means a volume of material behaving in a manner similar to tissue with respect to the attenuation and scattering of radiation.

(50) “Photofluorographic” means an x-ray system designed to superimpose a patient’s anatomical x-ray image from a fluoroscopic input phosphor onto a film strip through a system of lenses.

(51) “Phototimer” means a device for controlling radiation exposures to image receptors by the amount of radiation which reaches a radiation monitoring device. The radiation monitoring device is part of an electronic circuit which controls the duration of time the tube is energized. See also “Automatic exposure control”.

(52) “Portable x-ray equipment”. See “X-ray equipment”.

(53) “Position indicating device (PID)” means a device on dental x-ray equipment used to indicate the beam position and to establish a definite source-skin distance. It may or may not incorporate or serve as a beam-limiting device.

(54) “Primary dose monitoring system” means a system which will monitor the useful beam during irradiation and which will terminate irradiation when a preselected number of dose monitor units have been delivered.

(55) “Primary protective barrier”. See “Protective barrier”.

(56) “Protective apron” means an apron made of radiation absorbing material used to reduce radiation exposure.

(57) “Protective barrier” means a barrier containing radiation absorbing material used to reduce radiation exposure. The types of protective barriers are as follows:

(a) “Primary protective barrier” means the material, excluding filters, placed in the useful beam to reduce the radiation exposure by a required degree, for protection purposes.

(b) “Secondary protective barrier” means a barrier sufficient to attenuate the stray radiation by a required degree, for protection purposes.

(58) “Protective glove” means a glove made of radiation absorbing material and used to reduce radiation exposure.

(59) “Qualified person” means an individual who has the knowledge and training to measure ionizing radiation, to evaluate safety techniques and to advise regarding radiation protection needs.

(60) “Radiation detector” means a device which, in the presence of radiation, provides a signal or other indication suitable for use in measuring single or multiple quantities of incident radiation.

(61) “Radiation therapy simulation system” means a radiographic or fluoroscopic x-ray system, intended for localizing the volume to be exposed during radiation therapy, and confirming the position and size of the therapeutic irradiation field.

(62) “Radiograph” means an image receptor on which the image is created directly or indirectly by an x-ray pattern and results in a permanent record.

(63) “Radiographic imaging system” means any system whereby a permanent or semi-permanent image is recorded on an image receptor by the action of ionizing radiation.

(64) “Radiological physicist” means an individual who meets one of the following criteria:

(a) Is certified by the American Board of Radiology in therapeutic radiological physics, radiological physics, or x and gamma ray physics; or

(b) Has a bachelor’s degree in a physical science or engineering and 3 years full-time experience working in therapeutic radiological physics under the direction of a physicist certified by the American Board of Radiology. The work duties must include duties involving the calibration and spot checks of a medical accelerator or a sealed source teletherapy unit; or

(c) Has a master’s or a doctor’s degree in physics, biophysics, radiological physics, health physics or engineering; has had 1 year of full-time training in therapeutic radiological physics; and has had 1 year of full-time work experience in a radiotherapy facility where the individual’s duties involved calibration and spot checks of a medical accelerator or a sealed source teletherapy unit; or

(d) Has performed radiation physics work for a period of at least 10 years full time, prior to the effective date of these rules, in the field of therapeutic radiological physics in radiotherapy facilities where the individual’s duties involved calibration and spot checks of a medical accelerator or a sealed source teletherapy unit.

(65) “Rating” means the operating limits of a component as specified by the component manufacturer.

(66) “Recording” means producing a permanent form of an image resulting from x-ray photons, such as film or video tape.

(67) “Registrant”, as used in this part and in Parts IV, VI and VIII, means any person who possesses and administratively controls an x-ray system or other radiation producing machine and is required by the provisions in Part I to register with this Department.

(68) “Response time” means the time required for an instrument system to reach 90 percent of its final reading when the radiation-sensitive volume of the instrument system is exposed to a step change in radiation flux from zero sufficient to provide a steady state midscale reading.

(69) “Scattered radiation” means radiation that, during passage through matter, has been deviated in direction. See “Direct scattered radiation”.

(70) “Secondary dose monitoring system” means a system which will terminate irradiation in the event of failure of the primary system.

(71) “Secondary protective barrier”. See “Protective barrier”.

(72) “Shutter” means a device attached to the tube housing assembly which can totally intercept the useful beam and which has a lead equivalency not less than that of the tube housing assembly.

(73) “Source” means the focal spot of the x-ray tube.

(74) “Source-image receptor distance (SID)” means the distance from the source to the center of the input surface of the image receptor.

(75) “Spot check” means a procedure which is performed to assure that a previous calibration continues to be valid.

(76) “Spot film” means a radiograph which is made during a fluoroscopic examination to permanently record conditions which exist during that fluoroscopic procedure.

(77) “Spot-film device” means a device intended to transport or position a radiographic image receptor between the x-ray source and fluoroscopic image receptor. It includes a device intended to hold a cassette over the input end of an image intensifier for the purpose of making a radiograph.

(78) “SSD” means the distance between the source and the skin of the patient.

(79) “Stationary x-ray equipment”. See “X-ray equipment”.

(80) “Stray radiation” means the sum of leakage and scattered radiation.

(81) “Technique factors” means the conditions of operation. They are specified as follows:

(a) For capacitor energy storage equipment, peak tube potential in kV and quantity of charge in mAs;

(b) For field emission equipment rated for pulsed operation, peak tube potential in kV and number of x-ray pulses; and

(c) For all other equipment, peak tube potential in kV and either tube current in mA and exposure time in seconds, or the product of tube current and exposure time in mAs.

(82) “Termination of irradiation” means the stopping of irradiation in a fashion which will not permit continuance of irradiation without the resetting of operating conditions at the control panel.

(83) “Traceable to a national standard” means that a quantity or a measurement has been compared to a national standard directly or indirectly through one or more intermediate steps and that all comparisons have been documented.

(84) “Tube” means an x-ray tube, unless otherwise specified. See “X-ray tube”.

(85) “Tube housing assembly” means the tube housing with tube installed. It includes high voltage or filament transformers and other appropriate elements when such are contained within the tube housing.

(86) “Tube rating chart” means the set of curves provided by the manufacturer which specify the rated limits of operation of the tube in terms of the technique factors.

(87) “Useful beam” means the radiation emanating from the tube housing port or the radiation head and passing through the aperture of the beam-limiting device when the exposure controls are in a mode to cause the system to produce radiation.

(88) “Variable-aperture beam-limiting device” means a beam-limiting device which has capacity for stepless adjustment of the x-ray field size at a given SID.

(89) “Visible area” means that portion of the input surface of the image receptor over which incident x-ray photons are producing a visible image.

(90) “Wedge filter” means an added filter effecting continuous progressive attenuation on all or part of the useful beam.

(91) “X-ray control” means a device which controls input power to the x-ray high-voltage generator or the x-ray tube. It includes equipment such as timers, phototimers, automatic brightness stabilizers and similar devices, which control the technique factors of an x-ray exposure.

(92) “X-ray equipment” means an x-ray system, subsystem or component thereof.

(a) “Mobile” means x-ray equipment mounted on a permanent base with wheels or casters for moving while completely assembled.

(b) “Portable” means x-ray equipment designed to be hand carried.

(c) “Stationary” means x-ray equipment which is installed in a fixed location.

(d) “Special purpose” means x-ray equipment or a system designed for radiographic examinations of a specific anatomical area of the human body utilizing image receptors of more than one size; for example, the head or the spinal column.

(93) “X-ray field” means that area of the intersection of the useful beam and any one of the set of planes parallel to and including the plane of the image receptor, whose perimeter is the locus of points at which the exposure rate is one-fourth of the maximum in the intersection.

(94) “X-ray high-voltage generator” means a device which transforms electrical energy from the potential supplied by the x-ray control to the tube operating potential. The device may also include means for transforming alternating current to direct current, filament transformers for the x-ray tube, high-voltage switches, electrical protective devices and other appropriate elements.

(95) “X-ray system” means an assemblage of components for the controlled production of x-rays. It minimally includes an x-ray high-voltage generator, an x-ray control, a tube housing assembly, a beam-limiting device and the necessary supporting structures. Additional components which function with the system are considered integral parts of the system.

(96) “X-ray subsystem” means any combination of two or more components of an x-ray system.

(97) “X-ray tube” means any electron tube which is designed for the conversion of electrical energy into x-ray energy.

(98) “Medical physicist” means a person who practices the branch of physics that is associated with the practice of medicine.

(99) “Clinical image” means a radiograph.

(100) “Pulsed mode” means operation of the x-ray system such that the x-ray tube current is pulsed by the x-ray control to produce one or more exposure intervals of duration less than one-half second.

Rulemaking Authority 404.051 FS. Law Implemented 404.022, 404.031, 404.051(1), (4) FS. History–New 7-17-85, Amended 4-4-89, 11-20-94, 1-5-95, 5-15-96, Formerly 10D-91.602.

- 64E-5.502 General Requirements.

(1) Administrative Controls.

(a) Registrant. The registrant shall be responsible for directing the operation of the x-ray systems which are subject to registration as described in Rule 64E-5.511, F.A.C. The registrant or the registrant’s agent shall assure that the following requirements are met in the operation of the x-ray system.

1. Any x-ray system which does not meet the provisions of these regulations shall not be operated for diagnostic or therapeutic purposes unless the Department determines that such operation will not endanger the public health, safety and welfare.

2. An x-ray system operator shall receive instruction on and be competent in the safe use of the x-ray system. A medical x-ray system operator shall also be certified or authorized in accordance with Chapter 468, Part IV, F.S.

3. A chart shall be provided in the vicinity of the diagnostic x-ray system’s control panel, which specifies techniques and procedures to be used for all examinations which are performed by that system.

4. Except for patients who cannot be moved out of the room, only the staff and ancillary personnel required for the medical procedure or training shall be in the room during the radiographic exposure. Other than the patient being examined:

a. All individuals shall be positioned such that no part of the body will be struck by the useful beam unless protected by 0.5 millimeter lead equivalent.

b. Staff and ancillary personnel shall be protected from the direct scatter radiation by protective aprons or whole body protective barriers of not less than 0.25 millimeter lead equivalent.

c. Other patients who cannot be removed from the room shall be protected from the direct scatter radiation by whole body protective barriers of 0.25 millimeter lead equivalent or shall be so positioned that the nearest portion of the body is at least 2 meters from both the tube head and the nearest edge of the useful beam.

d. When a portion of the body of any staff or ancillary personnel is potentially subjected to stray radiation which could result in that individual receiving one-fourth of the maximum permissible dose as defined in Part III, additional protective devices may be required by the Department.

5. Gonad shields of not less than 0.25 millimeter lead equivalent shall be used for patients who have not passed the reproductive age, during radiographic procedures in which the gonads are in the useful beam, except for cases in which this would interfere with the diagnostic procedure.

6. Individuals shall not be exposed to the useful beam except for healing arts purposes and unless such exposure has been authorized by a licensed practitioner of the healing arts. This provision specifically prohibits the following:

a. Exposure of an individual for training, demonstration or other purposes unless there are also healing arts requirements and a proper prescription has been provided.

b. Exposure of an individual for the purpose of a healing arts self-referral program except when authorized by subparagraph 64E-5.502(1)(a)1., F.A.C.

c. Advertisement of free x-ray examinations unless the advertisement states that a determination of need will be made prior to the x-ray examination.

7. An x-ray system is exempt from the requirements of subparagraph 64E-5.502(1)(a)6., F.A.C., only if:

a. The system is a stationary, non-mobile system installed and used only in a jail or penal institution;

b. The system is used only on legal detainees in a jail or penal institution, and never on detainees’ family members, children, institution employees, contractors, visitors, the public, or any other persons;

c. The system is manufactured, maintained, and operated solely for security screening purposes in strict compliance with, and fully according to, the most restrictive standards found in ANSI/HPS N43.17-2009, “Radiation Safety for Personnel Security Screening Systems Using X-Ray or Gamma Radiation,” which is herein incorporated by reference. This ANSI publication may be examined and inspected at the Florida Department of Health, Bureau of Radiation Control at Building 4042 Bald Cypress Way, Suite 210, Tallahassee, Florida 32399-1741, and at the Florida Department of State, Room 701, The Capitol, Tallahassee, Florida 32399-0250, and is available from the American National Standards Institute, Inc., at http://www.ansi.org. The agency has determined that posting the publication on the internet for purposes of public inspection and examination would constitute a violation of federal copyright law; and,

d. The institution tracks radiation dose for each detainee, as specified by ANSI/HPS N43.17-2009, “Radiation Safety for Personnel Security Screening Systems Using X-Ray or Gamma Radiation,” to ensure the detainee does not exceed the recommended dose limit.

8. When a patient or film must be provided with auxiliary support during a radiation exposure:

a. Mechanical holding devices shall be used when the technique permits;

b. Written safety procedures shall be available to indicate the requirements for selecting a holder, list the individual projections where holding devices cannot be used and describe the procedure the holder shall follow;

c. The human holder shall be protected as required by subparagraph (1)(a)4., above; and,

d. No individual shall be used routinely to hold film or patients.

9. Exposure Procedures Designed to Minimize Patient and Personal Exposure.

a. The speed of film or screen and film combinations shall be the fastest speed consistent with the diagnostic objectives of the examination.

b. The radiation exposure to the patient shall be the minimum exposure required to produce images of good diagnostic quality.

c. Portable or mobile equipment shall be used only for examinations where it is impractical to transfer the patient to a stationary radiographic installation.

d. X-ray systems subject to Rule 64E-5.505, F.A.C., shall not be utilized in procedures where the source to patient distance is less than 30 centimeters.

e. A person shall not perform fluoroscopic imaging or otherwise expose a human to x-rays from a fluoroscopic system unless the person is a:

(I) Licensed practitioner as that term is defined in Section 468.301, F.S.; or

(II) Certified radiologist assistant practicing in accordance with the requirements of Chapter 468, Part IV, F.S.; or

(III) Certified general radiographer practicing in accordance with the requirements of Chapter 468, Part IV, F.S.; and

(A) The general radiographer has been trained and authorized in writing by the licensed practitioner in charge to perform the specified imaging; and

(B) The specified imaging does not rely upon the general radiographer to provide any diagnostic interpretation, or to determine suspicious areas for additional imaging, or to otherwise modify the scope of authorization for the imaging; and

(C) The specified imaging is designed to prevent or reduce exposure to patients by facilitating proper location and positioning for the authorized radiographic imaging.

10. Personnel Monitoring. All individuals who are associated with the operation of an x-ray system are subject to the occupational exposure limits and the requirements for the determination of the doses stated in Rules 64E-5.304 and 64E-5.308, F.A.C. In addition, when protective clothing or devices are worn on portions of the body and a personnel monitoring device is required, at least one such device shall be utilized as follows:

a. When a protective apron is worn, the monitoring device shall be worn at the collar outside of the apron.

b. The dose to the whole body shall be recorded in the records required by Rule 64E-5.339, F.A.C. If more than one device is used and a record is made of the data, each dose shall be identified with the area where the device was worn on the body.

11. Healing arts self-referral. Only healing arts self-referral programs for mammography screening will be authorized by the Department.

(b) Information and Maintenance Records and Associated Information. The registrant shall maintain at least the following information for each x-ray system:

1. Tube rating charts and cooling curves.

2. Record of surveys, calibrations, maintenance, modifications from the original schematics and drawings performed on the x-ray machine along with the names of persons who performed the service.

3. A copy of all correspondence with the Department regarding each x-ray system.

4. An x-ray log containing the patient’s name, the type of examination and the dates the examinations were performed. When the patient or film must be provided with human auxiliary support, the name of the human holder shall be recorded.

(c) Maintenance of X-ray Equipment. X-ray systems and accessory devices shall be maintained in good working condition, both mechanically and electrically, so that the clinical objectives may be fulfilled without risk of unproductive exposure due to equipment failure or malfunction.

(2) Shielding.

(a) Each x-ray facility shall have primary and secondary protective barriers as needed to assure that an individual will not receive a radiation dose in excess of the limits specified in Part III of these regulations.

(b) Structural shielding in walls and other vertical barriers required for personnel protection shall extend without breach from the floor to a height of at least 7 feet (2.1 m).

(c) Doors, door frames, windows and window frames shall have the same lead equivalent shielding as that required in the wall or other barrier in which they are installed.

(d) In computation of protective barrier requirements, the maximum anticipated workload, use factors, occupancy factors and the potential for radiation exposure from other sources shall be taken into consideration.

(e) Prior to construction, the floor plans and equipment arrangement of all new installations, or modifications of existing installations, utilizing x-ray energies of 200 keV and above for diagnostic or therapeutic purposes shall be submitted to the Department for review and approval.

1. The plans shall show, as a minimum, the following:

a. The normal location of the x-ray system’s radiation port; the port’s travel and traverse limits; general direction of the useful beam; locations of any windows and doors; the location of the operator’s booth; and the location of the x-ray control panel.

b. The structural composition and thickness or lead equivalent of all walls, doors, partitions, floor and ceiling of the room concerned.

c. The dimensions of the room concerned.

d. The type of occupancy of all adjacent areas inclusive of space above and below the room concerned. If there is an exterior wall, the distance to the closest area where it is likely that individuals may be present.

e. The make and model of the x-ray equipment and the maximum technique factors.

f. The type of examinations or treatments which will be performed with the equipment.

2. Information on the anticipated maximum workload of the x-ray system.

3. If the services of a qualified person have been utilized to determine the shielding requirements, a copy of the report, including all basic assumptions used, shall be submitted with the plans.

(3) X-ray Film Processing Facilities and Practices.

(a) Processing Facilities. Each installation using a radiographic x-ray system shall provide suitable equipment for handling and processing radiographic film in accordance with the following provisions:

1. The area in which undeveloped films are handled for processing shall be devoid of light with the exception of light in the wave lengths having no significant effect on the radiographic film.

2. Film pass boxes, if provided, shall be so constructed as to exclude light when film is placed in or removed from the boxes, and shall incorporate adequate shielding to prevent exposure of undeveloped film to stray radiation.

3. Darkrooms used by more than one individual shall be provided a positive method to prevent accidental entry while undeveloped films are being handled or processed.

4. Where film is developed manually,

a. At least one tri-sectional tank made of mechanically rigid, corrosion resistant material shall be utilized; and

b. The temperature of each solution shall be maintained within the range of 60º F to 80º F (16º C to 27º C). Film shall be developed in accordance with the time-temperature relationships specified by the film manufacturer, or, in the absence of such recommendations by the film manufacturer, with the following time-temperature chart:

Time-Temperature Chart

Thermometer Reading								Minimum Developing

(Degrees)								    Time (Minutes)

C				F

26.7			80						2

26.1			79						2

25.6			78						2 1/2

25.0			77						2 1/2

24.4			76						3

23.9			75						3

23.3			74						3 1/2

22.8			73						3 1/2

22.2			72						4

21.7			71						4

21.1			70						4 1/2

20.6			69						4 1/2

20.0			68						5

19.4			67						5 1/2

18.9			66						5 1/2

18.3			65						6

17.8			64						6 1/2

17.2			63						7

16.7			62						8

16.1			61						8 1/2

15.6			60						9 1/2

c. Devices shall be utilized which will:

(I) Indicate the actual temperature of the developer; and

(II) Signal the passage of a preset time as short as 2 minutes.

(b) Precautionary Practices.

1. Film shall be stored in a cool, dry place and shall be protected from exposure to stray radiation. Film in open packages shall be stored in a light tight container.

2. Film cassettes and intensifying screens shall be inspected periodically and shall be cleaned and replaced as necessary to best assure radiographs of good diagnostic quality.

3. Outdated x-ray film shall not be used for human diagnostic radiographs, unless the film has been stored in accordance with the manufacturer’s recommendations and a sample of the film passes a sensitometric test for normal ranges of base fog and speed.

4. Film developing solutions shall be prepared in accordance with the directions given by the manufacturer, and shall be maintained in strength by replenishment or renewal so that full development is accomplished within the time specified by the manufacturer.

5. Safe light and darkroom fog shall be such that, when a radiographic film is exposed to radiation to achieve a density of 1.0 and is exposed for 1 minute on any darkroom working surface, the film shall not have a density change greater than 0.1.

(c) Automatic Processors and Other Closed Processing Systems. Preventive maintenance shall be performed on the unit, except for extended periods of nonuse, on a frequency basis which is not less than that schedule recommended by the manufacturer. In the event that no schedule is available from the manufacturer, a maintenance schedule shall be established which will preserve good diagnostic film quality.

(d) Radiographic Film Quality.

1. Developed radiographs of patients or phantoms shall have an optical density of 0.5 to 2.0 in the area of clinical interest to allow for diagnostic interpretation of the image, unless justified due to special circumstances. Radiographs which provide the necessary diagnostic information shall not be repeated for the sole purpose of meeting the stated density range.

2. Radiographic film used for diagnostic purposes shall be free from light fog and artifacts.

Rulemaking Authority 404.051, 404.081, 404.141, 404.22 FS. Law Implemented 404.051, 404.081, 404.141, 404.22, 468.302 FS. History–New 7-17-85, Amended 4-4-89, 1-1-94, 11-20-94, 1-5-95, Formerly 10D-91.603, Amended 5-18-98, 8-16-07, 3-21-16.

- 64E-5.503 General Requirements for All Diagnostic X-Ray Systems.

In addition to other requirements of this part, all diagnostic x-ray systems shall meet the following requirements:

(1) Warning Label. The main control panel and all auxiliary control panels of the x-ray system shall bear the equivalent warning statement, legible and accessible to view, “WARNING: This x-ray unit may be dangerous to patient and operator unless safe exposure factors and operating instructions are observed.”

(2) Battery Charge Indicator. Visual means shall be provided on the control panel of battery-powered x-ray generators to indicate whether the battery is in a state of charge adequate for proper operation.

(3) Leakage Radiation from the Diagnostic Source Assembly. The leakage radiation from the diagnostic source assembly measured at a distance of 1 meter in any direction from the source shall not exceed 100 milliroentgens (25.8 uC per kg) in 1 hour when the x-ray tube is operated at its leakage technique factors. Compliance shall be determined by measurements averaged over an area of 100 square centimeters with no linear dimension greater than 20 centimeters.

(4) Radiation from Components Other than the Diagnostic Source Assembly. The radiation emitted by a component other than the diagnostic source assembly shall not exceed 2 milliroentgens (0.516 uC per kg) in 1 hour at 5 centimeters from any accessible surface of the component when it is operated in an assembled x-ray system under any conditions for which it was designed. Compliance shall be determined by measurements averaged over an area of 100 square centimeters with no linear dimension greater than 20 centimeters.

(5) Beam Quality.

(a) Half-value Layer. The half-value layer (HVL) of the useful beam for a given x-ray tube potential shall not be less than the values shown below. If it is necessary to determine such half-value layer at an x-ray tube potential which is not listed below, linear interpolation or extrapolation may be made.

Half-value

Design		Measured	layer

operating range	potential	(Millimeters of

(Kilovolts peak)	(Kilovolts peak)	aluminum)

Below 50	30	0.3

40	0.4

49	0.5

50 to 70	50	1.2

60	1.3

70	1.5

Above 70	71	2.1

80	2.3

90	2.5

100	2.7

110	3.0

120	3.2

130	3.5

140	3.8

150	4.1

1. The above HVL criteria will be considered to have been met if it can be demonstrated that the aluminum equivalent of the total filtration in the primary beam is not less than that shown below:

Filtration Required vs. Operating Voltage

Total Filtration

(inherent plus added)

Operating	(millimeters

Voltage	aluminum

(kVp)	equivalent)

Below 50	0.5 millimeters

50-70	1.5 millimeters

Above 70	2.5 millimeters

2. Beryllium window tubes shall have a minimum of 0.5 millimeters aluminum equivalent filtration permanently mounted in the useful beam.

3. For capacitor energy storage equipment, compliance shall be determined with the maximum quantity of charge per exposure.

4. The required minimum aluminum equivalent filtration shall include the filtration contributed by all materials which are always present between the focal spot of the tube and the patient.

5. In addition to the requirements of paragraph (5)(a)1., above, all intraoral dental radiographic systems manufactured on and after December 1, 1980, shall have a minimum half-value layer not less than 1.5 millimeters aluminum equivalent filtration permanently installed in the useful beam.

(b) Filtration Controls. For x-ray systems which have variable kVp and changeable filters, and which are used for low filtration techniques, a positive means shall be provided that will prevent an exposure unless the minimum required amount of filtration is in the useful beam for the selected kVp.

(6) Aluminum Equivalent of Material between Patient and Image Receptor. The aluminum equivalent of each of the items listed below, which are used between the patient and image receptor, shall not exceed the indicated limits. This requirement is applicable to the front panel of cassette holders and film changers provided by the manufacturer for purposes of patient support or to prevent foreign object intrusions. It does not apply to such items as a screen and its associated mechanical support panel or grids.

Maximum

Aluminum

Equivalent

Item	(millimeters)

Front panel of cassette holder (total of all)	1.0

Front panel of film changer (total of all)	1.0

Stationary tabletop	1.0

Moveable tabletop (including stationary subtop)	1.5

Cradle	2.0

(7) Multiple Tube Heads. Where two or more radiographic tubes are controlled by one exposure switch, the tube or tubes which have been selected shall be clearly indicated prior to initiation of the exposure. This indication shall be both on the x-ray control panel and at or near the selected tube housing assembly.

(8) Mechanical Support of Tube Head. The tube housing assembly supports shall be adjusted such that the tube housing assembly will remain stable during an exposure unless the tube housing movement is a designed function of the x-ray system.

(9) Technique Indicators.

(a) Each x-ray system shall be equipped with devices, such as labeled control settings or meters, correctly indicating the physical factors and modes of operation used for exposures. X-ray systems utilizing arbitrary number or letter designators for kVp, time and milliamperage shall be accompanied by a chart giving the value of physical factors for each arbitrary designator.

(b) The technique factors to be used during an exposure shall be indicated before the exposure begins, except when automatic exposure controls are used, in which case the technique factors which are set prior to the exposure shall be indicated.

(c) On equipment having fixed technique factors, the requirement in paragraph (9)(a), above, may be met by permanent markings. Indication of technique factors shall be visible from the operator’s position except in the case of fluoroscopy or spot films made by the fluoroscopist.

(d) Reproducible technique factor indicators may be relabeled to meet the accuracy requirements of this part. Such relabeling shall be conspicuous and clearly legible and shall be utilized by the registrant in setting technique factors.

(10) Accuracy of Technique Factors. Meters, labeled control settings, exposure time selectors and other physical factor indicators shall be accurate within the following tolerances:

(a) Milliamperage (mA) – Radiographic ± 10%

(b) Milliamperage (mA) – Fluoroscopic ± 0.2 mA

(c) Kilovolt peak (kVp) ± 5 percent

(d) Timer at settings:

1. Greater than 10 seconds ± 1 second

2. Ten seconds or less ± ten percent

(11) Timer Reproducibility. When four timer tests are performed at the same timer settings, the average time period (T) shall be greater than or equal to 12 times the maximum time period (Tmax) less the minimum time period (Tmin). T shall be less than or equal to 0.5 seconds. Expressed mathematically, T >= 12 (Tmax - Tmin).

(12) Exposure Reproducibility. The x-ray exposure produced by radiographic systems shall be reproducible to within the following criteria: When all technique factors are held constant and four or more exposures at the same technique factors are made, the value of the average exposure (E) shall be greater than or equal to 12 times the quantity of maximum exposure (Emax) minus the minimum exposure (Emin). Expressed mathematically, E >= 12 (Emax - Emin).

(13) Exposure Linearity. The x-ray output produced by radiographic systems utilizing means other than automatic exposure controls shall be linear to within the following criteria:

(a) When a choice of two or more current settings (mA) or current-time product settings (mAs) may be selected where X1 and X2 are the average mR per mAs values obtained from four exposures on each of two mA or mAs settings at a fixed tube potential (kVp) setting, within the range of 40 to 100 percent of the maximum tube rating, the average of four exposures (mR) for a given milliampere-second (mAs) product in mR per mAs shall not differ by more than:

1. Five hundredths times the sum of any 2 consecutive mA or mAs averaged settings; expressed mathematically,

|X1 - X2 |<= 0.05 (X1 + X2); and

2. One tenth times the sum of any other 2 mA or mAs averaged settings; expressed mathematically,

|X1 - X2 |<= 0.10 (X1 - X2)

(b) Equipment, which after calibration cannot be made to meet the requirements of paragraph (13)(a), above, may be relabeled to indicate the effective mA or mAs, 8providing that use of such relabeled stations will meet the requirements of paragraph (13)(a), above.

(c) Equipment, which after calibration cannot be made to meet the requirements of paragraph (13)(a) or (b), above, shall not be used unless written approval is obtained from the Department. Approval shall not be granted when the linearity determination exceeds the federal standard for certified systems.

(14) Automatic Exposure Controls. When automatic exposure control is provided:

(a) Indication shall be made on the control panel when this mode of operation is selected.

(b) When the x-ray tube potential is greater than or equal to 50 kVp, the minimum exposure time for field emission equipment rated for pulsed operation shall be less than or equal to a time interval equivalent to 2 pulses.

(c) The minimum exposure time for all equipment other than that specified in paragraph (14)(b), above, shall be less than or equal to 1/60 second or a time interval required to deliver 5 mAs, whichever is greater.

(d) Either the product of peak x-ray tube potential, current, and exposure time shall be limited to not more than 60 kWs per exposure, or the product of x-ray tube current and exposure time shall be limited to not more than 600 mAs per exposure except when the x-ray tube potential is less than 50 kVp, in which case the product of x-ray tube current and exposure time shall be limited to not more than 2,000 mAs per exposure.

(e) A visible signal shall indicate when an exposure has been terminated at the limits described in paragraph (14)(d), above, and manual resetting shall be required before further automatically timed exposures can be made.

(f) Exposure Linearity. When a choice of two or more tube current settings (mA) may be selected, the average of four consecutive exposures (E1) made at any one tube current setting minus the average of four consecutive exposures (En) made at any other tube current setting shall be less than or equal to 0.05 times the sum of the two averages. Expressed mathematically, |E1 - En |<= 0.05 (E1 + En). Measuring compliance for the above shall be based on the following criteria:

1. An attenuation block as described in subsection 64E-5.501(5), F.A.C., shall be in the useful beam.

2. Exposure (mR) shall be measured on the exit side of the attenuation block.

3. The tube potential (kVp) shall be maintained at a fixed setting within the range of 40 percent to 100 percent of the maximum tube rating.

(15) Beam Limiting Devices.

(a) Beam limiting devices capable of restricting the useful beam to the area of clinical interest shall be used during exposures.

(b) Beam limiting devices shall provide a degree of attenuation not less than that required for the tube housing.

(16) Remote Exposure Switches. Where an x-ray control is equipped with two or more remote exposure switches, each remote switch shall serve a single x-ray tube, and exposures with any tube shall be possible only by the remote switch with which that particular tube is associated.

(17) Electrical Power Supply. The electrical power supply and service lines to x-ray systems shall be of sufficient capacity to permit operation without significant variation in voltage or machine output.

Rulemaking Authority 404.051, 404.151, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.141, 404.22(1), (3) FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.604.

- 64E-5.504 Fluoroscopic X-Ray Systems.

All fluoroscopic x-ray systems shall meet the following requirements:

(1) Limitation of the Useful Beam.

(a) The fluoroscopic tube shall not produce x-rays unless the primary protective barrier is in position to intercept the entire cross section of the useful beam.

(b) A means shall be provided between the x-ray source and the patient for stepless adjustment of the size of the x-ray field.

(c) With the collimating shutters adjusted to the closed position, the minimum field size at the maximum SID shall not be greater than 5 × 5 centimeters when measured at the point where the beam enters the patient.

(d) Limitation to the Imaging Surface.

1. The x-ray field produced by nonimage-intensified fluoroscopic equipment shall not extend beyond the useable area of the largest image receptor at any SID.

2. The longitudinal and transverse dimensions of the x-ray field produced by image-intensified fluoroscopic equipment shall not extend beyond the corresponding dimensions of the visible area of the image receptor by more than 3 percent of the SID in either dimension in the plane of the image receptor and the sum of the excess shall be no greater than 4 percent of the SID. If the collimation is automatically accomplished, the x-ray field dimension criteria above shall apply to all film sizes and portions thereof that the spot film device accommodates and to the dimensions of the input phosphor, as appropriate. If collimation is not automatic, the x-ray field dimension criteria shall apply to the useful area of the input phosphor.

3. Compliance shall be determined with the beam axis perpendicular to the plane of the image receptor. For rectangular x-ray fields used with circular image receptors, the error in alignment shall be determined along the length and width dimensions of the x-ray field which passes through the center of the visible area of the image receptor.

4. The center of the x-ray field in the plane of the image receptor shall be aligned with the center of the selected portion of the image receptor to within 2 percent of the SID.

5. Adjustable automatic and manual collimators shall operate smoothly throughout the entire range of use.

6. For fluoroscopic systems with spot film capability, means shall be provided for adjustment of the x-ray field size in the plane of the film to a size smaller than the selected portion of the film.

(e) The requirements of paragraph (1)(b) and (c), above, are not applicable to mobile fluoroscopic systems.

(2) Activation of the Fluoroscopic Tube. A control of the dead-man type shall be incorporated into each fluoroscopic system such that x-ray production will be terminated at any time pressure is released from the switch except during the recording of serial fluoroscopic images with equipment in which means have been provided to permit completion of any single exposure of the series in progress.

(3) Allowable Entrance Exposure Rate Limits for Fluoroscopic Equipment.

(a) Fluoroscopic equipment manufactured after June, 1995, operable at any combination of tube potential and current that results in an exposure rate greater than 5 roentgens (1.29 × 10-3 C per kg) per minute at the point where the center of the useful beam enters the patient shall be equipped with automatic exposure control. Provision for manual selection of technique factors can be provided.

(b) Fluoroscopic equipment shall not be operable at any combination of tube potential and current that will result in an exposure rate in excess of 10 roentgens (2.58 × 10-3 C per kg) per minute at the point where the center of the useful beam enters the patient except:

1. During the recording of images from an x-ray image-intensifier tube using photographic film or a video camera when the x-ray source is operated in a pulsed mode.

2. When an optional high-level control is activated. When the high-level control is activated, the equipment shall not be operable at any combination of tube potential and current that will result in an exposure rate in excess of 20 roentgens (5.16 × 10-3 C per kg) per minute at the point where the center of the useful beam enters the patient. Special means to activate high-level controls shall be required. The high-level control shall only be operable when continuous manual activation is provided by the operator.

(c) Special means to activate high level controls such as additional pressure applied continuously by the operator shall be required to avoid accidental use.

(d) A continuous signal audible to the fluoroscopist shall indicate when the high level control is being employed.

(e) Compliance with the dose limits will be determined as follows:

1. Movable grids and compression devices will be removed from the useful beam during the measurement.

2. The fluoroscope’s radiation output will be maximized.

a. Systems with automatic exposure controls such as automatic brightness control will have sufficient lead or lead equivalent materials placed in the useful beam to produce the maximum output.

b. Systems without automatic exposure controls or systems with a manual mode in addition to automatic exposure control modes will have the current and potential set to produce the maximum output. Attenuating material will be placed in the useful beam to protect the imaging system. If the registrant has a written radiation protection program restricting the range of current and potential the tests will be performed within the range of allowed values.

c. Patient support device height and SID, where adjustable, will be varied to produce the maximum output. If the registrant has a written radiation protection program restricting the range of patient support device heights or SIDs the tests will be performed within the range of allowed values.

3. The exposure rate will be measured at the following points on the centerline of the beam unless the specified geometry is prohibited by a written radiation protection program.

a. At least one centimeter above the patient support device and corrected for distance to show the actual entrance exposure rate at the top surface of the patient support device for:

(I) Fluoroscopes where the x-ray tube is fixed under the patient support device.

(II) C-arm systems or stationary c-arm fluoroscopes where the x-ray tube can be rotated under the patient support device. The x-ray tube will be positioned as close to the patient support device as possible.

b. At 30 centimeters above the patient support device with the end of the beam-limiting device or spacer assembly positioned as close as possible to the point of measurement for:

(I) Fluoroscopes where the x-ray tube is fixed above the patient support device.

(II) C-arm systems or stationary c-arm fluoroscopes where the x-ray tube can be rotated above the patient support device.

c. At a point 15 centimeters laterally from the centerline of the patient support device or from the centerline of the patient if the registrant has a written radiation protection program specifying placement of the patient not on the centerline of the patient support device in the direction of the x-ray tube with the input surface of the fluoroscopic imaging assembly positioned as close to the edge of the patient support device as possible but no closer than 15 cm for:

(I) Fluoroscopes where the x-ray tube is fixed laterally to the patient support device.

(II) C-arm systems or stationary c-arm fluoroscopes where the x-ray tube can be rotated lateral to the patient support device.

d. At 30 centimeters from the input surface of the fluoroscopic imaging assembly, provided that the end of the beam-limiting device or spacer is no closer than 30 centimeters from the input surface of the fluoroscopic imaging assembly, for mobile c-arm fluoroscopes. Spacers or other attachments normally used can not be removed to allow measuring from a point closer to the actual input surface.

(f) Periodic Measurement of Entrance Exposure Rates. The entrance exposure rate shall be measured before use on humans after the completion of any initial or subsequent installation and after any maintenance of the system that might affect the exposure rate.

(g) For cinefluoroscopy, the maximum exposure at the face of the input phosphor with the grid removed and with an attenuation block in the beam shall not exceed 40 microroentgens (0.010 µC per kg) per frame. The maximum exposure shall be measured before use on humans after the completion of any initial or subsequent installation and after any maintenance of the system which might affect the maximum exposure.

(4) Barrier Transmitted Radiation Rate Limits.

(a) The exposure rate due to transmission through the primary protective barrier and frame assembly with the attenuation block in the useful beam combined with radiation from the image intensifier if provided shall not exceed 2 milliroentgens (0.516 µC per kg) per hour at 10 centimeters from any accessible surface of the fluoroscopic imaging assembly beyond the plane of the image receptor for each roentgen per minute of entrance exposure rate.

(b) Measuring Compliance with Barrier Transmission Limits.

1. The exposure rate due to transmission through the primary protective barrier combined with radiation from the image intensifier shall be determined by measurements averaged over an area no greater than 100 square centimeters with no linear dimension greater than 20 centimeters.

2. If the source is below the tabletop, the measurement shall be made with the input surface of the fluoroscopic imaging assembly positioned 30 centimeters above the tabletop.

3. If the source is above the tabletop and the SID is variable, the measurement shall be made with the end of the beam limiting device or spacer assembly as close to the tabletop as it can be placed but not closer than 30 centimeters.

4. Movable grids and compression devices shall be removed from the useful beam during the measurement.

5. The attenuation block shall be positioned in the useful beam 10 centimeters toward the input surface of the imaging assembly from the point at which the entrance exposure rate was measured.

6. The maximum beam size shall be used during measurements.

(5) Indication of Potential and Current. During fluoroscopy and cinefluorography, x-ray tube potential and current shall be continuously indicated.

(6) Source-to-Skin Distance. Positive means shall be provided to assure the source-to-skin distance shall not be less than:

(a) Thirty-eight centimeters on stationary fluoroscopes installed after January 1, 1977,

(b) Thirty-five and one-half centimeters on stationary fluoroscopes installed prior to January 1, 1977,

(c) Thirty centimeters on all mobile fluoroscopes,

(d) Twenty centimeters for image intensified fluoroscopes used for specific surgical applications. Written safety procedures must be provided to the operator of the fluoroscope and precautionary measures followed during the use of this device.

(e) Nineteen centimeters for extremity-use-only fluoroscopes.

(f) Ten centimeters for extremity-use-only fluoroscopes used for specific surgical applications. Written safety procedures must be provided to the operator of the fluoroscope and precautionary measures followed during the use of this device.

(7) Fluoroscopic Timer. A cumulative timing device activated by the fluoroscopic exposure switch shall be provided, the maximum cumulative time of which shall not exceed 5 minutes without resetting. The timer shall indicate the passage of the predetermined period of exposure by an audible signal or termination of the exposure. If such a signal is utilized, it shall continue while x-rays are produced until the timing device is reset.

(8) Mobile Fluoroscopes. In addition to the other requirements of this section, mobile fluoroscopes shall provide intensified imaging.

(9) Control of Scatter Radiation.

(a) Fluoroscopic table designs shall be such that scattered radiation which originates beneath the tabletop is attenuated by not less than 0.25 millimeters lead equivalent, and that no unprotected part of any staff or ancillary person’s body shall be exposed to unattenuated scattered radiation.

(b) Fluoroscopic equipment configuration shall be such that no portion of any staff or ancillary person’s body, except the extremities, shall be exposed to the unattenuated scattered radiation emanating from above the tabletop unless:

1. Such person is at least 120 centimeters from the center of the useful beam, or

2. The radiation has passed through not less than 0.25 millimeter lead equivalent material.

(c) Exceptions to paragraph (10)(b), above, may be made in some special procedures where a sterile field will not permit the use of the normal protective barriers. Where the use of prefitted sterilized covers for the barriers is practical, the Department shall not permit such exception.

(10) Photofluorographic Medical X-Ray Systems.

(a) In addition to other applicable sections of these regulations, photofluorographic x-ray systems shall conform with the following requirements:

1. Usage shall be limited to diagnostic radiography of the lungs and other soft tissues of the thoracic region.

2. Personnel monitoring shall be provided for all individuals who operate photofluorographic apparatus.

3. The average exposure, including backscatter, for chests measuring 25 centimeters in thickness shall not exceed 100 millirems (1.0 mSv) at the point where the x-ray beam enters the patient.

(b) Photofluorographic x-ray systems shall not be installed unless specifically approved by the Department.

(11) Radiation Therapy Simulation Systems. Radiation therapy simulation systems shall be exempt from all the requirements of subsections (1), (3), (4), (5) and (8), above, provided that:

(a) Such systems are designed and used in such a manner that no person other than the patient is in an unprotected area during periods of time when the system is producing x-rays; and

(b) Systems that do not meet the requirements of subsection (8), above, are provided with a means of indicating the cumulative time that an individual patient has been exposed to x-rays. In such cases, the timer shall be reset between examinations.

(c) The exposure rate measured at the point where the center of the useful beam enters the patient shall not exceed 20 roentgens (5.16 mC per kg) per minute, except during the recording of fluoroscopic images.

(12) For remotely operated fluoroscopic systems:

(a) The remote control panel shall be installed so as to require the operator to stand behind a permanent protective barrier meeting the requirements of paragraph 64E-5.502(2)(a)-(c), F.A.C. The barrier must be wide enough to prevent the secondary scatter radiation from striking the operator directly when the machine is operated from the remote control panel.

(b) The operator must be able to see and hear the patient when behind the barrier.

(c) The barrier shall be constructed of material of sufficient density to meet or exceed the barrier requirements of sub-subparagraph 64E-5.502(1)(a)4.b., F.A.C.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.051, 404.22 FS. History–New 7-17-85, Amended 4-4-89, 3-17-92, 1-5-95, Formerly 10D-91.605, Amended 5-18-98, 8-16-07, 5-8-13.

- 64E-5.505 Diagnostic Radiography Systems, Other than Fluoroscopic, Mammographic, Dental Intraoral or Veterinary Systems.

(1) Beam Limitation. The useful beam shall be limited to the area of clinical interest.

(a) General Purpose Stationary and Mobile X-Ray Systems.

1. A means for stepless adjustment of the size of the x-ray field shall be provided.

2. Means shall be provided for visually defining the perimeter of the x-ray field. The total misalignment of the edges of the visually defined field with the respective edges of the x-ray field along either the length or width of the visually defined field shall not exceed 2 percent of the distance from the source to the center of the visually defined field when the surface upon which it appears is perpendicular to the axis of the x-ray beam.

3. Mobile x-ray systems shall be equipped with an attached rule to accurately measure the SID at any distance up to 72 inches (183 cm).

(b) Stationary general purpose diagnostic x-ray systems shall be equipped with the following additional features:

1. Positive means shall be provided to indicate when the axis of the x-ray beam is perpendicular to the plane of the image receptor; to align the center of the x-ray field with the center of the image receptor to within 2 percent of the SID; and to indicate the SID to within 2 percent.

2. The beam-limiting device shall numerically indicate the field size in the plane of the image receptor to which it is adjusted.

3. Indication of field size dimensions and SIDs shall be specified in inches or centimeters, and shall be such that aperture adjustments result in x-ray field dimensions in the plane of the image receptor which correspond to those indicated by the beam limiting device to within 2 percent of the SID when the beam axis is perpendicular to the plane of the image receptor.

(c) X-Ray Systems Used for One Image Receptor Size. Radiographic equipment used for only one image receptor size shall have a fixed SID and shall be provided with positive means to limit the x-ray field at the plane of the image receptor to the area of the image receptor, and to align the center of the x-ray field with the center of the image receptor to within 2 percent of the SID.

(d) Special Purpose X-Ray Systems.

1. For x-ray systems with more than one image receptor size,

a. Means shall be provided to limit the x-ray field in the plane of the image receptor so that such field does not exceed each dimension of the image receptor by more than 2 percent of the SID when the axis of the x-ray beam is perpendicular to the plane of the image receptor.

b. Means shall be provided to align the center of the x-ray field with the center of the image receptor to within 2 percent of the SID, or means shall be provided to both size and align the x-ray field such that the x-ray field at the plane of the image receptor does not extend beyond any edge of the image receptor.

2. The requirements in this paragraph are met by a system that meets the requirements for a general purpose x-ray system as specified in (1)(a), above, or, when positive alignment means are also provided, may be met with either:

a. An assortment of removable, fixed-aperture, beam-limiting devices sufficient to meet the requirement for each combination of image receptor size and SID for which the unit is designed; each such device shall have clear markings to indicate the image receptor size and SID for which it is designed; or

b. A beam-limiting device having multiple fixed apertures sufficient to meet the requirements for each combination of image receptor size and SID for which the unit is designed. Permanent, clearly legible markings shall indicate the image receptor size and SID for which each aperture is designed and shall indicate which aperture is in position for use.

(2) Radiation Exposure Control Devices.

(a) Timers. Means shall be provided to terminate the exposure at a preset time interval, preset product of current and time (mAs), a preset number of pulses, or a preset radiation exposure to the image receptor.

1. Except for dental panoramic systems, termination of exposure shall cause automatic resetting of the timer to its initial setting or to 0.

2. It shall not be possible to make an exposure when the timer is set to a zero or off position, if either position is provided.

(b) X-Ray Exposure Control Switch Type and Location.

1. A control of the dead-man type shall be incorporated into each x-ray system such that an exposure will be terminated at any time pressure is released from the switch, except during serial radiography, when means have been provided to permit completion of any single exposure of the series in progress.

2. Each x-ray control shall be located in such a way as to meet the following requirements:

a. The operator’s station at the control panel shall be behind a protective barrier so positioned that leakage radiation and once scattered radiation will be intercepted.

b. For panoramic dental units with intensifying screens and a beam stop, the operator shall stand at least 4 feet (1.25 m) from the patient and the tube head or behind a protective barrier during exposures.

c. The operator’s protective barrier shall be equipped with a window or mirror system so arranged that the operator may keep the patient under constant visual surveillance during exposures. The window shall have lead equivalent shielding equal to that required in the operator’s protective barrier.

d. Each exposure switch, except those used in conjunction with fluoroscopic spot film devices and movable protective barriers, shall be securely fixed so that the operator cannot conveniently make exposures from an unshielded position.

e. Provision shall be made for aural communication with the patient from the control panel.

f. Mobile and portable x-ray systems which are:

(I) Used continually in a single location for a period greater than 1 week shall be considered a stationary radiographic system and shall meet the requirements for such an installation.

(II) Used at multiple locations shall be provided either with an adequate protective barrier or protective apron for the operator and with a method of control which will permit the operator to be at least 12 feet (3.75 m) from the tube head and the nearest edge of the useful beam during exposures.

3. The x-ray control shall provide a visual indication observable from the operator’s protected position whenever x-rays are produced.

4. A sound audible to the operator shall indicate that the exposure has terminated or is in progress.

(3) Source-to-Skin Distance Limitations. All mobile or portable radiographic systems shall be provided with a positive means to limit the source-to-skin distance to not less than 30 centimeters.

(4) Standby Radiation from Capacitor Energy Storage Equipment. Radiation emitted from the x-ray tube when the exposure switch or timer is not activated shall not exceed a rate of 2 milliroentgens (0.516 mC per kg) per hour at 5 centimeters from any accessible surface of the diagnostic source assembly, with the beam-limiting device fully open.

(5) Intracavitary X-Ray Systems. Use of intracavitary x-ray systems on humans is prohibited unless specific approval has been granted by the Department. Application for such use must include evidence attesting to the exclusive advantages to be gained in the use of intracavitary radiographic techniques as opposed to conventional radiographic procedures.

Rulemaking Authority 404.051, 404.141, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.141, 404.22(1), (3) FS. History–New 7-17-85, Amended 3-17-92, 1-5-95, Formerly 10D-91.606.

- 64E-5.506 Intraoral Dental Radiographic Systems.

(1) Source-to-Skin Distance. X-ray systems designed for use with an intraoral image receptor shall be provided with means to limit source-to-skin distance to not less than:

(a) Eighteen centimeters if operable at or above 50 kVp, or

(b) Ten centimeters if not operable above 50 kVp.

(2) Field Limitation. Radiographic systems designed for use with an intraoral image receptor shall be provided with means to limit the x-ray beam such that:

(a) If the minimum source-to-skin distance (SSD) is 18 centimeters or more, the x-ray field at the minimum SSD shall be containable in a circle having a diameter of no more than 7 centimeters; and

(b) If the minimum SSD is less than 18 centimeters, the x-ray field at the minimum SSD shall be containable in a circle having a diameter of no more than 6 centimeters.

(c) An open-ended position indicating device shall be used on machines procured after September 19, 1972. The attenuation shall be equivalent to that required for the diagnostic source assembly as described in subsection 64E-5.503(3), F.A.C.

(3) Timers. Means shall be provided to terminate the exposure at a preset time interval, preset product of current and time (mAs), a preset number of pulses, or a preset radiation exposure to the image receptor. In addition:

(a) Termination of exposure shall cause automatic resetting of the timer to its initial setting or to 0; and

(b) It shall not be possible to make an exposure when the time is set to a 0 or off position, if either position is provided.

(4) X-Ray Exposure Control Switch.

(a) A control shall be incorporated into each x-ray system such that an exposure can be terminated at any time. This switch shall be of the dead-man type.

(b) Each x-ray control shall be located in such a way as to meet the following criteria:

1. The operator shall observe the patient during an exposure.

2. The operator shall stand as far as practicable and at least 6 feet (1.8 m) from the patient and tube head and outside the useful beam or behind a protective barrier during exposures.

(c) The x-ray control shall provide visual indication observable at or from the operator’s position whenever x-rays are produced.

(d) A sound audible to the operator shall indicate that the exposure has terminated or is in progress.

(5) Operating Controls.

(a) The dentist, operator or assistant shall not hold the film in place for the patient during the exposure. Patient and film holding devices shall be used when the techniques permit.

(b) No person other than the patient shall be exposed to the useful beam.

(c) Neither the tube housing nor the position indicating device shall be held during an exposure.

(d) The x-ray system shall be arranged and operated in such a manner that the useful beam at the patient’s skin does not exceed the dimensions specified in (2)(a), above.

(e) Dental fluoroscopy without image intensification is prohibited.

(f) Each user of intraoral units that are specifically designed to be handheld shall:

1. Have and use individual monitoring devices to document safe use practices; and

2. Successfully complete training provided by the manufacturer using electronic media such as CD/DVD or a website.  Training on the safe use of the unit shall be documented and include at a minimum:

a. Proper positioning of the unit to ensure an adequate protected position;

b. Limitations on the use of position indicating devices that require longer distances to the patients face;

c. Diagrams (i.e., drawings, illustrations, schematics, etc.) of protected position and location in relationship to the unit;

d. Diagrams (i.e., drawings, illustrations, schematics, etc.) of the effect of improper distance or removal of shielding device; and

e. Diagrams (i.e., drawings, illustrations, schematics, etc.) of common examples of improper positioning of the unit and or location of the operator.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.022, 404.051, 404.22 FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.607, 8-16-07.

- 64E-5.507 Therapeutic X-Ray Systems of Less Than 1 MeV.

(1) Equipment Requirements.

(a) Leakage Radiation. When the tube is operated at its leakage technique factors, the leakage radiation shall not exceed the value specified at the distance specified for the following classification of that x-ray system:

1. Contact Therapy Systems. Leakage radiation shall not exceed 100 milliroentgens (25.8 uC per kg) per hour at 5 centimeters from the surface of the tube housing assembly.

2. Zero to One Hundred Fifty kVp Systems. Leakage radiation shall not exceed 100 milliroentgens (25.8 uC per kg) in 1 hour at 1 meter from the source.

3. One Hundred Fifty-one to Nine Hundred Ninety-nine kVp Systems. The leakage radiation shall not exceed 0.1 percent of the useful beam 1 meter from the source, for any of its operating conditions.

(b) Permanent Beam Limiting Devices. Permanent fixed diaphragms or cones used for limiting the useful beam shall provide the same or a higher degree of protection as required for the tube housing assembly.

(c) Removable and Adjustable Beam Limiting Devices.

1. Removable beam limiting devices shall, for the portion of the useful beam to be blocked by these devices, transmit not more than 1 percent of the useful beam at the maximum kilovoltage and maximum treatment filter. This requirement does not apply to auxiliary blocks or materials placed in the useful beam to shape the useful beam to the individual patient.

2. Adjustable beam limiting devices shall transmit not more than 5 percent of the useful beam at the maximum kilovoltage and with the maximum treatment filter in the useful beam.

(d) Filter System. The filter system shall be so designed that:

1. The filters cannot be accidentally displaced at any possible tube orientation;

2. The radiation at 5 centimeters from the filter insertion slot opening does not exceed 30 roentgens (7.7 mC per kg) per hour under any operating conditions;

3. Each filter is marked as to its material of construction and its thickness. For wedge filters, the wedge angle shall appear on the wedge or wedge tray; and

4. A filter indication system shall be used on all therapy machines using changeable filters. It shall be designed to permit easy recognition of any added filter in place. The presence or absence of any filter shall be discernible at the control panel.

(e) Tube Immobilization. The tube housing assembly shall be capable of being immobilized for stationary treatments.

(f) Focal Spot Marking. The tube housing assembly shall be so marked that it is possible to determine the location of the focal spot to within 5 millimeters, and such marking shall be readily accessible for use during calibration procedures.

(g) Beam Block. Contact therapy tube housing assemblies shall have a removable shield of at least 0.5 millimeter lead equivalency at 100 kVp that can be positioned over the entire useful beam exit port during periods when the beam is not in use.

(h) Beam Monitor System. Systems of greater than 150 kVp manufactured after January 1, 1985, shall be provided with a beam monitor system which:

1. Shall have the detector of the monitor system interlocked to prevent incorrect positioning;

2. Shall not allow irradiation until a pre-selected value of exposure has been made at the treatment control panel;

3. Shall independently terminate irradiation when the preselected exposure has been reached;

4. Shall be so designed that, in the event of a system malfunction or electrical power failure, the dose administered to a patient prior to the system malfunction or power failure can be accurately determined;

5. Shall have a display at the control panel from which the dose at a reference point in soft tissue can be calculated;

6. Shall have a control panel display which maintains the administered dose reading until intentionally reset to 0; and

7. Shall have a control panel display which does not have scale multiplying factors and utilizes a design such that increasing dose is displayed by increasing numbers.

(i) Timer.

1. A timer which has a display shall be provided at the treatment control panel. The timer shall have a pre-set time selector and an elapsed time indicator.

2. The timer shall be a cumulative timer which activates with the production of radiation and retains its reading after irradiation is interrupted or terminated. After irradiation is terminated and before irradiation can be reinitiated, it shall be necessary to reset the elapsed time indicator to 0.

3. The time shall terminate irradiation when a preselected time has elapsed if any dose monitoring system present has not previously terminated irradiation.

4. The timer shall permit accurate presetting and determination of exposure times as short as 1 second.

5. The timer shall not activate until the shutter is opened when irradiation is controlled by a shutter mechanism.

(j) Control Panel Functions. The control panel, in addition to the displays required in other provisions of this section, shall have:

1. An indication of whether electrical power is available at the control panel and if activation of the x-ray tube is possible;

2. An indication of whether x-rays are being produced;

3. Means for indicating x-ray tube potential and current;

4. Means for terminating an exposure at any time; and

5. A locking device which will prevent unauthorized use of the x-ray system.

(k) Multiple Tubes. When a control panel may energize more than one x-ray tube:

1. It shall be possible to activate only one x-ray tube at any time;

2. There shall be an indication at the control panel identifying which x-ray tube is energized; and

3. There shall be an indication at the tube housing assembly when that tube is energized.

(l) Source-to-Skin Distance (SSD). There shall be means of determining the SSD to within 1 centimeter.

(m) Low Filtration X-Ray Tubes. Each x-ray system equipped with a beryllium or other low-filtration window shall be clearly labeled as such upon the tube housing assembly and at the control panel.

(2) Facility Design Requirements for X-Ray Systems Operable Above 50 kVp.

(a) Aural Communication. Provision shall be made for two-way aural communication between the patient and the operator at the control panel. However, where excessive noise levels or treatment requirements make aural communication impractical, other methods of communication shall be used.

(b) Viewing Systems. Windows, mirrors, closed-circuit television, or an equivalent system shall be provided to permit continuous observation of the patient during irradiation and shall be so located that the operator can observe the patient from the control panel.

(c) Additional Requirements for X-Ray Systems Operable Above 150 kVp:

1. All protective barriers shall be fixed except for entrance doors or beam interceptors.

2. The control panel shall be located outside the treatment room.

3. Interlocks shall be provided such that all entrance doors must be closed before treatment can be initiated or continued. If the radiation beam is interrupted by any door opening:

a. The exposure at a distance of 1 meter from the source shall be reduced to less than 10 milliroentgens (2.58 uC per kg) per hour; and

b. It shall not be possible to restore the machine to operation without closing the door and reinitiating irradiation by manual action at the control panel.

4. Treatment rooms to which access is possible through more than one entrance shall be provided with flashing warning lights located in a readily observable position near the outside of all access doors to indicate when the useful beam is on.

(3) Surveys, Calibrations, Spot-Checks and Operating Procedures.

(a) Surveys.

1. All new facilities, and existing facilities not previously surveyed, shall have a survey made by, or under the direction of, a qualified person. In addition, such surveys shall be done after any change in the facility or equipment which might cause a significant increase in radiation hazard.

2. The registrant shall transmit a copy of the survey report to the Department within 30 days of the receipt of the report.

3. The survey and report shall indicate all instances where the installation, in the opinion of the qualified person, is in violation of applicable regulations.

(b) Calibrations.

1. The calibration of an x-ray system shall be performed at intervals not to exceed 1 year and after any change or replacement of components which could cause a change in the radiation output.

2. The calibration of the radiation output of an x-ray system shall be performed by or under the direction of a radiological physicist who is physically present at the facility during such calibration.

3. Calibration of the radiation output of an x-ray system shall be performed with a calibrated dosimetry system. The calibration of such system shall be traceable to a national standard. The system shall have been calibrated within the preceding 2 years.

4. The calibration shall be such that the dose at a reference point in soft tissue can be calculated to within an uncertainty not to exceed 5 percent.

5. The calibration of the x-ray system shall include, but not be limited to, the following determinations:

a. Verification that the x-ray system is operating in compliance with the design specifications;

b. The exposure rates as a function of field size, technique factors, filter and treatment distance used;

c. The degree of congruence between the radiation field and the field indicated by the localizing device if such device is present; and

d. An evaluation of the uniformity of the largest radiation field used.

6. Records of calibration shall be maintained by the registrant for 5 years after completion of the calibration.

7. A copy of the most recent x-ray system calibration shall be available at the facility for inspection by the Department.

(c) Spot-Checks. Spot-checks shall be performed on x-ray systems operable at greater than 150 kVp. Such spot-checks shall meet the following requirements:

1. The spot-check procedures shall be in writing and shall have been developed by a radiological physicist.

2. If a radiological physicist does not perform the spot-check measurement, the results of the spot-check measurements shall be reviewed by a radiological physicist within 15 days.

3. The spot-check procedures shall specify the frequency at which tests or measurements are to be performed. The spot-check procedures shall specify that the spot-check shall be performed during the calibration specified in paragraph (3)(b), above. The acceptable tolerance for each parameter measured in the spot-check when compared to the value for that parameter determined in the calibration specified in paragraph (3)(a), above, shall be stated.

4. The cause for a parameter exceeding a tolerance set by the radiological physicist shall be investigated and corrected before the system is used for patient irradiation.

5. Whenever a spot-check indicates a significant change in the operating characteristics of a system, the system shall be recalibrated as required in paragraph (3)(b), above.

6. Records of spot-check measurements and any necessary corrective actions shall be maintained by the registrant for 2 years.

7. Where a spot-check involves a radiation measurement, such measurement shall be obtained using a system satisfying the requirements of paragraph (3)(b), above, or which has been compared within the previous year with a system meeting those requirements.

(d) Operating Procedures.

1. X-ray systems shall not be left unattended unless the system is secured against unauthorized use.

2. When a patient must be held in position for radiation therapy, mechanical supporting or restraining devices shall be used.

3. The tube housing assembly shall not be held by hand during operation unless the system is designed to require such holding and the peak tube potential of the system does not exceed 50 kVp. In such cases, the holder shall wear protective gloves and apron of not less than 0.5 millimeter lead equivalency at 100 kVp.

4. No individual other than the patient shall be in the treatment room unless such individual is protected by a barrier sufficient to meet the requirements of Rule 64E-5.304, F.A.C. No individual other than the patient shall be in the treatment room during exposures from x-ray systems operating above 150 kVp.

5. Machines capable of having an output of more than 1,000 roentgens (258 mC per kg) per minute at any accessible place shall not be left unattended without the power being shut off at the disconnect switch in addition to the control panel switch.

Rulemaking Authority 404.051, 404.081, 404.141, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), 404.081(1), 404.22(1), (3) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.608.

- 64E-5.508 X-Ray and Electron Therapy Systems with Energies of 1 MeV and Above.

(1) Definitions. In addition to the definitions in Rule 64E-5.501, F.A.C., the following definitions shall apply to this section:

(a) “Applicator” means a structure which determines the extent of the treatment field at a given distance from the virtual source.

(b) “Beam scattering filter” means a filter used in order to scatter a beam of electrons.

(c) “Central axis of the beam” means a line passing through the virtual source and the center of the plane figure formed by the edge of the first beam limiting device.

(d) “Dose monitoring system” means a system of devices for the detection, measurement and display of quantities of radiation.

(e) “Dose monitor unit” means a unit response from the dose monitoring system from which the absorbed dose can be calculated.

(f) “Field-flattening filter” means a filter used to provide dose uniformity over the area of a useful beam of x-rays at a specified depth.

(g) “Field size” means the dimensions along the major axes of an area in a plane perpendicular to the specified direction of the beam of incident radiation at the normal treatment distance and defined by the intersection of the major axes and the 50% percent isodose line. Material shall be placed in the beam such that dose maximum is produced at the normal treatment distance when field size is being determined.

(h) “Gantry” means that part of the system supporting and allowing possible movements of the radiation head.

(i) “Interruption of irradiation” means the stopping of irradiation with the possibility of continuing irradiation without resetting of operating conditions at the control panel.

(j) “Isocenter” means a fixed point in space located at the center of the smallest sphere through which the central axis of the beam passes in all conditions.

(k) “Moving beam therapy” means radiation therapy with relative displacement of the useful beam and the patient during irradiation. It includes arc therapy, skip therapy and rotational therapy.

(l) “Normal treatment distance” means:

1. For electron irradiation, the virtual source to surface distance along the central axis of the useful beam as specified by the manufacturer for the applicator.

2. For x-ray irradiation, the virtual source to isocenter distance along the central axis of the useful beam. For nonisocentric equipment, this distance shall be that specified by the manufacturer.

(m) “Radiation head” means the structure from which the useful beam emerges.

(n) “Shadow tray” means a device attached to the radiation head to support auxiliary beam limiting material.

(o) “Stationary beam therapy” means radiation therapy without relative displacement of the useful beam or the patient during irradiation.

(p) “Target” means that part of a radiation head which by design intercepts a beam of accelerated particles with subsequent emission of other radiation.

(q) “Virtual source” means a point from which radiation appears to originate.

(2) Requirements for Equipment.

(a) Leakage Radiation to the Patient Area.

1. New equipment shall meet the following requirement: For operating conditions producing maximum leakage radiation, the absorbed dose in rads (grays) due to leakage radiation, including x-rays, electrons, and neutrons, at any point in a circular plane of 2 meters radius centered on and perpendicular to the central axis of the beam at the isocenter or normal treatment distance and outside the maximum useful beam size, shall not exceed 0.1% percent of the maximum absorbed dose in rads (grays) of the unattenuated useful beam measured at the point of intersection of the central axis of the beam and the plane surface. Measurements, excluding those for neutrons, shall be averaged over an area up to, but not exceeding, 100 square centimeters at the positions specified. Measurements of the portion of the leakage radiation dose contributed by neutrons may be obtained from the manufacturer and shall be averaged over an area up to, but not exceeding, 200 square centimeters.

2. The registrant shall determine or obtain from the manufacturer the leakage radiation existing at the positions specified and for the specified operating conditions. Records on radiation leakage shall be maintained for inspection by the Department.

(b) Leakage of Radiation Outside the Patient Area for New Equipment.

1. The absorbed dose in rads (grays) due to leakage radiation, except in the area specified in paragraph (2)(a), above, when measured at any point 1 meter from the path of the charged particle, before the charged particle strikes the target or window, shall not exceed 0.1 percent for x-ray leakage nor 0.05 percent for neutron leakage of the maximum absorbed dose in rads (grays) of the unattenuated useful beam measured at the point of intersection of the central axis of the beam and the circular plane specified in paragraph (2)(a), above.

2. The registrant shall determine or obtain from the manufacturer the actual leakage radiation existing at the positions specified and for specified operating conditions. Radiation measurements, excluding neutrons, shall be averaged over an area up to, but not exceeding, 100 square centimeters. Neutron measurements shall be averaged over an area up to, but not exceeding, 200 square centimeters.

(c) Beam Limiting Devices. Adjustable or interchangeable beam-limiting devices shall be provided, and such devices shall transmit no more than 5 percent of the useful beam at the normal treatment distance for the portion of the useful beam which is to be attenuated by the beam limiting device. The neutron component of the useful beam shall not be included in this requirement.

(d) Filters.

1. Each filter which is removable from the system shall be clearly marked with an identification number. Documentation available at the control panel shall contain a description of the filter. For wedge filters, the wedge angle shall appear on the wedge or wedge tray.

2. If the absorbed dose rate data indicated at the control panel relates exclusively to operation with a field flattening or beam scattering filter in place, such filter shall be removable only by the use of tools.

3. For new equipment which utilizes a system of wedge filters, interchangeable field flattening filters or interchangeable beam scattering filters:

a. Irradiation shall not be possible until a selection of a filter has been made at the treatment control panel;

b. An interlock system shall be provided to prevent irradiation if the filter selected is not in the correct position;

c. A display shall be provided at the treatment control panel showing the filter in use; and

d. An interlock shall be provided to prevent irradiation if any filter selection operation carried out in the treatment room does not agree with the filter selection operation carried out at the treatment control panel.

(e) Beam Quality. The registrant shall determine, or obtain from the manufacturer, data sufficient to assure that the following beam quality requirements are met:

1. The absorbed dose resulting from x-rays in a useful electron beam at a point on the central axis of the beam 10 centimeters greater than the practical range of the electrons shall not exceed the values stated below. Linear interpolation shall be used for values not stated.

X-Ray Absorbed Dose

Maximum Energy of				as a Fraction of

Electron Beam				Maximum Absorbed

in MeV					Dose

1					0.03

15					0.05

35					0.10

50					0.20

2. Compliance with subparagraph (2)(e)1., above, shall be determined using:

a. A measurement within a phantom with the incident surface of the phantom at the normal treatment distance and normal to the central axis of the beam:

b. The largest field size available which does not exceed 15 by 15 centimeters; and

c. A phantom whose cross-sectional dimensions exceed the measurement radiation field by at least 5 centimeters and whose depth is sufficient to perform the required measurement.

(f) Beam Monitors. All therapy systems shall be provided with two radiation detectors in the radiation head.

1. The detectors shall be incorporated into two separate dose monitoring systems.

2. The detector and the system into which that detector is incorporated shall meet the following requirements:

a. Each detector shall be removable only with tools and shall be designed to prevent incorrect positioning.

b. Each detector shall form part of a dose monitoring system from whose readings in dose monitor units the absorbed dose at a reference point in the treatment volume can be calculated.

c. Each dose monitoring system shall be capable of independently monitoring, interrupting and terminating irradiation.

d. For new equipment, the design of the dose monitoring systems shall assure that:

(I) The malfunctioning of one system shall not affect the correct functioning of the second system; and

(II) The failure of any element common to both systems which could affect the correct function of both systems shall terminate irradiation.

3. Each dose monitoring system shall have a legible display at the treatment control panel. For new equipment, each display shall:

a. Maintain a reading until intentionally reset to 0;

b. Have only one scale and no scale multiplying factors;

c. Utilize a design such that increasing dose is displayed by increasing numbers and in the event of an overdosage of radiation, the absorbed dose may be accurately determined; and,

d. In the event of power failure, the dose monitoring information required to be displayed at the control panel at the time of failure shall be retrievable in at least one system for a 20-minute period of time.

(g) Beam Symmetry. In equipment inherently capable of producing useful beams with asymmetry exceeding 5% percent, the asymmetry of the radiation beam in two orthogonal directions shall be monitored before the beam passes through the beam limiting device. Facilities shall be provided so that, if the difference in dose rate between one region and another region symmetrically displaced from the central axis of the beam exceeds 5% percent of the central axis dose rate, indication of this condition is made at the control panel; and if this difference exceeds 10% percent, the irradiation is terminated.

(h) Selection and Display of Dose Monitor Units.

1. Irradiation shall not be possible until a selection of a number of dose monitor units or exposure time has been made at the treatment control panel.

2. The pre-selected number of dose monitor units or exposure time shall be displayed at the treatment control panel until reset manually for the next irradiation.

3. After termination of irradiation, it shall be necessary to reset the dosimeter display to 0 before subsequent treatment can be initiated.

4. After termination of irradiation, it shall be necessary to manually reset the pre-selected dose monitor units before irradiation can be initiated.

(i) Termination of Irradiation by the Dose Monitoring System or Systems During Stationary Beam Therapy.

1. Each primary system shall terminate irradiation when the pre-selected number of dose monitor units has been detected by the system.

2. The system shall be capable of terminating irradiation when not more than 10% percent or 30 dose monitoring units above the pre-selected number of dose monitor units set at the control panel has been detected by the second dose monitoring system.

3. An indicator on the control panel shall show which dose monitoring system has terminated irradiation.

(j) Interruption Switches. It shall be possible to interrupt irradiation and equipment movements at any time from the operator’s position at the treatment control panel. Following an interruption, it shall be possible to restart irradiation by operator action without any reselection of operating conditions. If any change is made of a pre-selected value during an interruption, irradiation and equipment movements shall be automatically terminated.

(k) Termination Switches. It shall be possible to terminate irradiation and equipment movements, or go from an interruption condition to termination conditions, at any time from the operator’s position at the treatment control panel.

(l) Timer.

1. A timer which has a display shall be provided at the treatment control panel. The timer shall have a pre-set time selector and an elapsed time indicator.

2. The timer shall be a cumulative timer which activates with the production of radiation and retains its reading after irradiation is interrupted or terminated. After irradiation is terminated and before irradiation can be reinitiated, it shall be necessary to reset the elapsed time indicator to 0.

3. After termination of irradiation and before irradiation can be reinitiated, it shall be necessary to manually reset the preset time selector.

4. The timer shall terminate irradiation when a pre-selected time has elapsed if the dose monitoring systems have not previously terminated irradiation.

(m) Selection of Radiation Type. Equipment capable of both x-ray therapy and electron therapy shall meet the following:

1. Irradiation shall not be possible until a selection of radiation type has been made at the treatment control panel.

2. An interlock system shall be provided to ensure that the equipment can emit only the radiation type which has been selected.

3. An interlock system shall be provided to prevent irradiation if any selected operations carried out in the treatment room do not agree with the selected operations carried out at the treatment control panel.

4. An interlock system shall be provided to prevent irradiation with x-rays except to obtain a port film when electron applicators are fitted.

5. An interlock system shall be provided to prevent irradiation with electrons when accessories specific for x-ray therapy are fitted.

6. The radiation type selected shall be displayed at the treatment control panel before and during irradiation.

(n) Selection of Energy. Equipment capable of generating radiation beams of different energies shall meet the following:

1. Irradiation shall not be possible until a selection of energy has been made at the treatment control panel.

2. An interlock system shall be provided to prevent irradiation if any selected operations carried out in the treatment room do not agree with the selected operations carried out at the treatment control panel.

3. The nominal energy value selected shall be displayed at the treatment control panel before and during irradiation.

(o) Selection of Stationary Beam Therapy or Moving Beam Therapy. Equipment capable of both stationary beam therapy and moving beam therapy shall meet the following:

1. Irradiation shall not be possible until a selection of stationary beam therapy or moving beam therapy has been made at the treatment control panel.

2. An interlock system shall be provided to ensure that the equipment can operate only in the mode which has been selected.

3. An interlock system shall be provided to prevent irradiation if any selected operations carried out in the treatment room do not agree with the selected operations carried out at the treatment control panel.

4. The mode of operation shall be displayed at the treatment control panel.

5. An interlock system shall be provided to terminate irradiation if:

a. Movement of the gantry occurs during stationary beam therapy; or

b. Movement of the gantry stops during moving beam therapy unless such stoppage is a preplanned function.

6. Moving beam therapy shall be controlled to obtain the selected relationships between incremental dose monitor units and incremental angle of movement.

a. An interlock system shall be provided to terminate irradiation if the number of dose monitor units delivered in any 10o degrees of arc differs by more than 20% percent from the selected value.

b. Where gantry angle terminates the irradiation in arc therapy, the dose monitor units shall differ by less than 5% percent from the value calculated from the absorbed dose per unit angle relationship.

7. Where the dose monitor system terminates the irradiation in arc therapy, the termination of irradiation shall be as required by paragraph (2)(i), above.

(p) Absorbed Dose Rate. A system shall be provided from whose readings the absorbed dose rate at a reference point in the treatment volume can be calculated. In addition:

1. The dose monitor unit rate shall be displayed at the treatment control panel; and,

2. The radiation detectors specified in paragraph (2)(f), above, may form part of this system.

(q) Location of Virtual Source and Beam Orientation. The registrant shall determine, or obtain from the manufacturer, the location with reference to an accessible point on the radiation head of:

1. The x-ray target or the virtual source of x-rays; and

2. The electron window or the virtual source of electrons if the system has a electron beam capabilities.

(r) System Checking Facilities. Capabilities shall be provided so that all radiation safety interlocks can be checked for correct operation. When preselection of any of the operating conditions requires action in the treatment room and at the treatment control panel, selection at one location shall not give a display at the other location until the requisite selected operations in both locations have been completed.

(3) Facility and Shielding Requirements. In addition to Part III, the following design requirements shall apply:

(a) Protective Barriers. All protective barriers shall be fixed except for entrance doors or beam interceptors.

(b) Control Panel. The control panel shall be located outside the treatment room.

(c) Viewing System. Windows, mirrors, closed-circuit television, or an equivalent system shall be provided to permit continuous observation of the patient during irradiation and shall be so located that the operator may observe the patient from the control panel.

(d) Aural Communication. Provision shall be made for two-way aural communication between the patient and the operator at the control panel. However, where excessive noise levels or treatment requirements make aural communication impractical, other methods of communication shall be used.

(e) Room Entrance. Treatment room entrances shall be provided with warning lights in readily observable positions near the outside of all access doors or other entrances to indicate when the useful beam is on.

(f) Entrance Interlocks. Interlocks shall be provided such that all entrance doors must be closed before treatment can be initiated or continued. If the radiation beam is interrupted by any barrier penetration or door opening, it shall not be possible to restore the machine to operation without closing the door and reinitiating irradiation by manual action at the control panel.

(4) Surveys, Calibrations, Spot-Checks and Operating Procedures.

(a) Surveys.

1. All new facilities, and existing facilities not previously surveyed, shall have a survey made by, or under the direction of, a qualified person. In addition, such surveys shall be done after any change in the facility or equipment which might cause a significant increase in radiation hazard.

2. The registrant shall obtain a written report of the survey from the qualified person, and a copy of the report shall be transmitted by the registrant to the Department within 30 days of receipt of the report.

3. The survey and report shall indicate all instances where the installation, in the opinion of the qualified person, is in violation of applicable regulations.

(b) Calibration.

1. The calibration of systems subject to Rule 64E-5.508, F.A.C., shall be performed in accordance with an established calibration protocol acceptable to the Department, such as the calibration protocol published by the American Association of Physicists in Medicine, before the system is first used for irradiation of a patient and thereafter at time intervals which do not exceed 12 months and after any change which might significantly alter the calibration, spatial distribution or other characteristics of the therapy beam.

2. The calibration shall be performed under the direct supervision of a radiological physicist who is physically present at the facility during the calibration.

3. Calibration radiation measurements required by paragraph (4)(b), above, shall be performed using a dosimetry system:

a. Having a calibration factor for cobalt 60 gamma rays traceable to a national standard;

b. Which has been calibrated within the previous 2 years and after any servicing that may have affected its calibration;

c. Which has been calibrated in such a fashion that an uncertainty can be stated for the radiation quantities monitored by the system; and

d. Which has had constancy checks performed on the system as specified by a radiological physicist.

4. Calibrations shall be in sufficient detail that the dose at a reference point in soft tissue may be calculated to within an uncertainty of 5 percent.

5. The calibration of the therapy beam shall include the following determinations:

a. Verification that the equipment is operating in compliance with the design specifications concerning the light localizer, side light, and back-pointer alignment with the isocenter when applicable, variation in the axis of rotation for the table, gantry, and jaw system, and beam flatness and symmetry at the specified depth.

b. The absorbed dose rate at various depths of water for the range of field sizes used, for each effective energy, that will verify the accuracy of the dosimetry of all therapy procedures utilized with that therapy beam.

c. The uniformity of the radiation field and any dependency upon the direction of the useful beam.

d. Verification that existing depth-dose data and isodose charts applicable to the specific machine continue to be valid or are updated to existing machine conditions.

e. Verification of transmission and electron buildup factors for all accessories such as wedges, shadow trays, and compensators.

6. Records of calibration measurements and dosimetry system calibrations required in paragraph (4)(b), above, shall be maintained for 5 years after completion of the full calibration.

7. A copy of the latest calibration performed shall be available in the facility for inspection by the Department.

(c) Spot-Checks. Spot-checks shall be performed on systems subject to this section during calibrations and thereafter at intervals not to exceed 1 month. Such spot-checks shall meet the following requirements:

1. The spot-check procedures shall be in writing and shall have been developed by a radiological physicist. Acceptable tolerance for each parameter measured in the spot-check shall not exceed manufacturer’s recommendations.

2. If a radiological physicist does not perform the spot-check measurements, the results of the spot-check measurements shall be reviewed by a radiological physicist within 15 days.

3. The spot-check procedures shall specify the frequency at which tests or measurements are to be performed and the acceptable tolerance for each parameter measured in the spot-check when compared to the value for that parameter determined in the calibration.

4. At intervals established in the spot-check procedures, spot-checks shall be made of absorbed dose measurements at a minimum of 2 depths in a phantom.

5. Where a system has built-in devices which provide a measurement of any parameter during irradiation, such measurement shall not be utilized as a spot-check measurement.

6. The cause for a parameter exceeding a tolerance set by the radiological physicist shall be investigated and corrected before the system is used for patient irradiation.

7. Whenever a spot-check indicates a significant change in the operating characteristics of a system, as specified in the radiological physicist’s spot-check procedures, the system shall be recalibrated as required in paragraph (4)(b), above.

8. Records of spot-check measurements and any necessary corrective actions shall be maintained by the registrant for a period of 2 years.

9. Where a spot-check involves an absolute radiation measurement, such measurement shall be obtained using a system satisfying the requirements of subparagraph (4)(b)3., above, or which has been compared with a system meeting those requirements within the previous year.

(d) Additional Operating Procedures.

1. No individual other than the patient shall be in the treatment room during treatment of a patient.

2. If a patient must be held in position during treatment, mechanical supporting or restraining devices shall be used.

3. The system shall not be used in the administration of radiation therapy unless the requirements of paragraphs (4)(a), (4)(b) and (4)(c), above, have been met.

Rulemaking Authority 404.031, 404.051, 404.071, 404.081, 404.141, 404.22 FS. Law Implemented 404.022, 404.031, 404.051(1), (4), (5), (6), 404.071(1), 404.081(1), 404.141, 404.22(1), (3) FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.609, Amended 6-3-15.

- 64E-5.509 Veterinary Medicine X-Ray Operations.

(1) Applicable Regulations. Veterinary medical x-ray operations shall conform with requirements of the following sections of these regulations:

(a) Rule 64E-5.502, F.A.C., General Requirements, except subparagraphs 64E-5.502(1)(a)5., 6., 7., 8., F.A.C.

(b) Rule 64E-5.503, F.A.C., General Requirements for all Diagnostic X-Ray Systems.

(c) Rule 64E-5.504, F.A.C., Fluoroscopic X-Ray Systems.

(2) Additional Requirements.

(a) Positive means of beam alignment shall be provided in the form of accurate linear rulings, beam defining or beam centering lights, optical viewing devices or the equivalent. Such alignment means or devices shall be adjusted to indicate the beam center or beam area to within 2 percent of the SID.

(b) Means shall be provided to limit the useful beam to the area of diagnostic interest or to the area of the image receptor used in each particular case. Beam limitation may be accomplished by any of the means described in subsection 64E-5.505(1), F.A.C.

(c) Each x-ray system shall be equipped with a device which will terminate the exposure after a preset time or exposure.

(d) Each exposure switch shall be of the dead-man type.

(e) Each exposure switch shall be located in such a way as to meet the following criteria:

1. The operator shall stand as far as practicable and at least 6 feet (1.8 m) from the animal and tube head and outside the useful beam or behind a protective barrier during exposures.

2. In lieu of distance or a protective barrier the operator shall wear a protective apron and monitoring device as provided in (3)(c), below.

(3) Operating Procedures.

(a) The operator shall stand in a protected position as indicated in paragraph (2)(e), above, during radiographic exposures with no other individuals in the x-ray room unless assistance of the nature described in paragraph (3)(c), below, is required.

(b) To the greatest practicable extent, animals must be immobilized by anesthetic, straps, sandbags, foam wedges, and other supporting or restraining devices.

(c) If an animal must be held by an individual, that individual shall be protected by appropriate shielding devices such as a protective apron and gloves, and the holder shall be so positioned that no part of his body will be struck by the useful beam. The exposure of that individual shall be monitored when engaged in such purposes.

Rulemaking Authority 404.051, 404.141, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.141, 404.22(1),(3) FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.610.

- 64E-5.510 Mammographic Systems.

(1) Mammographic medical x-ray systems shall meet the requirements of Rules 64E-5.502 and 64E-5.503, F.A.C. Registrants who provide mammography services shall:

(a) Have a written quality assurance program specific to mammography imaging that includes an equipment quality control program for performance monitoring and an evaluation of all components of the equipment from the x-ray generator to the image processor.

(b) Establish standards for clinical image evaluations that include breast positioning, compression and overall image quality.

(c) Assign qualified and trained personnel to each part of the quality assurance program.

(d) Conduct a general review of the effectiveness of the quality assurance program annually and maintain a written report of the review.

(e) Have available the services of a medical physicist to furnish diagnostic x-ray physics support who is able to establish and conduct the equipment quality control program and who meets the requirements specified in subsection (12), below. The specific duties of the medical physicist must include:

1. Monitoring equipment performance or verifying the qualifications and training of others to monitor equipment performance.

2. Evaluating the monitoring results to identify problems.

3. Verifying that corrections are effective and meet regulatory requirements.

(2) The film processor shall be optimized for the specific mammography film used by the facility. Its performance shall be checked for consistency of speed, contrast, and base plus fog prior to processing patient films and after being idle more than 6 hours.

(a) Performance checks shall be plotted and compared to established limits. If these limits are exceeded, documented corrective actions including an image quality check as specified in subsection (3), below, are required.

(b) Corrective action shall be taken when:

1. Optical density deviates by more than 0.15 from established operating levels for readings of mid-density or density difference on the sensitometric control charts.

2. Base plus fog exceeds the established operating level by more than 0.03 optical density.

(c) These records for processor optimization, performance, image quality checks and documented corrective actions shall be maintained for inspection by the department for at least 1 year.

(3) Mammographic x-ray systems shall be monitored and evaluated using the following standards:

(a) The image quality shall be checked using a standard phantom approved by the FDA which meets the criteria below at least weekly and whenever service which could affect image quality is performed on the x-ray system or the film processor. The image quality shall be scored on the ability to image fibers, specks, and low density masses. If quality control limits are exceeded, image quality checks also must be performed after any corrective actions have been taken. This standard phantom must be designed to evaluate image quality in the 1.2 to 1.8 optical density range, shall not change more than 0.2 optical density from its previous reading, and must be composed of material that is equivalent to a nominal 4.5 centimeter compressed breast of average density of approximately 50 percent adipose and 50% percent glandular tissue. It shall contain the following objects:

1. Nylon fibers with thicknesses of 1.56 millimeters, 1.12 millimeters, 0.89 millimeters, 0.75 millimeters, 0.54 millimeters, and 0.40 millimeters.

2. Aluminum oxide specks with diameters of 0.54 millimeters, 0.40 millimeters, 0.32 millimeters, 0.24 millimeters, and 0.16 millimeters.

3. Phenolic plastic spherical masses with thicknesses of 2.00 millimeters, 1.00 millimeters, 0.75 millimeters, 0.50 millimeters, and 0.25 millimeters.

(b) Phantom checks which indicate a decrease in image quality shall require immediate investigation of possible corrective actions.

(c) The minimum acceptable image quality of a standard phantom described in paragraph (3)(a), above, shall demonstrate the ability to image at least 0.75 millimeter fibers; 0.32 millimeter specks; and 0.75 millimeter spherical masses. Mammographic examinations shall not be performed on systems which do not meet the minimum image quality standard.

(d) The registrant must document in the annual review required in subsection (1), above, that all equipment quality control items were performed under the direction and approval of the medical physicist.

(e) Mammography system performance must be evaluated regularly. Those components and parameters of the equipment quality control program tested for performance daily, weekly, monthly or quarterly shall be evaluated quarterly. The annual onsite survey by the medical physicist must include a summary of the quarterly evaluations.

(4) The following requirements apply to personnel involved in any aspect of mammography, including the production, processing, and interpretation of mammograms and related quality assurance activities.

(a) Interpreting physicians shall be licensed to practice medicine in the State of Florida, as specified in Chapters 458 and 459, F.S.

(b) Radiologic technologists shall be certified as a general radiographer in the State of Florida as specified in Chapter 64E-3, F.A.C.

(c) A medical physicist qualified to conduct surveys of mammography facilities and provide oversight of the facility quality assurance program shall be licensed in Florida as a medical physicist as specified in Chapter 483, F.S.

(5) Documentation, records and surveys. Each facility shall maintain records, policies, procedures and documentation to demonstrate compliance with these requirements, including corrective actions taken.

(a) Clinical images. Each facility shall establish and maintain a clinical image quality control program, including:

1. Monitoring of mammograms repeated because of poor image quality; and

2. Maintaining records, analysis of results, and a description of any remedial action taken as a result of this monitoring.

(b) Clinical image interpretation. To ensure that quality clinical images are produced routinely at the facility, each facility shall submit clinical images to the Department for review as required by the Department.

(c) Surveys. A medical physicist who meets the qualifications specified in subsection (4), above, and who establishes, monitors, evaluates, and directs the equipment quality control program must perform an on-site survey of the facility to assure that it meets quality control and equipment standards. These surveys shall be performed at least annually and shall be available for inspection by the department. Each survey report shall be retained by the facility until the next annual survey is completed satisfactorily.

(6) In addition to the above requirements, effective October 1, 1994, no facility can conduct mammography procedures unless the facility also obtains and maintains a certificate issued by the FDA as described in Public Law 102-539, the Mammography Quality Standards Act of 1992, and complies with  all  requirements of 21 CFR Part 900, April 1, 2014 edition, which is incorporated herein by reference and available from the internet at http://www.gpo.gov/fdsys/pkg/CFR-2014-title21-vol8/pdf/CFR-2014-title21-vol8-chapI-subchapI.pdf, and at https://www.flrules.org/Gateway/reference.asp?No=Ref-05440.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.051(1), (4), 404.141, 404.22(1), (3), (6) FS. History–New 3-17-92, Amended 1-1-94, 11-20-94, Formerly 10D-91.611, Amended 5-18-98, 6-3-15.

- 64E-5.511 Registration of Radiation Machines.

(1) Exemptions.

(a) Electronic equipment that produces radiation incidental to its operation for other purposes is exempt from registration registration requirements if the dose equivalent rate averaged over an area of 10 square centimeters does not exceed 0.5 mR (5 µSv) per hour at 5 centimeters from any accessible surface of the equipment. The production, testing or factory servicing of such equipment shall not be exempt.

(b) Radiation machines that are non-operational and under the control of a registered vendor prior to final installation are exempt from the registration and fee requirements of this section.

(2) Application and Fees for Registration of Radiation Machines.

(a) Each person who acquires a radiation machine or an additional radiation machine shall:

1. Apply for registration of the radiation machine with the Department within 30 days after acquisition and before use. Application for registration shall be on DH Form 1107, 10/15, “Radiation Machine Facility Registration,” which is herein incorporated by reference and available from the internet at http://www.floridahealth.gov/radiation, or at https://www.flrules.org/Gateway/reference.asp?No=Ref-06528.

2. Designate an individual who will be responsible for radiation protection.

3. Prohibit any person who is not registered with the department as a provider of services as specified in (3), below, from furnishing radiation machine servicing or services to his radiation machine.

(b) Registration fees are due within 30 days after acquiring a radiation machine. If the machine is acquired within 120 days before the October 28 annual renewal date, the registration fee will be due on or before October 28 and will be the annual renewal fee. Otherwise, the renewal fee is due annually on or before October 28.

(c) An annual fee for the registration and inspection of radiation machines shall be paid according to the following schedule:

First Tube/Unit				Each Additional Tube/Unit

Medical or Chiropractic or Osteopathic		$145					$85

Veterinary One Tube		$50					$34

Educational or Industrial		$47					$23

Dental or Podiatry		$31					$11

Medical Accelerator		$258					$148

Non-Medical Accelerator		$81					$48

(3) Application for Registration of Servicing and Services.

(a) Each person who installs or offers to install radiation machines or furnishes or offers to furnish radiation machine servicing or services in Florida shall apply to the Department to register such services before furnishing or offering to furnish such services.

(b) Application for registration shall be completed on DH Form 1113, 09/14, “Radiation Machine Vendor Registration Form,” which is herein incorporated by reference and is available from the internet at http://www.floridahealth.gov/radiation, and at https://www.flrules.org/Gateway/reference.asp?No=Ref-05442.

(c) Services include the installation or servicing of radiation machines and associated radiation machine components.

(4) Report of Changes. The registrant shall report in writing within 30 days any changes to the information in the Certificate of Registration. The report shall include name, address of installation change, receipt, sale, transfer, or disposal of any radiation machine or major component.

(5) Assembler or Transferor Obligation.

(a)1. Any registrant or person who sells, leases, transfers, relocates, lends or disposes of one or more radiation machines or major components of one or more such machines shall notify the Department within 15 days after such action. Notification shall be made on DH Form 1107, 10/15, “Radiation Machine Facility Registration.”

2. Any assembler who relocates, assembles, installs or disposes of one or more radiation machines or major components of one or more such machines shall notify the Department within 15 days after such action. Notification shall be made on DH Form 1114, 10/15, “Report of Assembly of Certified or Non-Certified X-Ray Systems,” which is herein incorporated by reference and available from the internet at http://www.floridahealth.gov/radiation and at https://www.flrules.org/Gateway/reference.asp?No=Ref-06529, or a similar form which captures all the information required on DH Form 1114, 10/15.

(b) No person shall sell, offer to sell, lease, transfer, lend or install radiation machines unless such machines meet the requirements of these regulations.

(6) Out-of-State Radiation Machines.

(a) Any person proposing to bring a radiation machine into Florida shall notify the Department in writing at least 10 days before the machine is to be used in the state. The notice shall include the type of radiation machine; the nature, duration and scope of use; and the exact location where the radiation machine will be used.

(b) Any person proposing to bring a radiation machine into Florida shall register the machine with the department and pay the registration fee.

(c) Any out-of-state person using a radiation machine in Florida shall notify the department when the use of the machine has been completed.

(7) Enforcement. The General Statement of Policy and Procedure for Radiation Machine Enforcement Actions, September 2014, which is herein incorporated by reference, will be used to determine enforcement actions to be taken. This publication can be obtained from the internet at http://www.floridahealth.gov/radiation, and at https://www.flrules.org/Gateway/reference.asp?No=Ref-05444.

Rulemaking Authority 404.051, 404.131, 404.22 FS. Law Implemented 404.071, 404.091, 404.101, 404.131, 404.141, 404.161, 404.162, 404.163, 404.22 FS. History–New 12-12-96, Formerly 10D-91.612, Amended 8-16-07, 6-3-15, 3-21-16.

PART VI USE OF RADIONUCLIDES IN THE HEALING ARTS

- 64E-5.601 License Required.

(1) Radioactive materials shall not be manufactured, produced, acquired, received, possessed, prepared, used, or transferred for medical use except as provided in a specific license.

(2) Any licensee who is licensed for one or more of the medical uses in Rule 64E-5.626, 64E-5.627, 64E-5.630, or 64E-5.632, F.A.C., also is authorized to use radioactive material under a general license in subsection 64E-5.206(8), F.A.C., for specified in vitro uses without filing the certificate required by paragraph 64E-5.206(8)(b), F.A.C., but is subject to the other provisions of subsection 64E-5.206(8), F.A.C.

(3)(a) Unless prohibited by license condition, a physician, in training may receive, possess, acquire, prepare, use, or transfer radioactive materials as provided in these regulations under the supervision of an authorized user as provided in subsections 64E-5.608(1) and 64E-5.608(3), F.A.C.

(b) Current and active certified radiologic technologists as authorized in Part IV, Chapter 468, F.S., may receive, possess, acquire, prepare, use, or transfer radioactive materials as provided in these regulations under the supervision of an authorized user as provided in paragraph 64E-5.607(3)(e) and subsection 64E-5.608(3), F.A.C.

(c) Unless prohibited by license condition, a medical physicist in training may receive, acquire, prepare, use, possess, or transfer radioactive materials as provided in these regulations under the supervision of an authorized medical physicist as provided in subsections 64E-5.608(2) and 64E-5.608(3), F.A.C.

(4) Unless authorized by the department, no individual shall manufacture, produce, acquire, receive, possess, prepare, use, or transfer radioactive materials for medical use unless:

(a) That individual is listed on the licensee’s specific license as an authorized user, authorized medical physicist, or an authorized nuclear pharmacist;

(b) Authorized by Rule 64E-5.609, F.A.C.;

(c) Authorized by subsection 64E-5.601(2), F.A.C., with approval of the radiation safety committee at medical institutions or by management for licensees that are not medical institutions; or

(d) That individual is in training, authorized by subsection 64E-5.601(3), F.A.C., and subpart I of Part VI.

(5) Provisions for the protection of human research subjects are:

(a) A licensee may conduct research involving human research subjects only if it uses the radioactive materials specified on its license for the uses authorized on its license.

(b) If the research is conducted, funded, supported, or regulated by another federal agency that has implemented the “Federal Policy for the Protection of Human Subjects (Federal Policy)”, as described in 45 CFR Part 46, dated 11/9/2009, which is herein incorporated by reference, and may be accessed at http://www.doh.state.fl.us/environment/radiation/, or requested in writing from the Department of Health, Bureau of Radiation Control, Bin #C21, 4052 Bald Cypress Way, Tallahassee, FL 32399-1741, the licensee shall, before conducting research:

1. Obtain review and approval of the research from an “Institutional Review Board (IRB),” as defined and described in the Federal Policy; and

2. Obtain “informed consent,” as defined and described in the Federal Policy, from the human research subject.

(c) If the research will not be conducted, funded, supported, or regulated by another federal agency that has implemented the Federal Policy, the licensee shall, before conducting research, apply for and receive a specific amendment to its radioactive materials medical use license. The amendment request must include a written commitment that the licensee will, before conducting research:

1. Obtain review and approval of the research from an IRB as defined and described in the Federal Policy; and

2. Obtain “informed consent”, as defined and described in the Federal Policy, from the human research subject.

(d) Nothing in this section relieves licensees from complying with the other requirements in this part.

(6) Authorized nuclear pharmacists must be actively licensed as a nuclear pharmacist by the Department of Health, Division of Medical Quality Assurance as specified in Rule 64B16-28.903, F.A.C., and authorized medical physicists must have an active medical physicist license, in the area they are practicing, issued by the Department of Health, Division of Medical Quality Assurance.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141, 483.901 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.707, Amended 8-6-01, 2-11-10.

- 64E-5.6011 Definitions.

(1) “Authorized medical physicist” means an individual who meets the requirements:

(a) Specified in subsection 64E-5.656(1) and Rule 64E-5.658, F.A.C.; or

(b) Is identified as an authorized medical physicist or teletherapy physicist on:

1. A specific medical use license issued by the NRC or an agreement state;

2. A medical use permit issued by a NRC master material licensee;

3. A permit issued by a NRC or agreement state broad scope medical use licensee; or

4. A permit issued by a NRC master material license broad scope medical use permittee.

(2) “Authorized user” means:

(a) A physician, dentist, or podiatrist who meets the requirements in Rule 64E-5.658 and subsection 64E-5.649(1), 64E-5.660(1), 64E-5.661(1), 64E-5.662(1), 64E-5.652(1), 64E-5.654(1) or 64E-5.655(1), F.A.C.; or

(b) An individual identified for medical use of radioactive materials on:

1. A NRC or agreement state license that authorizes the medical use of radioactive material;

2. A permit issued by a NRC master material licensee that is authorized to permit the medical use of radioactive material;

3. A permit issued by a NRC or agreement state specific licensee of broad scope that is authorized to permit the medical use of radioactive material; or

4. A permit issued by a NRC master material license broad scope permittee that is authorized to permit the medical use of radioactive material.

(3) “Brachytherapy” means a method of radiation therapy in which sources are used to deliver a radiation dose by surface, intracavitary, intralumimnal or interstitial application.

(4) “Brachytherapy source” means a radioactive source or a manufacturer-assembled source train or a combination of these sources that is designed to deliver a therapeutic dose within a distance of a few centimeters.

(5) “Diagnostic clinical procedures manual” means a collection of written procedures that describes each method by which the licensee shall perform diagnostic clinical procedures, and provides other instructions and precautions related thereto. Each diagnostic clinical procedure shall be approved by the authorized user and shall include the radiopharmaceutical, dosage, and route of administration.

(6) “High dose-rate remote afterloader,” as used in this part, means a brachytherapy device that remotely delivers a dose rate in excess of 12 gray (1200 rads) per hour at the point or surface where the dose is prescribed.

(7) “Low dose-rate remote afterloader,” as used in this part, means a brachytherapy device that remotely delivers a dose rate of less than or equal to 2 gray (200 rads) per hour at the point or surface where the dose is prescribed.

(8) “Manual brachytherapy,” as used in this part, means a type of brachytherapy in which the brachytherapy sources (e.g., seeds, ribbons) are manually delivered.

(9) “Medical use” means the intentional internal or external administration of radioactive material, or the radiation therefrom, to patients or humans research subjects under the supervision of an authorized user.

(10) “Medium dose-rate remote afterloader,” as used in this part, means a brachytherapy device that remotely delivers a dose rate of greater than 2 gray (200 rads) per hour, but less than or equal to 12 gray (1200 rads) per hour at the point or surface where the dose is prescribed.

(11) “Mobile medical service” means the ability to transport and use radioactive materials for medical use at the client’s address.

(12) “Output” means the exposure rate, dose rate, or a quantity related in a known manner to these rates from a brachytherapy source or a teletherapy, remote afterloader, or gamma stereotactic radiosurgery unit for a specified set of exposure conditions.

(13) “Preceptor” means an individual who provides, directs, or verifies training and experience required for an individual to become an authorized user under Chapter 64E-5, Part VI, F.A.C., an authorized medical physicist, an authorized nuclear pharmacist or a RSO under Chapter 64E-5 Part VI, F.A.C.

(14) “Pulsed dose-rate remote afterloader,” as used in this part, means a special type of remote afterloading brachytherapy device that uses a single source capable of delivering dose rates in the “high dose-rate” range, provided that the source is:

(a) Approximately one-tenth of the activity of typical high dose-rate remote afterloader sources; and

(b) Used to simulate the radiobiology of a low dose-rate treatment by inserting the source for a given fraction of each hour.

(15) “Radiation Safety Officer” or “RSO” means an individual who:

(a) Meets the requirements in subsection 64E-5.648(1) or paragraph 64E-5.648(3)(a) and Rule 64E-5.658, F.A.C.; or

(b). Is identified as a RSO on a specific medical use license issued by the NRC or an agreement state or a medical use permit issued by a NRC master material licensee.

(16) “Teletherapy physicist” means an individual identified as the qualified teletherapy physicist on a department license.

(17) “Therapeutic dosage” means a dosage of unsealed radioactive materials that is intended to deliver a radiation dose to a patient or human research subject for palliative or curative treatment.

(18) “Therapeutic dose” means a radiation dose delivered from a source containing radioactive materials to a patient or human research subject for palliative or curative treatment.

(19) “Treatment site” means the anatomical description of the tissue intended to receive a radiation dose, as described in a written directive.

(20) “Unit dosage” means a dosage prepared for medical use for administration as a single dosage to a patient or human research subject without any further manipulation of the dosage after it is initially prepared.

Rulemaking Authority 404.051, 404.061 FS. Law Implemented 404.031, 404.061(2), 404.20, 404.22, 404.30 FS. History–New 2-11-10, Amended 12-26-13.

- 64E-5.602 License Amendments.

A licensee shall apply for and receive a license amendment or departmental approval:

(1) Before using radioactive material for a method or type of medical use not permitted by the license;

(2) Before permitting anyone, except a visiting authorized user, visiting authorized medical physicist, or visiting authorized nuclear pharmacist described in Rule 64E-5.609, F.A.C., to work as an authorized user, authorized nuclear pharmacist, or authorized medical physicist.

(3) Before changing a RSO or authorized medical physicist;

(4) Before ordering or receiving radioactive material in excess of the amount, in a different form, or receiving a different radionuclide than is authorized on the license;

(5) Before adding to or changing the areas of use or address or addresses of use identified in the application or on the license; and

(6) Before changing statements, representations, and procedures which are incorporated into the license.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.708, Amended 2-11-10.

- 64E-5.603 Notification.

A licensee shall notify the department in writing within 30 days when the licensee changes its mailing address or when an authorized user, RSO, authorized nuclear pharmacist, or authorized medical physicist permanently discontinues performance of their duties under the licensee.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.709, Amended 8-6-01, 2-11-10.

PART VI USE OF RADIONUCLIDES IN THE HEALING ARTS

SUBPART A GENERAL ADMINISTRATIVE REQUIREMENTS

- 64E-5.604 ALARA Program.

(1) Each licensee shall develop and implement a written program to maintain radiation doses and releases of radioactive material in effluents to unrestricted areas as low as reasonably achievable as provided in Rule 64E-5.303, F.A.C.

(2) The management, radiation safety officer, and all authorized users shall participate in the establishment, implementation, and operation of the program as required by these regulations or the radiation safety committee.

(3) For licensees that are not required to have a radiation safety committee, medical institutions, management and all authorized users shall participate in the program as required by the RSO.

(4) The ALARA program shall include an annual review by the radiation safety committee for medical licensees required to have a radiation safety committee, or by management and the RSO for licensees that are not required to have a radiation safety committee. The review shall include summaries of the types, amounts and purposes of radioactive material used; occupational dose reports; and continuing education and training of all personnel who work with or in the vicinity of radioactive material. The purpose of the review is to ensure that individuals make every reasonable effort to maintain occupational doses, doses to the general public, and releases of radioactive material as low as reasonably achievable, taking into account the state of technology and the cost of improvements in relation to benefits.

(5) The licensee shall retain a current written description of the ALARA program for the duration of the license. The written description shall include:

(a) A commitment by management to keep occupational doses as low as reasonably achievable;

(b) A requirement that the radiation safety officer annually report to management in writing on the radiation safety program; and

(c) Categories of personnel exposure levels that, when exceeded, will initiate investigation by the radiation safety officer of the cause of the exposure and actions taken to reduce the probability of recurrence.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.710, Amended 2-11-10.

- 64E-5.605 Radiation Safety Officer.

(1) A licensee shall appoint a RSO who agrees in writing to be responsible for implementing the radiation safety program. The licensee, through the RSO, shall ensure that radiation safety activities are performed with approved procedures and regulatory requirements in the daily operation of the licensee’s radioactive materials program.

(2) The radiation safety officer shall promptly investigate and implement corrective actions as necessary regarding:

(a) Overexposures;

(b) Accidents;

(c) Spills;

(d) Losses;

(e) Thefts;

(f) Unauthorized receipts, uses, transfers, and disposals; and

(g) Other deviations from approved radiation safety practice. A written report of these investigations and the corrective actions taken shall be given to management.

(3) The radiation safety officer shall implement written policies and procedures to:

(a) Authorize the purchase of radioactive material;

(b) Receive and open packages of radioactive material;

(c) Store radioactive material;

(d) Keep an inventory record of radioactive material;

(e) Use radioactive material safely;

(f) Take emergency action if control of radioactive material is lost;

(g) Perform periodic radiation surveys;

(h) Perform checks of survey instruments and other safety equipment;

(i) Dispose of radioactive material;

(j) Train personnel who work in or frequent areas where radioactive material is used or stored; and

(k) Keep a copy of all records and reports required by department regulations, a copy of these regulations, a copy of each licensing request and license including amendments, and the written policies and procedures required by the regulations.

(4) The radiation safety officer shall approve radiation safety program changes for medical use not at a medical institution with the consent of management prior to sending to the department for licensing action.

(5) The radiation safety officer shall assist the radiation safety committee for medical use at a medical institution.

(6) The RSO shall review, sign and date, at least every 3 months, the occupational radiation exposure records of all personnel working with radioactive material.

(7) The licensee shall retain a copy of both authority, duties, and responsibilities of the RSO and a signed copy of each RSO’s agreement to be responsible for implementing the radiation safety program for the duration of the license. The records must include the signature of the RSO and licensee management.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.711, Amended 2-11-10.

- 64E-5.606 Radiation Safety Committee.

(1) Each license listed below shall establish a radiation safety committee to oversee the use of radioactive materials;

(a) Medical institutions as defined in Rule 64E-5.101, F.A.C.; or

(b) Other licenses authorized for any of the following medical uses:

1. Subsection 64E-5.627(2), F.A.C., and any subsection of Rule 64E-5.632 or 64E-5.634, F.A.C.;

2. Subsection 64E-5.627(3), F.A.C., and any subsection of Rule 64E-5.632 or 64E-5.634, F.A.C.;

3. Subsection 64E-5.627(4), F.A.C., and any subsection of Rule 64E-5.632 or 64E-5.634, F.A.C.;

4. Any subsection of Rule 64E-5.630, F.A.C., and any subsection of Rule 64E-5.632 or 64E-5.634, F.A.C.;

5. Subsections 64E-5.634(1) and 64E-5.634(2), F.A.C.;

6. Subsections 64E-5.634(1) and 64E-5.634(3), F.A.C.; or

7. Subsections 64E-5.634(2) and 64E-5.634(3), F.A.C.

(2) Membership of the radiation safety committee shall include an authorized user of each type of use permitted by the license, the RSO, a representative of the nursing service, and a representative of management who is neither an authorized user nor a RSO.  Other members who are experienced in the assay of radioactive material and protection against radiation, such as an authorized medical physicist or a nuclear medicine technologist employed by or working under contract with the institution may be included as appropriate.

(3) The committee shall meet at least every 6 months. To establish a quorum and to conduct business, one-half of the committee’s membership shall be present, including the RSO, or designee and the management representative, or designee.

(4) The minutes of each radiation safety committee meeting shall include:

(a) The date of the meeting;

(b) Members present;

(c) Members absent;

(d) Summary of deliberations and discussions;

(e) Recommended actions and the numerical results of all ballots; and

(f) Documentation of any reviews required in Rules 64E-5.604 and 64E-5.606, F.A.C.

(5) The committee shall provide each member with a copy of the meeting minutes and shall retain a copy for 5 years or until the department authorizes its disposition.

(6) The committee shall be responsible for monitoring the institutional program to maintain occupational doses as low as reasonably achievable.

(7) The committee shall review and approve any individual to be an authorized user, an authorized nuclear pharmacist, the RSO, or an authorized medical physicist based on safety and the training and experience standards of this part before sending a license application or request for amendment or renewal.

(8) The committee shall review and approve each proposed method of use of radioactive material based on safety.

(9) The committee shall review and approve procedures and radiation safety program changes based on safety and with the advice of the RSO and the management representative prior to sending to the department for licensing action.

(10) The committee shall review occupational radiation exposure records of all personnel working with radioactive material and all incidents involving radioactive material at least every 6 months, with the assistance of the RSO, to determine cause and review subsequent actions taken.

(11) The committee shall review the radioactive materials program at least every 12 months with the assistance of the RSO as described in subsection 64E-5.604(4), F.A.C.

(12) The committee shall establish levels for occupational dose that will result in investigations and considerations of action by the RSO when exceeded.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.712, Amended 8-6-01, 2-11-10.

- 64E-5.607 Authority and Responsibilities.

(1) A licensee shall provide sufficient authority and organizational freedom to the radiation safety officer and the radiation safety committee to:

(a) Identify radiation safety problems;

(b) Initiate, recommend, or provide solutions;

(c) Require and verify implementation of corrective actions; and

(d) Stop unsafe operations.

(2) A licensee shall establish in writing and keep current the authority, duties, responsibilities, and radiation safety activities of the radiation safety officer and the radiation safety committee.

(3) Authorized users shall have the following special responsibilities:

(a) For written directives:

1. A written directive must be dated and signed by an authorized user before the administration of I-131 as sodium iodide greater than 1.11 megabecquerels (MBq) (30 microcuries ([micro]Ci)), any therapeutic dosage of unsealed radioactive material or any therapeutic dose of radiation from material; or

2. Due to the emergent nature of the patient’s condition, a delay in order to provide a written directive would jeopardize the patient’s health, an oral directive is acceptable provided:

a. The information contained in the oral directive must be documented as soon as possible in writing in the patient’s record; and

b. A written directive must be prepared within 48 hours of the oral directive.

3. The written directive must contain the patient or human research subject’s name and the following information:

a. For any administration of quantities greater than 1.11 MBq (30 [micro]Ci) of sodium iodide I-131: the dosage;

b. For an administration of a therapeutic dosage of unsealed radioactive material other than sodium iodide I-131: the radioactive drug, dosage, and route of administration;

c. For gamma stereotactic radiosurgery: the total dose, treatment site, and values for the target coordinate settings per treatment for each anatomically distinct treatment site;

d. For teletherapy: the total dose, dose per fraction, number of fractions, and treatment site;

e. For high dose-rate remote afterloading brachytherapy: the radionuclide, treatment site, dose per fraction, number of fractions, and total dose; and

f. For all other brachytherapy;

(I) Before implantation: treatment site, the radionuclide, and dose; and

(II) After implantation but before completion of the procedure: the radionuclide, treatment site, number of sources, and total source strength and exposure time (or the total dose).

4. A written revision to an existing written directive may be made if the revision is dated and signed by an authorized user before the administration of the dosage of unsealed radioactive material, the brachytherapy dose, high dose remote afterloader dose, the gamma stereotactic radiosurgery dose, the teletherapy dose, or the next fractional dose; or

5. Due to the emergent nature of the patient’s condition, a delay in order to provide a written directive would jeopardize the patient’s health, an oral directive is acceptable provided:

a. The information contained in the oral directive must be documented as soon as possible in the patient’s record; and

b. A written directive must be prepared within 48 hours of the oral directive.

(b) Review personally the patient’s case to assure that the therapeutic radiation procedure is appropriate;

(c) Review personally the patient’s case or develop and implement adequate written procedures to assure that the diagnostic radiation procedure is appropriate.

(d) Prior to administration, the authorized user must document deviations from the diagnostic clinical procedures manual for each patient.

(e) Use radioactive material or direct technologists and physicians in training in using radioactive material;

(f) Interpret results of diagnostic procedures; and

(g) Review regularly the progress of the patient receiving therapy and modify the originally prescribed dose if needed.

(4) The licensee shall retain a copy of the written directives specified in paragraph 64E-5.607(3)(a), F.A.C., for three years.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.713, Amended 2-11-10, 12-26-13.

- 64E-5.608 Supervision.

(1) Supervision of a physician in training to become an authorized user:

(a) A licensee who permits the receipt, acquisition, possession, use, preparation, or transfer of radioactive material by a physician in training under the supervision of an authorized user as allowed by paragraph 64E-5.601(3)(a), F.A.C., shall:

1. Instruct the supervised individual in the principles of radiation safety appropriate to that individual’s use of radioactive material and in the licensee’s written quality management program;

2. Review the supervised individual’s use of radioactive material, provide reinstruction as needed and review records kept to reflect this use;

3. Require the preparation of radioactive materials use only under the supervision of an authorized user or authorized nuclear pharmacist;

4. Require the authorized user to be immediately available to communicate with the supervised individual; and

5. Require that only those individuals specifically designated by the authorized user be permitted to administer radionuclides or radiation to patients.

(b) A licensee shall require the supervised individual receiving, possessing, acquiring, preparing, using or transferring radioactive material specified in paragraph 64E-5.601(3)(a), F.A.C., to:

1. Follow the instructions of the supervising authorized user;

2. Follow the written radiation and quality management program procedures established by the licensee; and

3. Comply with these regulations and the license conditions regarding the use of radioactive material.

(c) The licensee’s management or radiation safety committee shall provide written approval prior to any training of a physician to receive, acquire, prepare, possess or use radioactive material under the supervision of an authorized user. After the training has been completed, the licensee shall provide documentation to the supervised individual that the individual received the training and experience required by this section. The licensee shall maintain records that identify physicians currently in training and the physicians who have completed training for 7 years after the last date training was received.

(2) Supervision of an individual in training to become an authorized medical physicist:

(a) A licensee who permits the receipt, preparation, acquisition, possession, use, or transfer of radioactive material to an individual in training under the supervision of an authorized medical physicist as allowed by paragraph 64E-5.601(3)(c), F.A.C., shall:

1. Instruct the supervised individual in the principles of radiation safety appropriate to that individual’s use of radioactive material and in the licensee’s written quality management program;

2. Review the supervised individual’s use of radioactive material, provide reinstruction as needed and review records kept to reflect this use; and

3. Require the authorized medical physicist to be immediately available to communicate with the supervised individual.

(b) A licensee shall require the supervised individual receiving, acquiring or preparing, possessing, using or transferring radioactive material specified in paragraph 64E-5.601(3)(c), F.A.C., to:

1. Follow the instructions of the supervising authorized medical physicist;

2. Follow the written radiation and quality management program procedures established by the licensee; and

3. Comply with these regulations and the license conditions regarding the use of radioactive material.

(c) The licensee’s management or radiation safety committee shall provide written approval prior to any individual to receive, possess or use radioactive material under the supervision of an authorized medical physicist. After the training has been completed, the licensee shall provide documentation to the supervised individual that the individual received the training and experience required by this section. The licensee shall maintain records that identify individuals currently in training and the individuals who have completed training for 7 years after the last date training was received.

(3) A licensee that permits any supervised activities regarding the use of radioactive materials or radiation from radioactive materials is responsible for the acts and omissions of the supervised individual.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.714, Amended 2-11-10.

- 64E-5.609 Visiting Authorized User, Visiting Authorized Medical Physicist, or Visiting RSO.

(1) A licensee may permit any visiting authorized user to use licensed material for medical use under the terms of the licensee’s license for 60 days each year if:

(a) The licensee has a copy of a license issued by the department, the NRC, or an agreement state that identifies the visiting authorized user by name as an authorized user for medical use; and

(b) The visiting authorized user performs only those procedures for which he is specifically authorized by the license described in paragraph 64E-5.609(1)(a), F.A.C., above.

(2) For up to 60 days each year, a licensee may permit an authorized medical physicist or an individual qualified under Rules 64E-5.656 and 64E-5.658, F.A.C., to function as a visiting authorized medical physicist as authorized by the license.

(3) For up to 60 days each year, a licensee may permit an authorized user or an individual qualified to be a RSO, under Rules 64E-5.648 and 64E-5.658, F.A.C., to function as a visiting RSO and to perform the functions of a RSO, as provided in Rule 64E-5.605 and subsection 64E-5.607(1), F.A.C.

(4) A license amendment is not needed to permit a visiting authorized user, visiting authorized medical physicist, or visiting RSO to use licensed material or perform functions in accordance with this section.

(5) The visiting authorized user, visiting authorized medical physicist, or visiting RSO shall have the prior written permission of the licensee’s management and, if the use or function occurs on behalf of a medical institution, the institution’s radiation safety committee.

(6) Licensee records shall include a copy of the record described in Rule 64E-5.657, F.A.C., or some other form of documentation that verifies the individual has met the respective training and experience requirements listed in Subpart I. A licensee shall retain copies of the records specified in Rule 64E-5.609, F.A.C., for 3 years after the last visit.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.715, Amended 2-11-10, 12-26-13.

- 64E-5.610 Mobile Medical Service Requirements.

The department shall license mobile medical services or clients of such services. The mobile medical service shall be licensed if the service receives, uses or possesses radioactive material. The client of the mobile medical service shall be licensed if the client receives or possesses radioactive material to be used by a mobile medical service.

(1) The mobile medical licensee shall obtain a letter signed by the management of each client for which services are rendered that permits the use of radioactive materials at the client’s address and clearly delineates the authority and responsibility of the licensee and the client.  A licensee providing mobile medical services shall retain this letter for 3 years after the provision of service.

(2) Mobile medical service licensees shall secure or keep under constant surveillance and immediate control all radioactive material when in transit or at a location of use.

(3) The mobile medical licensee shall check instruments used to measure the activity of unsealed or sealed radioactive materials for proper function before medical use at each client’s address or on each day of use, whichever is more frequent. At a minimum, the check for proper function required by this paragraph must include a constancy check and shall perform all daily quality control tests on all equipment used to obtain images or information from radionuclide studies before medical use at each location of use.

(4) Before leaving a client location, mobile medical service licensees shall perform a survey of all areas where radioactive materials are used with a radiation survey instrument in order to ensure that they have complied with the requirements in Rule 64E-5.621, F.A.C., that radiation dose rates are at background levels, and that removable contamination is below 2000 disintegrations per minute per 100 square centimeters sampled. A licensee shall check each survey instrument for proper operation with a dedicated check source before each use at each location. The licensee is not required to keep records of these dedicated source survey instrument checks.

(5) Mobile medical service licensees shall retain a record of each survey required for 3 years. The record must include the date of the survey, a diagram of each area that was surveyed, the measured dose rate at several points in each area of use in millirems (microsieverts) per hour, the model and serial number of the instrument used to make the survey, and the initials of the individual who performed the survey.

(6) A physician shall be on site at each client’s address at the time radioactive materials are administered. An authorized user shall be immediately available to communicate with the supervised individuals or individuals under their direction.

(7) Radioactive material will be received at the permanent location of the mobile medical service or delivered directly to an authorized individual in the vehicle at a place of use. A mobile medical service may not have radioactive materials delivered from the manufacturer or the distributor to the client unless the client has a radioactive materials license allowing possession of the radioactive material. Radioactive material delivered to the client must be received and handled in conformance with the client’s license.

(8) Restrooms contained in mobile vehicles shall not routinely be used by patients who have been administered radioactive material.

(9) Radioactive gases or aerosols shall not be used by mobile medical service licenses.

(10) Prior to administration, the mobile medical service licensee shall assure that individuals or human research subjects meet the patient release criteria specified in Rule 64E-5.622, F.A.C.

(11) A licensee authorized to use mobile remote afterloaders for medical use shall follow the requirements specified in Rule 64E-5.6423, F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.716, Amended 2-11-10.

- 64E-5.611 Quality Management Program and Notifications, Records and Reports of Medical Events.

(1) Each applicant or licensee under this part, as applicable, shall establish and maintain a written quality management program to provide a high confidence that radioactive material or radiation from radioactive material will be administered as directed by the authorized user. The quality management program must include written policies and procedures to meet the following objectives:

(a) Except where a delay to provide a written directive would jeopardize the patient’s health as specified in paragraphs (b) and (c) of this section, a written directive is prepared prior to administration for the following:

1. Any teletherapy radiation dose;

2. Any gamma stereotactic radiosurgery radiation dose;

3. Any brachytherapy radiation dose;

4. Any administration of iodine 131 as sodium iodide in quantities greater than 30 microcuries (1.11 megabecquerels);

5. Any therapeutic administration of a radiopharmaceutical other than iodine 131 as sodium iodide; or

6. Any high dose rate remote afterloader radiation dose.

(b) An oral directive is acceptable when a delay to provide a written directive would jeopardize the patient’s health because of the emergent nature of the patient’s condition. The information contained in the oral directive must be documented immediately in the patient’s record and a written directive prepared within 48 hours of the oral directive.

(c) An oral revision to an existing written directive is acceptable when a delay to provide a written revision to an existing written directive would jeopardize the patient’s health. The oral revision must be documented immediately in the patient’s record and a revised written directive must be signed by the authorized user within 48 hours of the oral revision.

(d) A written directive which changes an existing written directive can be made for any diagnostic or therapeutic procedure if the revision is dated and signed by an authorized user prior to the administration of the radiopharmaceutical dosage, the brachytherapy dose, the gamma stereotactic radiosurgery dose, high dose rate remote afterloader dose, the teletherapy dose, or the next fractional dose.

(e) The patient’s or human research subject’s identity is verified by more than one method as the individual named in the written directive prior to administration;

(f) The final plans of treatment and related dose calculations, manually or computer generated, for brachytherapy, teletherapy, high dose rate remote afterloader, and gamma stereotactic radiosurgery agree with the respective written directives:

(g) Verify that any computer-generated calculations are correctly transferred into the consoles of therapeutic medical units authorized by Rule 64E-5.634, F.A.C.;

(h) Each administration agrees with the written directive; and

(i) Any unintended deviation from the written directive is identified and evaluated and appropriate action is taken.

(2) The licensee shall develop procedures for and conduct a review of the quality management program including an evaluation of the following:

(a) A representative sample of patient administrations within the review period;

(b) All recordable events within the review period; and

(c) All medical events within the review period to verify compliance with all aspects of the quality management program.

(3) The review of the quality management program specified in subsection (2) above shall be conducted at intervals not to exceed 12 months. A record of each review shall be maintained for inspection by the department in an auditable form for 3 years and shall include evaluations and findings of the review.

(4) The licensee shall evaluate each of these reviews to determine the effectiveness of the quality management program and make modifications to meet the objectives in subsection 64E-5.611(1), F.A.C.

(5) Within 30 days of discovery of each recordable event, the licensee shall:

(a) Assemble the relevant facts including the cause;

(b) Identify any corrective action required to prevent recurrence;

(c) Retain a record in an auditable form for 3 years of the relevant facts and any corrective action taken.

(6) The licensee shall retain in an auditable form for 3 years each written directive and a record of each administered radiation dose or radiopharmaceutical dosage where a written directive is required by subsection 64E-5.611(1), F.A.C.

(7) Each applicant for a new license shall submit to the department a quality management program as part of the application for a license and implement the program upon issuance of the license by the department.

(8) Each licensee shall maintain copies of the quality management program for the duration of the license.

(9) Each licensee shall submit and maintain records and reports of medical events as required by subsections 64E-5.345(4) and (5), F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.717, Amended 2-11-10.

- 64E-5.612 Suppliers.

A licensee shall use for medical use only:

(1) Radioactive material manufactured, labeled, packaged, and distributed as specified in a license issued by the department or by another agreement state, a licensing state or the U.S. Nuclear Regulatory Commission;

(2) Generators and reagent kits that have been manufactured, labeled, packaged, and distributed as specified in an approval issued by the U.S. Department of Health and Human Services, Food and Drug Administration unless the kits are not subject to the Federal Food, Drug, and Cosmetics Act and the Public Health Services Act.

(3) Teletherapy sources manufactured and distributed as specified in a license issued by the department or by another agreement state, a licensing state or the NRC; or

(4) Sealed sources or devices containing radioactive materials that are either;

(a) Manufactured, labeled, packaged, and distributed as specified in a license issued by the department or by another agreement state, a licensing state or the NRC; or

(b) Noncommercially transferred from a medical use licensee authorized by Chapter 64E-5, Part VI, F.A.C., or equivalent medical use license issued by another agreement state or the NRC.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.718, Amended 2-11-10.

- 64E-5.613 Quality Control of Diagnostic Instrumentation.

Each licensee shall establish written quality control procedures for all equipment used to obtain images or information from radionuclide studies. The procedures shall be recommended by equipment manufacturers or be approved by the department. The licensee shall perform quality control as specified in written procedures and retain a copy of the quality control results for 3 years.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-13-93, Formerly 10D-91.719.

SUBPART B GENERAL TECHNICAL REQUIREMENTS

- 64E-5.614 Possession, Use, Calibration, and Check of Dose Calibrators in the Use of Unsealed Radiopharmaceuticals.

(1) A medical use licensee authorized to administer radiopharmaceuticals shall possess a dose calibrator and use it to measure the amount of activity administered to each patient or human research subject.

(2) A licensee shall check each dose calibrator before use each day of use, or during an assigned shift for facilities operating continuously, for constancy with a dedicated check source. The check shall be performed on a frequently used setting with a sealed source of not less than 10 microcuries (370 kBq) of radium 226 or 50 microcuries (1.85 MBq) of any other photon-emitting radionuclide with a half-life greater than 90 days. A record shall be made of each check, which shall include:

(a) The model and serial number of the dose calibrator;

(b) The identity and decay corrected activity of the radionuclide contained in the check source;

(c) The date of the check;

(d) The activity measured;

(e) The percent error;

(f) The instrument settings; and

(g) The name or initials of the individual performing the check.

(3) The licensee shall test each dose calibrator for accuracy at the time of installation and at least every 12 months. The test shall be completed by assaying at least two sealed sources containing different radionuclides, the activity of which has been determined by the National Institute of Standards and Technology (NIST) or by the manufacturer who has compared their source to a source calibrated by the NIST. The sources shall have a minimum activity of 10 microcuries (370 kBq) for radium 226 and 50 microcuries (1.85 MBq) for any other photon-emitting radionuclide. At least one of the sources shall have a principal photon energy between 100 kiloelectron volts and 500 kiloelectron volts. A record shall be made of each test, which shall include:

(a) The model and serial number of the dose calibrator;

(b) The model and serial number of each source used and the identity of the radionuclide contained in the source and its activity;

(c) The date of the test;

(d) The results of the test;

(e) The instrument settings; and

(f) The name of the individual performing this test.

(4) The licensee shall test each dose calibrator for linearity at the time of installation and at least every 3 months over the range of use between 10 microcuries (370 kBq) and the highest dosage that will be administered. A record shall be made of each test, which shall include:

(a) The model and serial number of the dose calibrator;

(b) The calculated activities;

(c) The measured activities;

(d) The date of the test; and

(e) The name of the individual performing this test.

(5) The licensee shall test each dose calibrator for geometry dependence at the time of installation over the range of volumes and volume configurations for which it will be used. The licensee shall keep a record of this test for the duration of the use of the dose calibrator. A record shall be made of each test, which shall include:

(a) The model and serial number of the dose calibrator;

(b) The configuration of the source measured;

(c) The activity measured and the instrument setting for each volume measured;

(d) The date of the test; and

(e) The name of the individual performing this test.

(6) A licensee shall correct mathematically dosage readings for any geometry or linearity error that exceeds 10 percent if the dosage is greater than 10 microcuries (370 kBq) and shall repair or replace the dose calibrator if the accuracy or constancy error exceeds 10 percent.

(7) A licensee shall also perform checks and tests required by Rule 64E-5.614, F.A.C., following adjustment or repair of the dose calibrator.

(8) A licensee shall retain a record of each check and test required by Rule 64E-5.614, F.A.C., for 3 years.

(9) A licensee may calibrate instrumentation used in Rule 64E-5.614, F.A.C., using nationally recognized standards or the manufacturer’s instructions. The standards or instructions used by the licensee must be available for inspection by the department.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.720, Amended 2-11-10, 12-26-13.

- 64E-5.615 Use, Calibration and Check of Survey Instruments.

A licensee shall ensure that the survey instruments used to comply with this part have been calibrated before first use, at least every 12 months, and after repair.

(1) A record shall be made of each calibration, which shall include:

(a) A description of the source used;

(b) The certified dose rates from the source;

(c) The rates indicated by the instrument being calibrated;

(d) The correction factors deduced from the calibration data;

(e) The name of the individual who performed the calibration;

(f) The date of calibration;

(g) The model number and serial number of the instrument being calibrated; and

(h) The results of the calibration.

(2) The licensee shall:

(a) Calibrate all required scale readings up to 1,000 millirems (10 mSv) per hour with a radiation source;

(b) Calibrate each linear scale instrument at two points located approximately 1/3 and 2/3 of full-scale, calibrate each logarithmic scale instrument at midrange of each decade and at two points of at least one decade, and calibrate each digital instrument at appropriate points; and

(c) Conspicuously note on the instrument the date of calibration.

(3) The licensee shall:

(a) Consider a point as calibrated if the indicated dose rate differs from the calculated dose rate by not more than 10 percent; and

(b) Consider a point as calibrated if the indicated dose rate differs from the calculated dose rate by not more than 20 percent and a correction chart or graph is attached conspicuously to the instrument.

(4) A licensee shall check each survey instrument for proper operation with a dedicated check source before each use. The licensee is not required to keep records of these checks.

(5) The licensee shall retain a record of each calibration required in subsection 64E-5.615(1), F.A.C., for 3 years.

(6) The licensee may use persons licensed by the department, the U.S. Nuclear Regulatory Commission, an agreement state, or a licensing state to perform calibrations of survey instruments. Records of calibrations required by subsection 64E-5.615(1), F.A.C., shall be maintained by the licensee.

(7) A licensee authorized to use radioactive material for uptake, dilution, and excretion studies or sealed sources for diagnostic purposes shall possess a portable radiation survey instrument with a range from 0.1 millirem (1.0 µSv) per hour to at least 1,000 millirem (10 mSv) per hour.

(8) A licensee authorized to use radioactive material for imaging and localization studies, radiopharmaceutical therapy or implant therapy shall possess portable radiation survey instruments with a range from 0.1 millirem (1.0 µSv) per hour to at least 1,000 millirem (10 mSv) per hour.

(9) A licensee authorized to use radioactive material in Rule 64E-5.634, F.A.C., shall possess a radiation survey instrument as described in subsection (7) or (8), above.

(10) The licensee shall not use survey instruments if the difference between the indicated exposure rate and the calculated exposure rate is more than 20 percent.

(11) A licensee may calibrate instrumentation used in Rule 64E-5.615, F.A.C., using nationally recognized standards or the manufacturer’s instructions. The standards or instructions used by the licensee must be available for inspection by the department.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), (3), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-15-96, Formerly 10D-91.721, Amended 2-11-10.

- 64E-5.616 Determination of Dosages of Unsealed Radioactive Material for Medical Use.

(1) The licensee shall determine by assay or direct measurement within 30 minutes before each radiopharmaceutical dosage and record the activity of each dosage before medical use. A record of the assay shall be made which shall include:

(a) The generic name, trade name, or abbreviation of the radiopharmaceutical; its lot number; expiration date; and the radionuclide;

(b) The patient’s or human research subject’s name or identification number if one has been assigned;

(c) The prescribed dosage and activity of the dosage at the time of assay or a notation that the total activity assayed is less than 10 microcuries (370 kBq);

(d) The date and time of the assay and administration; and

(e) The name of the individual who performed the assay.

(2) Unless directed by the authorized user, a licensee may not use a dosage if the dosage does not fall within the prescribed dosage range or if the dosage differs from the prescribed dosage by more than 20 percent.

(3) A licensee shall retain a record of the assays listed in Rule 64E-5.616, F.A.C., for 3 years.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.722, Amended 2-11-10.

- 64E-5.617 Authorization for Calibration, Transmission and Reference Sources.

Any person authorized by Rule 64E-5.601, F.A.C., for medical use of radioactive material may receive, possess, and use the following radioactive material for check, calibration, transmission and reference use:

(1) Sealed sources that:

(a) Do not exceed 1.11 GBq (30 mCi) each, manufactured and distributed by a person licensed by the department, the NRC, an agreement state; or

(b) Do not exceed 1.11 GBq (30 mCi) each, which are redistributed by a licensee that is authorized to redistribute sealed sources that are manufactured and distributed by a person licensed by the department, the NRC, or an agreement state, provided the redistributed sealed sources are in the original packaging and shielding, and are accompanied by the manufacturer’s approved instructions;

(2) Any radioactive material with a half-life of 120 days or less in individual amounts not to exceed 15 millicuries (555 MBq) each;

(3) Any radioactive material with a half-life greater than 120 days in individual amounts not to exceed 200 microcuries (7.4 MBq) each; and

(4) Unless approved by the department, the maximum possession limit of radioactive materials described in subsections 64E-5.617(1), (2) and (3), F.A.C., above, shall not exceed a combined activity of 1 curie (37 GBq), This includes radioactive materials as waste in storage.

(5) Unless approved by the department, the maximum possession limit for Technetium 99m in individual amounts shall not exceed 300 millicuries (11.1 GBq) each and a combined activity of 900 millicuries (33.3 GBq).

Rulemaking Authority 404.051, 404.061, 404.141 FS. Law Implemented 404.051(1), (4), (6), (10), 404.061(2), 404.141 FS. History–New 8-25-91, Formerly 10D-91.723, Amended 5-18-98, 2-11-10.

- 64E-5.618 Requirements for Possession of Sealed Sources and Brachytherapy Sources.

(1) A licensee who possesses any sealed source or brachytherapy source shall follow the radiation safety and handling instructions supplied by the manufacturer or equivalent instructions approved by the department and shall maintain the instructions for the duration of source use in a legible form and convenient to users.

(2) A licensee in possession of a sealed source shall assure that:

(a) The source is tested for leakage before its first use unless the licensee has a certificate from the supplier indicating that the source was tested within 6 months before transfer to the licensee; and

(b) The source is tested for leakage at least every 6 months or at intervals approved by the department, another agreement state, a licensing state or the U.S. Nuclear Regulatory Commission.

(c) Leak tests are capable of detecting the presence of 0.005 microcurie (185 Bq) of radioactive material on the test sample, or, in the case of radium, the escape of radon at the rate of 0.001 microcurie (37 Bq) each 24 hours;

(d) Test samples are taken from the source or from the surfaces of the device in which the source is mounted or stored on which radioactive contamination might be expected to accumulate; and

(e) Teletherapy and other device source samples are taken when the source is in the off position.

(f) Leak tests are analyzed by individuals who are licensed by the department, U.S. Nuclear Regulatory Commission, an agreement state or licensing state to perform leak test services.

(3) A licensee shall retain leak test records for 3 years. The records shall contain the model number and serial number if assigned of each source tested, the identity of each source radionuclide and its estimated activity, the measured activity of each test sample expressed in microcuries (becquerels), the date of the test, and the name of the individual who performed the test analysis.

(4) If the leak test reveals the presence of 0.005 microcurie (185 Bq) or more of removable contamination, the licensee shall:

(a) Immediately withdraw the sealed source from use and cause it to be decontaminated and repaired or to be disposed of in accordance with these regulations; and

(b) File a written report with the department within 5 days of receiving the leak test results describing the equipment involved, the test results, the action taken, the model number and serial number or the leaking source if assigned, the radioisotope and its estimated activity, and the date of the test.

(5) A leak test is not required on the following sources:

(a) Sources containing only radioactive material with a half-life of less than 30 days;

(b) Sources containing only radioactive material as a gas;

(c) Sources containing 100 microcuries (3.7 MBq) or less of beta or photon-emitting material or 10 microcuries (370 kBq) or less of alpha-emitting material; and

(d) Seeds of iridium 192 encased in nylon ribbon.

(6) Leak tests are not required on calibration and reference sources stored and not being used. The licensee shall, however, clearly indicate on the inventory records that these sources are for storage only and the date placed in storage. The licensee shall test each such source for leakage before any use or transfer unless it has been tested for leakage within 6 months before the date of use or transfer.

(7) Leak tests are not required on brachytherapy and teletherapy sources that are listed on a department license for storage only. The licensee shall test each such source for leakage before any use or transfer unless it has been tested for leakage within 6 months before the date of use or transfer.

(8) A licensee who possesses sealed sources or brachytherapy sources, except gamma stereotactic radiosurgery sources, shall conduct a physical inventory of all such sources at intervals not to exceed six months. The licensee shall retain each inventory record for 3 years.  The inventory records shall contain the model number of each source and serial number if one has been assigned, the identity of each source radionuclide and its estimated activity, the location of each source, the date of the inventory, and the name of the individual who performed the inventory.

(9) A licensee who possesses a sealed source or brachytherapy source shall survey all areas where such sources are stored with a radiation survey instrument at least every 3 months. This does not apply to teletherapy sources in teletherapy units or sealed sources in diagnostic devices.

(10) A licensee shall retain a record of each survey required in subsection 64E-5.618(9), F.A.C., for 3 years. The record shall include the date of the survey, a sketch of each area that was surveyed, the measured dose rate at several points in each area expressed in millirems (microsieverts) per hour, the model number and serial number of the survey instrument used to make the survey, and the name of the individual who performed the survey.

(11) Sealed sources designated as radioactive waste and held for decay in storage as in Rule 64E-5.624, F.A.C., are not required to be leak tested or inventoried as required by this section.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.724, Amended 2-11-10.

- 64E-5.619 Syringe Shields and Labels.

(1) A licensee shall keep syringes that contain radioactive material to be administered in a radiation shield. Each individual who prepares or administers radiopharmaceuticals shall use a syringe radiation shield unless the use of the shield is contraindicated for that patient.

(2) Unless used immediately, a licensee shall label conspicuously each syringe or syringe radiation shield that contains a syringe with a radiopharmaceutical with the patient’s name or the radiopharmaceutical name or its abbreviation and the type of diagnostic study or therapy procedure to be performed.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), (3), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-15-96, Formerly 10D-91.725.

- 64E-5.620 Vial Shields and Labels.

A licensee shall require each individual preparing or handling a vial that contains a radiopharmaceutical to keep the vial in a vial radiation shield and conspicuously label each vial with the radiopharmaceutical name or its abbreviation.

Rulemaking Authority 404.022, 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), (3), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-15-96, Formerly 10D-91.727.

- 64E-5.621 Surveys for Contamination and Ambient Radiation Dose Rate.

(1) A licensee shall survey with a radiation survey instrument at the end of each day of use, or during an assigned shift for facilities operating continuously, all areas where radiopharmaceuticals are routinely prepared for use or administered.

(2) A licensee shall survey all areas where radiopharmaceuticals or radioactive wastes are stored with a radiation survey instrument at least once each week.

(3) A licensee shall conduct the surveys required by subsections 64E-5.621(1) and (2), F.A.C., with an instrument capable of measuring dose rates as low as 0.1 millirem (1 µSv) per hour.

(4) A licensee shall establish dose rate action levels for the surveys required by subsections 64E-5.621(1) and (2), F.A.C., and shall require that the individual performing the survey immediately notify the radiation safety officer if a dose rate exceeds an action level.

(5) A licensee shall perform a wipe survey for removable contamination weekly of all areas where radiopharmaceuticals are routinely prepared for use, administered or stored.

(6) A licensee shall analyze the wipe surveys required by subsection 64E-5.621(5), F.A.C., with an instrument capable of detecting contamination of 2,000 disintegrations per minute (33.3 Bq) or shall monitor each wipe sample in a low background area with a radiation survey instrument using a probe with a maximum window thickness of 2.0 mg/cm2 and a minimum probe diameter of 1.5 inches.

(7) A licensee shall establish removable contamination action levels for the wipe surveys required by subsection 64E-5.621(5), F.A.C., and shall require that the individual performing the survey immediately notify the radiation safety officer if contamination exceeds action levels.

(8) A licensee shall retain a record of each survey required by subsections 64E-5.621(1), (2), and (5), F.A.C., for 3 years. The record shall include:

(a) The date of the survey;

(b) A diagram of each area surveyed;

(c) Action levels established for each area;

(d) The measured dose rate at several points in each area expressed in millirems (microsieverts) per hour or the removable contamination in each area expressed in disintegrations per minute (becquerels) per 100 square centimeters, or counts per minute if performed with a radiation survey instrument as described in subsection 64E-5.621(6), F.A.C.;

(e) The serial number and the model number of the instrument used to make the survey or analyze the samples; and

(f) The name of the person who performed the survey.

(9) The licensee does not need to perform the radiation surveys in subsection 64E-5.621(1) or (2), F.A.C., in areas where patients or human research subjects are currently confined when such patients or subjects cannot be released under Rule 64E-5.622, F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.729, Amended 2-11-10.

- 64E-5.622 Release of Patients or Human Research Subjects Treated with Radiopharmaceuticals, Implants or Remote Afterloader Units.

(1) Except as authorized by subsection 64E-5.622(4), F.A.C., a licensee shall not authorize release from confinement for medical care any patient administered a radiopharmaceutical until:

(a) The dose rate from the patient is less than 5 millirems (50 µSv) per hour at a distance of 1 meter; or

(b) The activity in the patient is less than 30 millicuries (1.11 GBq).

(2) Except as authorized by subsection 64E-5.622(4), F.A.C., a licensee shall not authorize release from confinement for medical care any patient administered a permanent implant until the dose rate from the patient is less than 5 millirems (50 µSv) per hour at a distance of 1 meter.

(3) Immediately after removing the last temporary implant source from a patient, the licensee shall make a radiation survey of the patient with a radiation survey instrument to confirm that all sources have been removed. The licensee shall not release a patient treated by temporary implant from confinement for medical care until all sources have been removed.

(4) Licensees and license applicants whose proposed procedures to release individuals who have been administered radiopharmaceuticals or permanent implants containing radioactive material from the control of licensees differ from those specified in subsections (1) and (2), above, must submit their proposed procedures to the department for approval. The procedures must:

(a) Demonstrate that the total effective dose equivalent to any other individual from exposure to the released individual is not likely to exceed 500 millirem (5 µSv);

(b) Contain a copy of the instructions including written instructions to be given to the released individual, or the individual’s parent or guardian, on actions recommended to maintain doses to other individuals as low as is reasonably achievable if the total effective dose equivalent to another individual is likely to exceed 100 millirem (1 µSv). If the dose to a breast-feeding infant or child could exceed 100 millirem (1 µSv) if there were no interruption of breast-feeding, the instructions also shall include:

1. Guidance on the interruption or discontinuance of breast-feeding, and

2. Information on the consequences of failing to follow the guidance.

(c) Specify that the licensee shall maintain a record of the basis for authorizing the release of an individual from their control who has been administered radiopharmaceuticals or permanent implants containing radioactive material for 3 years after the date of release.

(5) A licensee shall maintain a record of patient surveys which demonstrates compliance with subsections 64E-5.622(3) and (6), F.A.C., for 3 years. Each record shall include the date of the survey, the name of the patient, the dose rate from the patient expressed as millirems (microsieverts) per hour and measured within 1 meter from the patient, and the initials of the individual who performed the survey.

(6) Before releasing a patient or human research subject from licensee control, a licensee shall survey the patient or the human research subject and the remote afterloader unit with a portable radiation detection survey instrument to confirm that the source(s) has been removed from the patient or human research subject and returned to the safe shielded position.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), (3), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-15-96, Formerly 10D-91.730, Amended 10-8-00, 2-11-10.

- 64E-5.623 Storage of Volatiles and Gases.

A licensee shall store volatile radiopharmaceuticals and radioactive gases in the shippers’ radiation shield and container or an equivalent shield and container. A licensee shall store and use a multidose container in a properly functioning fume hood.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.731.

- 64E-5.624 Decay in Storage.

(1) A licensee shall hold radioactive material with a physical half life of less than 120 days for decay in storage before disposal as ordinary trash. A licensee is exempt from the requirements of paragraph 64E-5.331(1)(c), F.A.C., of these regulations if:

(a) The radioactive material is held for decay a minimum of 10 half-lives;

(b) The radioactive material is monitored at the container surface before disposal as ordinary trash and its radioactivity cannot be distinguished from the background radiation level in a low background radiation area with a radiation survey instrument set on its most sensitive scale and with no interposed shielding;

(c) All radiation labels are removed or obliterated; and

(d) Each generator column is separated and monitored individually with all radiation shielding removed to ensure that its contents have decayed to background radiation levels before disposal.

(2) The licensee shall retain a record of each disposal for 3 years. The record shall include:

(a) The date of the disposal;

(b) The date on which the radioactive material was placed in storage;

(c) The radionuclides disposed;

(d) The model and serial number of the radiation survey instrument used;

(e) The background dose rate;

(f) The radiation dose rate measured at the surface of each waste container; and

(g) The name of the individual who performed the disposal.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.732, Amended 2-11-10.

- 64E-5.625 Safety Instructions and Precautions for Liquid Iodine, Radiopharmaceutical Therapy, Manual Brachytherapy, Remote Afterloader Units, Teletherapy Units and Gamma Stereotactic Radiosurgery.

(1) A licensee shall provide oral and written radiation safety instructions to all personnel caring for patients or human research subjects, who cannot be released under Rule 64E-5.622, F.A.C., undergoing radiopharmaceutical therapy or manual brachytherapy. This training shall be provided initially prior to caring for patients and refresher training shall be provided at least every 12 months. The instruction shall describe the licensee’s procedures for notification of the RSO and an authorized user in case of the patient’s death or medical emergency.

(2) The instruction for radiopharmaceutical therapy shall be commensurate with the duties of the personnel and describe the procedures for:

(a) Patient or human research subject control;

(b) Visitor control, including;

1. Routine visitation to hospitalized individuals in accordance with paragraph 64E-5.312(1)(a), F.A.C.; and

2. Visitation authorized in accordance with subsection 64E-5.312(5), F.A.C.

(c) Contamination control; and

(d) Waste control.

(3) The instruction for manual brachytherapy shall be commensurate with the duties of the personnel and describe:

(a) Size and appearance of the brachytherapy sources;

(b) Safe handling and shielding instructions;

(c) Procedures for patient or human research subject control; and

(d) Procedures for visitor control, including;

1. Routine visitation to hospitalized individuals in accordance with paragraph 64E-5.312(1)(a), F.A.C.; and

2. Visitation authorized in accordance with paragraph 64E-5.312(5), F.A.C.

(4) A licensee shall provide instruction for remote afterloader units, teletherapy units, and gamma stereotactic radiosurgery units as described in Rule 64E-5.636, F.A.C.

(5) A licensee shall keep a record of individuals receiving instruction required by subsections (1), (2), (3) and (4), above, which includes a list of topics covered, the date of instruction, and the name of the individual who gave the instruction. Such record shall be maintained for inspection by the department for 3 years.

(6) A licensee shall take the following safety precautions for each patient or human research subject receiving manual brachytherapy or radiopharmaceutical therapy who cannot be released under Rule 64E-5.622, F.A.C.

(a) Post the patient’s or human research subject’s door with a “Caution: Radioactive Material” sign and note on the door or on the patient’s chart where and how long visitors may stay in the patient’s room.

(b) Measure promptly, after administration of the dosage, the dose rates in contiguous restricted and unrestricted areas with a radiation measurement survey instrument to demonstrate compliance with the requirements of Rule 64E-5.312, F.A.C. Retain for 3 years a record of each survey that includes the time and date of the survey, a plan of the area or list of points surveyed, the measured dose rate at several points expressed in millirems per hour, the instrument used to make the survey, and the initials of the individual who made the survey.

(c) Provide the patient with radiation safety guidance before authorizing release of the patient that will help to keep radiation dose to household members and the public as low as reasonably achievable.

(d) Notify the RSO and an authorized user immediately if the patient dies or has a medical emergency.

(7) Individuals receiving radiopharmaceutical therapy shall be provided a private room with a private sanitary facility or a room with another individual who is receiving unsealed radioactive materials who cannot be released under Rule 64E-5.622, F.A.C. Individuals receiving manual brachytherapy shall be provided a private room or a room with another individual who is receiving manual brachytherapy and cannot be released under Rule 64E-5.622, F.A.C. The licensee shall not place an individual receiving manual brachytherapy in the same room with a patient who is not receiving manual brachytherapy.

(8) A licensee shall take these additional safety precautions for radiopharmaceutical therapy patients or human research subjects who cannot be released by Rule 64E-5.622, F.A.C.:

(a) Monitor material and items removed from the patient’s or human research subject’s room to determine that any contamination cannot be distinguished from the natural background radiation level with a radiation detection survey instrument set on its most sensitive scale and with no interposed shielding, or handle these materials and items as radioactive waste.

(b) Survey the patient’s room and private sanitary facility for removable contamination before assigning another patient to the room. The room shall not be reassigned until removable contamination is less than 200 disintegrations per minute (3.33 Bq) per 100 square centimeters or the wipe samples are equal to background when surveyed with an instrument using a probe with a maximum window thickness of 2.0 mg/cm and a minimum probe diameter of 1.5 inches.

(9) For manual brachytherapy patients or human research subjects who cannot be released by Rule 64E-5.622, F.A.C., the licensee shall have the applicable emergency response equipment available near each treatment room to respond to the following:

(a) A source that is dislodged from the patient or human research subject; and

(b) A sealed source lodged within the patient following removal of the source applicators.

(10) The licensee shall establish a bioassay program to measure the thyroid burden of each individual who helps prepare, prepares or administers a dosage of unsealed iodine 131 or iodine 125 in accordance with Rule 64E-5.1320, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), (3), 404.081, 404.141 FS. History–New 5-15-96, Formerly 10D-91.7321, Amended 2-11-10.

- 64E-5.6251 Manual Brachytherapy Therapy Related Computer Systems.

The licensee shall perform acceptance testing on the treatment planning system of manual brachytherapy therapy-related computer systems in accordance with published protocols accepted by nationally recognized bodies. At a minimum, the acceptance testing must include, as applicable, verification of:

(1) The source-specific input parameters required by the dose calculation algorithm;

(2) The accuracy of dose, dwell time, and treatment time calculations at representative points;

(3) The accuracy of isodose plots and graphic displays; and

(4) The accuracy of the software used to determine sealed source positions from radiographic images.

The licensee shall maintain records of this acceptance testing and protocols used in performing these tests for inspection by the department.

Rulemaking Authority 404.051, 404.061, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (6), (10), (11), 404.061(2), (3), 404.081, 404.141 FS. History–New 2-11-10, Amended 12-26-13.

SUBPART C

UPTAKE, DILUTION, AND EXCRETION

- 64E-5.626 Use of Radiopharmaceuticals for Uptake, Dilution, or Excretion Studies.

A licensee is allowed to use any unsealed radioactive material in a radiopharmaceutical for a diagnostic use involving measurements of uptake, dilution, or excretion for medical use under the following conditions:

(1) When a written directive is not required by subsection 64E-5.607(3), F.A.C., the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by a NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol, or a Notice of Claimed Investigational Exemption for a New Drug (IND) protocol accepted by U.S. Food and Drug Administration (FDA); or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application, or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. Except for sodium iodide I-131 in quantities greater than 30 microcuries (1.11 MBq), a physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, and sub-subparagraph 64E-5.650(3)(a)2.g., F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.626(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e), or subsection 64E-5.608(1), F.A.C., or;

(e) The authorized user must satisfy the training and experience specified in Rule 64E-5.649 or 64E-5.657, F.A.C.

(2) When a written directive is required by subsection 64E-5.607(3), F.A.C., the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by an NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application, or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.626(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the applicable training and experience specified in Rule 64E-5.657, 64E-5.660, 64E-5.661, 64E-5.662 or 64E-5.663, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.733, Amended 8-6-01, 2-11-10, 12-26-13.

SUBPART D

IMAGING AND LOCALIZATION

- 64E-5.627 Use of Unsealed Radiopharmaceuticals, Generators, and Reagent Kits for Imaging and Localization Studies.

A licensee is allowed to use any radioactive material in a diagnostic radiopharmaceutical, or any generator, or reagent kit, for preparation and diagnostic use of a radiopharmaceutical containing radioactive material for medical use under the following conditions:

(1) When a written directive is not required by subsection 64E-5.607(3), F.A.C., the licensee must satisfy the following:

(a) Obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent U.S. Nuclear Regulatory Commission or Agreement State regulations; or

(b) Radioactive material is obtained from and prepared by a NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. Except for sodium iodide I-131 in quantities greater than 30 microcuries (1.11 MBq), a physician who is an authorized user and meets the training requirements specified in Rules 64E-5.650 or 64E-5.660 and sub-subparagraph 64E-5.650(3)(a)2.g., F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.626(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.; or

(e) The authorized user must satisfy the training and experience specified in Rule 64E-5.650 or 64E-5.657, F.A.C.

(2) When a written directive is required by subsection 64E-5.607(3), F.A.C., the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by an NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.627(2)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the applicable training and experience specified in Rules 64E-5.657, 64E-5.660, 64E-5.661, 64E-5.662 or 64E-5.663, F.A.C.

(3) For oral administration of sodium iodide I-131 in quantities less than or equal to 33 millicuries (1.22 gigabecquerels) and when a written directive is required by subsection 64E-5.607(3), F.A.C., the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by a NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.;

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.627(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the applicable training and experience specified in Rules 64E-5.657, 64E-5.660 or 64E-5.661, F.A.C.

(4) A licensee shall use radioactive aerosols or gases only if application on DH Form 1322 12/09 is made to and approved by the department and the requirements of Rule 64E-5.629, F.A.C., are met.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.735, Amended 8-6-01, 2-11-10, 12-26-13.

- 64E-5.628 Generators.

(1) Permissible Molybdenum/Technetium Concentration.

(a) A licensee shall not administer a radiopharmaceutical containing more than 0.15 microcurie of molybdenum 99 per millicurie of technetium 99m (5.55 kilo-becquerel of molybdenum 99 per 37 megabecquerel of technetium 99m).

(b) A licensee preparing technetium 99m radiopharmaceuticals from molybdenum 99/technetium 99m generators shall measure the molybdenum 99 concentration in each eluate or extract.

(c) A licensee who is required to measure molybdenum concentrations shall retain a record of each measurement for 3 years.  The record shall include for each elution or extraction of technetium 99m:

1. The measured activity of the technetium expressed in millicuries (megabecquerels);

2. The measured activity of molybdenum expressed in microcuries (kilobecquerels);

3. The ratio of the measures expressed as microcuries of molybdenum per millicurie of technetium (kilobecquerels of molybdenum per megabecquerel of technetium);

4. The date of the test; and

5. The initials of the individual who performed the test.

(d) A licensee shall report immediately to the department each occurrence of molybdenum 99 concentration exceeding the limits specified in subsection 64E-5.628(1), F.A.C.

(2) Permissible Strontium/Rubidium Concentration.

(a) A licensee shall not administer a radiopharmaceutical containing more than 0.02 microcurie of strontium 82 per millicurie of rubidium 82 (0.74 kilobecquerel of strontium 82 per 37 megabecquerel of rubidium 82) or more than 0.2 microcurie of strontium 85 per millicurie of rubidium 82 (7.4 kilobecquerel of strontium 85 per 37 megabecquerel of rubidium 82).

(b) A licensee preparing rubidium 82 radiopharmaceuticals from strontium 82/rubidium 82 generators shall measure and calculate the strontium 82 and strontium 85 concentration on each day of use prior to the use of rubidium chloride for injection.

(c) A licensee who is required to measure strontium 82 and strontium 85 concentrations shall retain a record of each measurement for 3 years.  The record shall include for each day of use assay:

1. The measured activity of the rubidium 82 expressed in millicuries (megabecquerels);

2. The measured activity of strontium 82 expressed in microcuries (kilobecquerels);

3. The calculated activity of strontium 85 expressed in microcuries (kilobecquerels);

4. The ratio of the measures expressed as microcuries of strontium 82 per millicurie of rubidium 82 (kilobecquerels of strontium 82 per megabecquerel of rubidium 82) and the ratio of the measures expressed as microcuries of strontium 85 per millicurie of rubidium 82 (kilobecquerels of strontium 85 per megabecquerel of rubidium 82);

5. The date of the test; and

6. The initials of the individual who performed the test.

(d) A licensee shall report immediately to the department each occurrence of strontium 82 or strontium 85 concentrations exceeding the limits specified in subsection 64E-5.628(2), F.A.C.

(3) Other Permissible Parent/Daughter Concentration.

(a) If a licensee seeks to utilize a Parent/Daughter concentration other than those listed in subsection (1) or (2) above, the licensee must submit a license amendment to the department for review and approval of the maximum parent isotope or other contaminate concentrations breakthrough per daughter isotope concentration allowed for administration to patients or human research subjects, and the instrumentation and procedures used in determining parent isotope or other contaminate breakthrough concentrations;

(b) Each license must perform the determination listed in paragraph (3)(a), above, on each day of use prior to the administration to patients or human research subjects;

(c) Retain a record of each measurement for 3 years. The record shall include for each day of use assay:

1. The measured activity of the daughter isotope expressed in millicuries (megabecquerels);

2. The measured activity of parent isotope(s) and other contaminates expressed in microcuries (kilobecquerels);

3. The calculated activity of parent isotope(s) and other contaminates expressed in microcuries (kilobecquerels) as applicable;

4. The ratio of the measures expressed as microcuries of parent isotope(s) and other contaminates per millicurie of daughter isotope (kilobecquerels of parent isotope(s) per megabecquerel of daughter isotope);

5. The date of the test; and

6. The initials of the individual who performed the test.

(d) A licensee shall report immediately to the department each occurrence of parent isotope(s) or other contaminates concentrations exceeding the limits specified in paragraph 64E-5.628(3)(a), F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.736, Amended 2-11-10.

- 64E-5.629 Control of Aerosols and Gases.

(1) A licensee shall only administer radioactive aerosols or gases when airborne concentrations are within the limits prescribed by State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table I, Column 3, and Table II.

(2) The system shall either be directly vented to the atmosphere through an air exhaust or provide for collection and decay or disposal of the aerosol or gas in a shielded container.

(3) A 1icensee shall only administer radioactive gases in rooms that are at negative pressure compared to surrounding rooms.

(4) Before receiving, using, or storing radioactive gas, the licensee shall calculate the time needed after a release to reduce the concentration in the area of use to the occupational limit listed in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012. The calculation shall be based on the highest activity of gas handled in a single container and the measured available air exhaust rate.

(5) A licensee shall post the time calculated in subsection 64E-5.629(4), F.A.C., at the area of use and require that individuals evacuate the room until the posted time has elapsed if a gas spill occurs.

(6) A licensee shall check the operation of collection systems prior to use each month of use and measure the ventilation rates in areas of use every 6 months. Records of these checks and measurements shall be maintained for 3 years.

(7) A copy of the calculations required in subsection 64E-5.629(4), F.A.C., shall be recorded and retained for the duration of the license.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.737, Amended 2-11-10, 12-26-13.

SUBPART E

RADIOPHARMACEUTICALS FOR THERAPY

- 64E-5.630 Use of Radiopharmaceuticals for Therapy.

A licensee is allowed to use any unsealed radioactive material in a radiopharmaceutical that requires a written directive as described in subsection 64E-5.607(3), F.A.C., and for a therapeutic medical use provided the following is met:

(1) For any unsealed radiopharmaceutical including parenteral use listed in subsection 64E-5.630(4), F.A.C., and sodium iodide I-131 use listed in subsections 64E-5.630(2) and (3), F.A.C., the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by an NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.630(1)(d)2., F.A.C., as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the applicable training and experience specified in Rule 64E-5.660 or 64E-5.657, F.A.C.

(2) For oral administration of sodium iodide I-131 in quantities less than or equal to 33 millicuries (1.22 gigabecquerels) the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by an NRC or agreement state licensee use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.626(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the training and experience specified in Rule 64E-5.661 or 64E-5.657, F.A.C.

(3) For oral administration of sodium iodide I-131 in quantities greater than 33 millicuries (1.22 gigabecquerels) the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by an NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.626(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the training and experience specified in Rule 64E-5.662 or 64E-5.657, F.A.C.

(4) For parenteral use of radioactive materials the licensee must satisfy the following:

(a) Radioactive material is obtained from a manufacturer or pharmacy licensed as specified in subsection 64E-5.210(10), F.A.C., or in equivalent NRC or agreement state regulations; or

(b) Radioactive material is obtained from and prepared by an NRC or agreement state licensee for use in research in accordance with a Radioactive Drug Research Committee-approved protocol or an IND protocol accepted by FDA; or

(c) Radioactive material is prepared by the licensee for use in research in accordance with a Radioactive Drug Research Committee-approved application or an IND protocol accepted by FDA; or

(d) Radioactive material is prepared by:

1. An authorized nuclear pharmacist;

2. A physician who is an authorized user and meets the training requirements specified in Rule 64E-5.650 or 64E-5.660, F.A.C.; or

3. An individual under the supervision of a physician who is an authorized user under subparagraph 64E-5.626(1)(d)2., F.A.C., and as specified in paragraphs 64E-5.601(3)(a) and (b), 64E-5.607(3)(e) or subsection 64E-5.608(1), F.A.C.

(e) The authorized user must satisfy the training and experience specified in Rule 64E-5.663 or 64E-5.657, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.739, Amended 8-6-01, 2-11-10, 12-26-13.

SUBPART F

SEALED SOURCES FOR DIAGNOSIS

- 64E-5.631 Use of Sealed Sources for Diagnosis.

The licensee is allowed to use the sealed sources listed below, provided they are approved by and used as specified in, the Sealed Source and Device Registry, for diagnostic medical uses, or in research in accordance with an active IDE application accepted by the FDA and the requirements of Rule 64E-5.612, F.A.C., are met.

(1) Iodine 125 as a sealed source in a device for bone mineral analysis;

(2) Iodine 125 as a sealed source in a portable device for imaging;

(3) Gadolinium 153 as a sealed source in a device for bone mineral analysis;

(4) Americium 241 as a sealed source in a device for bone mineral analysis; or

(5) For isotopes or uses not listed in subsections 64E-5.631(1) through (4), F.A.C., above, the licensee must amend their radioactive materials license.

In order to use isotopes in accordance this Rule, an authorized user must satisfy the training and experience requirements specified in Rule 64E-5.654 or 64E-5.657, F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.743, Amended 2-11-10.

SUBPART G

SOURCES FOR BRACHYTHERAPY

- 64E-5.632 Use of Sources for Manual Brachytherapy.

The licensee is allowed to use the brachytherapy sources listed below, provided they are approved by and used as specified in, the Sealed Source and Device Registry for medical uses, or in research in accordance with an active IDE application accepted by the FDA and the requirements of Rule 64E-5.612, F.A.C., are met.

(1) Cobalt 60 as a sealed source in needles and applicator cells for topical, interstitial, and intracavitary treatment of cancer;

(2) Strontium 90 as a sealed source in an applicator for treatment of superficial eye conditions;

(3) Palladium 103 as a sealed source in seeds for interstitial treatment of cancer;

(4) Iodine 125 as a sealed source in seeds for interstitial treatment of cancer;

(5) Cesium 137 as a sealed source in needles and applicator cells for topical, interstitial, and intracavitary treatment of cancer;

(6) Iridium 192 as seeds encased in nylon ribbon for interstitial treatment of cancer;

(7) Gold 198 as a sealed source in seeds for interstitial treatment of cancer;

(8) Radon 222 as seeds for interstitial treatment of cancer;

(9) Radium 226 as a sealed source in needles or applicator cells for topical, interstitial, and intracavitary treatment of cancer;

(10) Cesium 131 as a sealed source in seeds for interstitial treatment of cancer; or

(11) For isotopes or uses not listed in subsections 64E-5.632(1) through (10), F.A.C., above, the licensee must amend their radioactive materials license.

In order to use isotopes in accordance with Rule 64E-5.632, F.A.C., an authorized user must satisfy the training and experience requirements specified in Rule 64E-5.652 or 64E-5.657, F.A.C. An authorized user of only Strontium 90 as a sealed source in an applicator for treatment of superficial eye conditions listed in subsection 64E-5.632(2), F.A.C., above must satisfy the training and experience specified in Rule 64E-5.652, 64E-5.653 or 64E-5.657, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.745, Amended 2-11-10, 12-26-13.

- 64E-5.633 Manual Brachytherapy Sources Inventory and Surveys.

(1) The licensee shall maintain accountability at all times for all manual brachytherapy sources in storage or use.

(a) As soon as possible each time brachytherapy sources are returned to an area of storage from an area of use, the licensee shall immediately count or otherwise verify the number returned to ensure that all sources taken from the storage area have been returned; and

(b) As soon as possible after removing the sources from a patient or a human research subject, the licensee shall immediately count or otherwise verify the number of sources and return them to a secure storage area.

(2) A licensee shall make a record of the use of manual brachytherapy sources which includes:

(a) For temporary implants;

1. The number and activity of sources removed from storage, the room number of use and patient’s name, the time and date the sources were removed from storage, the number and activity of sources in storage after the removal, and the name of the individual who removed the sources from storage; and

2. The number and activity of sources returned to storage, the room number of use and patient’s name, the time and date the sources were returned to storage, the number and activity of sources in storage after the return, and the name of the individual who returned the sources to storage.

(b) For permanent implants;

1. The number and activity of sources removed from storage, the room number of use and patient’s name, the time and date the sources were removed from storage, the number and activity of sources in storage after the removal, and the name of the individual who removed the sources from storage;

2. The number and activity of sources returned to storage, the room number of use and patient’s name, the time and date the sources were returned to storage, the number and activity of sources in storage after the return, and the name of the individual who returned the sources to storage; and

3. The number and activity of sources permanently implanted in the patient or human research subject.

(3) Immediately after implanting sources in a patient or human research subject and immediately after removal of sources from a patient or human research subject, the licensee shall make a radiation survey of the patient or human research subject and the area of use to confirm that no sources have been misplaced. The licensee shall make a record of each survey. This record shall contain the date and results of the survey, the survey instrument used and the name of the individual who performed the survey.

(4) A licensee shall maintain the records required in subsections 64E-5.633(2) and (3), F.A.C., for 3 years.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.748, Amended 2-11-10, 12-26-13.

- 64E-5.6331 Calibration Measurements of Manual Brachytherapy Sources.

(1) Before the first medical use of a brachytherapy source, the licensee shall, using published protocols currently accepted by nationally recognized bodies, determine the following:

(a) Source output or activity using a dosimetry system that meets the requirements of subsection 64E-5.640(1), F.A.C.; and

(b) Source positioning accuracy within applicators.

(2) Instead of a licensee making its own measurements as required in subsection 64E-5.6331(1), F.A.C., the licensee may use measurements provided by the source manufacturer or by a calibration laboratory accredited by the American Association of Physicists in Medicine (AAPM) that are made in accordance with subsection 64E-5.6331(1), F.A.C.

(3) A licensee shall mathematically correct the outputs or activities determined in subsection 64E-5.6331(1), F.A.C., for physical decay at intervals consistent with 1 percent physical decay.

(4) For each brachytherapy source the licensee shall retain the following records for three years after the last use of the source:

(a) The date of calibration;

(b) The manufacturer’s name, model number, and serial number for the source and the instruments used to calibrate the source;

(c) The source output or activity;

(d) The source positioning accuracy within the applicators; and

(e) The name of the individual, the source manufacturer, or the calibration laboratory that performed the calibration.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

- 64E-5.6332 Decay of Strontium-90 Sources for Ophthalmic Treatments.

(1) Only an authorized medical physicist or authorized user qualified to perform procedures described in subsection 64E-5.632(2), F.A.C., shall calculate the activity of each strontium-90 source that is used to determine the treatment times for ophthalmic treatments. The decay must be based on the activity determined under Rule 64E-5.6331, F.A.C.

(2) For each Strontium 90 source the licensee shall retain the following records for the life of the source:

(a) The date and activity of the source as determined under Rule 64E-5.6331, F.A.C.; and

(b) For each decay calculation, the date and the source activity as determined under Rule 64E-5.6332, F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

SUBPART H

PHOTON EMITTING REMOTE AFTERLOADER UNITS, TELETHERAPYH UNITS,

AND GAMMA STEREOTACTIC RADIOSURGERY UNITS

- 64E-5.634 Use of a Sealed Source in a Remote Afterloader Unit, Teletherapy Unit, or Gamma Stereotactic Radiosurgery Unit.

(1) A licensee shall use sealed sources in photon emitting gamma stereotactic radiosurgery units for therapeutic medical uses:

(a) As approved in the Sealed Source and Device Registry; or

(b) In research in accordance with an active IDE application accepted by the FDA provided the requirements of Rule 64E-5.612, F.A.C. are met.

(2) A licensee shall use sealed sources in photon emitting remote afterloader units for therapeutic medical uses:

(a) As approved in the Sealed Source and Device Registry; or

(b) In research in accordance with an active IDE application accepted by the FDA provided the requirements of Rule 64E-5.612, F.A.C., are met.

(3) A licensee shall use sealed sources in photon emitting teletherapy units for therapeutic medical uses:

(a) As approved in the Sealed Source and Device Registry; or

(b) In research in accordance with an active IDE application accepted by the FDA provided the requirements of Rule 64E-5.612, F.A.C., are met.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.751, Amended 2-11-10.

- 64E-5.635 Installation, Adjustment, Maintenance and Repair Restrictions.

(1) Only a person specifically licensed by the NRC or an agreement state shall install, maintain, adjust, or repair a remote afterloader unit, teletherapy unit, or gamma stereotactic radiosurgery unit that involves work on the source(s) shielding, the source(s) driving unit, or other electronic or mechanical component that could expose the source(s), reduce the shielding around the source(s), or compromise the radiation safety of the unit or the source(s).

(2) Except for low dose-rate remote afterloader units, only a person specifically licensed by the NRC or an agreement state shall install, replace, relocate, or remove a sealed source or source contained in other remote afterloader units, teletherapy units, or gamma stereotactic radiosurgery units.

(3) For a low dose-rate remote afterloader unit, only a person specifically licensed by the NRC or an agreement state or an authorized medical physicist shall install, replace, relocate, or remove a sealed source(s) contained in the unit.

(4) A licensee shall retain a record of the installation, maintenance, adjustment, and repair of remote afterloader units, teletherapy units, and gamma stereotactic radiosurgery units for 3 years. For each installation, maintenance, adjustment and repair, the record must include the date, description of the service, and name(s) of the individual(s) who performed the work.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.752, Amended 2-11-10.

- 64E-5.636 Safety Procedures and Instructions for Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic Radiosurgery Units.

(1) Listed below are the safety and instruction requirements for a licensee:

(a) Secure the unit, the console, the console keys, and the treatment room when not in use or unattended;

(b) Permit only individuals approved by the authorized user, RSO, or authorized medical physicist to be present in the treatment room during treatment with the source(s);

(c) Prevent dual operation of more than one radiation producing device in a treatment room if applicable; and

(d) Develop, implement, and maintain written procedures for responding to an abnormal situation when the operator is unable to place the source(s) in the shielded position, or remove the patient or human research subject from the radiation field with controls from outside the treatment room. These procedures must include the following:

1. Instructions for responding to equipment failures and the names of the individuals responsible for implementing corrective actions;

2. The process for restricting access to and posting of the treatment area to minimize the risk of inadvertent exposure; and

3. The names and telephone numbers of the authorized users, the authorized medical physicist, and the RSO to be contacted if the unit or console operates abnormally.

(2) A copy of the procedures required by paragraph 64E-5.636(1)(d), F.A.C., of this section must be physically located at the unit console.

(3) A licensee shall post instructions at the unit console to inform the operator of the following:

(a) The location of the procedures required by paragraph (4)(a) of this section; and

(b) The names and telephone numbers of the authorized users, the authorized medical physicist, and the RSO to be contacted if the unit or console operates abnormally.

(4) A licensee shall provide instruction, initially and at least annually, to all individuals who operate the unit, as appropriate to the individual’s assigned duties, in the following:

(a) The procedures identified in paragraph 64E-5.636(1)(d), F.A.C., of this section; and

(b) The operating procedures for the unit.

(5) A licensee shall ensure that operators, authorized medical physicists, and authorized users participate in drills of the emergency procedures, initially and at least annually.

(6) A licensee shall retain a record of individuals receiving instruction required by paragraph 64E-5.636(4), F.A.C., of this section.  These records shall be maintained for 3 years and must include the list of topics covered, the date of the instruction, the name(s) of the attendee(s) and the name(s) of the individual(s) who provided the instruction.

(7) A licensee shall retain a copy of the procedures required by paragraphs 64E-5.636(1)(d) and 64E-5.636(4)(b), F.A.C., until the licensee no longer possesses the remote afterloader, teletherapy unit or gamma stereotactic radiosurgery unit.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.753, Amended 2-11-10.

- 64E-5.637 Safety Precautions for Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic Radiosurgery Units.

(1) A licensee shall control access to the treatment room by a door at each entrance.

(2) A licensee shall equip each entrance to the treatment room with an electrical interlock system that shall:

(a) Prevent the operator from initiating the treatment cycle unless each treatment room entrance door is closed;

(b) Cause the source(s) to be shielded when an entrance door is opened; and

(c) Prevent the source(s) from being exposed following an interlock interruption until all treatment room entrance doors are closed and the source(s) on-off control is reset at the console.

(3) A licensee shall require any individual entering the treatment room to assure, through the use of appropriate radiation monitors, that radiation levels have returned to ambient levels.

(4) Except for low-dose remote afterloader units, a licensee shall construct or equip each treatment room with viewing and intercom systems to permit continuous observation of the patient or the human research subject from the treatment console during irradiation.

(5) For licensed activities where sources are placed within the patient’s or human research subject’s body, a licensee shall only conduct treatments which allow for expeditious removal of a decoupled or jammed source.

(6) In addition to the requirements specified in subsections 64E-5.637(1) through (5), F.A.C., of this section, a licensee shall:

(a) For medium dose-rate and pulsed dose-rate remote afterloader units, require:

1. An authorized medical physicist and either, an authorized user or a physician, under the supervision of an authorized user, who have been trained in the operation and emergency response for the unit, to be physically present during the initiation of all patient treatments involving the unit; and

2. An authorized medical physicist and either, an authorized user or an individual under, the supervision of an authorized user, who have been trained to remove the source applicator(s) in the event of an emergency involving the unit, to be immediately available during continuation of all patient treatments involving the unit.

(b) For high dose-rate remote afterloader units, require:

1. An authorized user and an authorized medical physicist to be physically present during the initiation of all patient treatments involving the unit; and

2. An authorized medical physicist and either, an authorized user or a physician, under the supervision of an authorized user, who have been trained in the operation and emergency response for the unit, to be physically present during continuation of all patient treatments involving the unit.

(c) For gamma stereotactic radiosurgery units, require an authorized user and an authorized medical physicist to be physically present throughout all patient treatments involving the unit.

(d) Notify the RSO, or his/her designee, and an authorized user as soon as possible if the patient or human research subject has a medical emergency or dies.

(7) A licensee shall have applicable emergency response equipment available near each treatment room in order to respond to a source remaining in the unshielded position or lodged within the patient following completion of the treatment.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.755, Amended 2-11-10.

- 64E-5.638 Radiation Monitoring Devices.

(1) A licensee shall have a permanent radiation monitor in each teletherapy, medium or high dose rate remote afterloader, or gamma stereotactic radiosurgery room capable of continuously monitoring radiation levels.

(2) Each radiation monitor shall be capable of providing visible notice of a teletherapy unit, medium or high dose rate remote afterloader unit, or gamma stereotactic radiosurgery unit malfunction that results in an exposed or partially exposed source. The visible indicator of high radiation levels shall be observable by an individual entering the teletherapy, medium or high dose rate remote afterloader, or gamma stereotactic radiosurgery room.

(3) Each radiation monitor shall be equipped with a backup power supply separate from the power supply to the teletherapy unit, medium or high dose rate remote afterloader unit, or gamma stereotactic radiosurgery unit. This backup power supply may be a battery system.

(4) Each radiation monitor shall be checked daily with a dedicated check source for proper operation before the teletherapy unit, medium or high dose rate remote afterloader unit, or gamma stereotactic radiosurgery unit is used.

(5) A licensee shall maintain a record of the check required by subsection 64E-5.638(4), F.A.C., for 3 years. The record shall include the date of the check, notation what the monitor indicates when its detector is and is not exposed to the source, and the initials of the individual who performed the check.

(6) If a radiation monitor is inoperable, the licensee shall require any individual entering the teletherapy, medium or high dose rate remote afterloader, or gamma stereotactic radiosurgery room to use a radiation survey instrument or audible alarm personal dosimeter to monitor for any malfunction of the source exposure mechanism that may result in an exposed or partially exposed source. The radiation survey instrument or dosimeter shall be checked with a dedicated check source for proper operation at the beginning of each day of use. The licensee shall keep a record as described in subsection 64E-5.638(5), F.A.C.

(7) A licensee shall promptly repair or replace the radiation monitor if it is inoperable.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.757, Amended 2-11-10.

- 64E-5.639 Viewing Systems.

A licensee shall construct or equip each teletherapy, medium or high dose rate remote afterloader, or gamma stereotactic radiosurgery room to permit continuous observation of the patient, or human research subject from the teletherapy unit console during irradiation.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.758, Amended 2-11-10.

- 64E-5.640 Dosimetry Equipment Used With Remote Afterloading Units, Teletherapy Units, or Gamma Stereotactic Radiosurgery Units.

(1) Except for low dose-rate remote afterloader source output or where the activity is determined by the manufacturer, a licensee shall have a dosimetry system available for use calibrated by paragraph (a) or (b) below.

(a) The system must have been calibrated using a system or source traceable to the NIST and published protocols accepted by nationally recognized bodies; or by a calibration laboratory accredited by the AAPM. The calibration must have been performed within the previous 2 years and after any servicing that may have affected system calibration.

(b) The system shall have been calibrated within the previous 4 years and shall have been intercompared 18 to 30 months after the calibration at an intercomparison meeting with another dosimetry system that has been calibrated within the previous 2 years by the NIST or by a calibration laboratory accredited by the AAPM. The intercomparison meeting shall be sanctioned by a calibration laboratory or radiologic physics center accredited by the AAPM. The calibration factor of the licensee’s system shall not have changed by more than 2 percent. The licensee shall not use the intercomparison result to change the calibration factor. When intercomparing dosimetry systems to be used for calibrating sealed sources for therapeutic units, the licensee shall use a comparable unit with beam attenuators or collimators, as applicable, and sources of the same radionuclide as the source used at the licensee’s facility.

(2) The licensee shall have available for use a dosimetry system for spot-check measurements. The spot-check system shall be the same system used to meet the requirement in subsection 64E-5.640(1), F.A.C., or shall be a system that has been compared with a system that has been calibrated as provided in subsection 64E-5.640(1), F.A.C. This comparison shall have been performed within the previous year and after each servicing that may have affected system calibration.

(3) The licensee shall maintain a record of each calibration, intercomparison, and comparison for the duration of the license. For each calibration, intercomparison, or comparison, the record shall include:

(a) The date, the manufacturer’s name, model numbers and serial numbers of the instruments that were calibrated, intercompared, or compared as required by subsections 64E-5.640(1) and (2), F.A.C.;

(b) The correction factors that were determined;

(c) The names of the individuals who performed the calibration, intercomparison, or comparison; and

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404,081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.759, Amended 2-11-10.

- 64E-5.641 Full Calibration Measurements on Teletherapy Units.

(1) A licensee authorized to use a teletherapy unit for medical use shall perform full calibration measurements on each teletherapy unit:

(a) Before the first medical use of the unit;

(b) Before medical use whenever spot-check measurements indicate that the output differs by more than 5 percent from the output obtained at the last full calibration corrected mathematically for radioactive decay;

(c) Before medical use following replacement of the source or following reinstallation of the teletherapy unit in a new location;

(d) Before medical use following any repair of the teletherapy unit that includes removal of the source or major repair of the components associated with the source exposure assembly; and

(e) At least every 12 months.

(2) Full calibration measurements shall include the determination of:

(a) The output within 3 percent for the range of field sizes and for the distance or range of distances used for medical use;

(b) The coincidence of the radiation field and the field indicated by the light beam localizing device;

(c) The uniformity of the radiation field and its dependence on the orientation of the useful beam;

(d) Timer constancy and linearity over the range of use;

(e) On-off error; and

(f) The accuracy of all distance measuring and localization devices in medical use.

(3) A licensee shall use the dosimetry system described in Rule 64E-5.640, F.A.C., to measure the output for one set of exposure conditions. The remaining radiation measurements required in paragraph 64E-5.6411(2)(a), F.A.C., may be made using a dosimetry system that indicates relative dose rates.

(4) A licensee shall make full calibration measurements required by subsection 64E-5.6411(1), F.A.C., using the manufacturer’s published protocols, published protocols as accepted by nationally recognized bodies or equivalent procedures that have been submitted to the department. An example of a nationally recognized body is the American Association of Physicists in Medicine.

(5) A licensee shall correct mathematically the outputs determined in paragraph 64E-5.6411(2)(a), F.A.C., for physical decay monthly for cobalt 60 and at least every 6 months for cesium 137.

(6) Full calibration measurements required by subsection 64E-5.6411(1), F.A.C., and physical decay corrections required by subsection 64E-5.6411(5), F.A.C., shall be performed by the authorized medical physicist.

(7) A licensee shall maintain a record of each calibration of each teletherapy unit for three years. The record shall include:

(a) The date of the calibration;

(b) The manufacturer’s name, model number, and serial number for both the teletherapy unit and the source;

(c) The model numbers and serial numbers of the instruments used to calibrate the teletherapy unit;

(d) The results and an assessment of the full calibration to include the following:

1. The tables that describe the output of the unit over the range of field sizes and for the range of distances used in radiation therapy;

2. A determination of the coincidence of the radiation field and the field indicated by the light beam localizing device;

3. The measured timer accuracy for a typical treatment time;

4. The calculated on-off error;

5. The estimated accuracy of each distance measuring or localization device; and

6. The signature of the authorized medical physicist.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 5-12-93, Formerly 10D-91.760, Amended 2-11-10.

- 64E-5.6411 Full Calibration Measurements on Remote Afterloader Units.

(1) A licensee authorized to use a remote afterloader unit for medical use shall perform full calibration measurements on each remote afterloader unit:

(a) Before the first medical use of the unit;

(b)1. Before medical use following replacement of the source or following reinstallation of the unit in a new location outside the facility; and

2. Before medical use following any repair of the unit that includes removal of the source or major repair of the components associated with the source exposure assembly; and

(c) At intervals not exceeding 1 quarter for high dose-rate, medium dose-rate, and pulsed dose-rate remote afterloader units with sources whose half-life exceeds 75 days; and

(d) At intervals not exceeding 1 year for low dose-rate remote afterloader units.

(2) Full calibration measurements of remote afterloader unit shall include the determination of:

(a) The output within 5 percent;

(b) Source positioning accuracy to within 1 millimeter;

(c) Source retraction with backup battery upon power failure;

(d) Timer constancy and linearity over the range of use;

(e) Length of the source transfer tubes;

(f) Length of the applicators; and

(g) Function of the source transfer tubes, applicators, and transfer tube-applicator interfaces.

(3) A licensee shall use the dosimetry system described in Rule 64E-5.640, F.A.C., to measure the output for one set of exposure conditions. The remaining radiation measurements required in paragraph 64E-5.6411(2)(a), F.A.C., may be made using a dosimetry system that indicates relative dose rates.

(4) A licensee shall make full calibration measurements required by subsection 64E-5.6411(1), F.A.C., in accordance with published protocols accepted by nationally recognized bodies.

(5) A licensee shall correct mathematically the outputs determined in paragraph 64E-5.6411(2)(a), F.A.C., for physical decay at intervals consistent with 1 percent physical decay.

(6) Full calibration measurements required by subsection 64E-5.6411(1), F.A.C., and physical decay corrections required by subsection 64E-5.6411(5), F.A.C., shall be performed by the authorized medical physicist.

(7) In addition to the requirements for full calibrations for low dose-rate remote afterloader units in subsection 64E-5.6411(2), F.A.C., a licensee shall perform an autoradiograph of the source(s) to verify inventory and source(s) arrangement at intervals not exceeding 1 quarter.

(8) For low dose-rate remote afterloader units, a licensee may use measurements provided by the source manufacturer that are made in accordance with subsections 64E-5.6411(1)-(5), F.A.C.

(9) A licensee shall maintain a record of each remote afterloader unit calibration for three years. The record shall include the following:

(a) The date of the calibration;

(b) The manufacturer’s name, model number, and serial number for both the remote afterloader unit and the source;

(c) The model numbers and serial numbers of the instruments used to calibrate the remote afterloader unit;

(d) The results and an assessment of the full calibrations.

(e) The results of the autoradiograph required for low dose-rate remote afterloaders; and

(f) The signature of the authorized medical physicist.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

- 64E-5.6412 Full Calibration Measurements on Gamma Stereotactic Radiosurgery Units.

(1) A licensee authorized to use a gamma stereotactic radiosurgery unit for medical use shall perform full calibration measurements on each gamma stereotactic radiosurgery:

(a) Before the first medical use of the unit;

(b)1. Before medical use whenever spot-check measurements indicate that the output differs by more than 5 percent from the output obtained at the last full calibration corrected mathematically for radioactive decay;

2. Before medical use following replacement of the source or following reinstallation of the gamma stereotactic radiosurgery unit in a new location;

3. Following any repair of the gamma stereotactic radiosurgery unit that includes removal of the sources or major repair of the components associated with the source assembly; and

(c) At intervals not exceeding 1 year, with the exception that relative helmet factors need only be determined before the first medical use of a helmet and following any damage to a helmet.

(2) Full calibration measurements of the gamma stereotactic radiosurgery unit shall include the determination of:

(a) The output within 3 percent;

(b) Relative helmet factors;

(c) Isocenter coincidence;

(d) Timer constancy and linearity over the range of use;

(e) On-off errors;

(f) Trunnion centricity;

(g) Treatment table retraction mechanism, using backup battery power or hydraulic backups with the unit off;

(h) Helmet microswitches;

(i) Emergency timing circuits; and

(j) Stereotactic frames and localizing devices (trunnions).

(3) A licensee shall use the dosimetry system described in Rule 64E-5.640, F.A.C., to measure the output for one set of exposure conditions.  The remaining radiation measurements required in paragraph 64E-5.6412(2)(a), F.A.C., may be made using a dosimetry system that indicates relative dose rates.

(4) A licensee shall make full calibration measurements required by subsection 64E-5.6412(1), F.A.C., in accordance with published protocols accepted by nationally recognized bodies.

(5) A licensee shall correct mathematically the outputs determined in paragraph 64E-5.6412(2)(a), F.A.C., at intervals not exceeding 1 month for cobalt-60 and at intervals consistent with 1 percent physical decay for all other radionuclides.

(6) Full calibration measurements required by subsection 64E-5.6412(1), F.A.C., and physical decay corrections required by subsection 64E-5.6412(5), F.A.C., shall be performed by the authorized medical physicist.

(7) A licensee shall maintain a record of each gamma stereotactic radiosurgery unit calibration for three years. The record shall include:

(a) The date of the calibration;

(b) The manufacturer’s name, model number, and serial number for both the gamma stereotactic radiosurgery unit and the source;

(c) The model numbers and serial numbers of the instruments used to calibrate the gamma stereotactic radiosurgery unit;

(d) The results and an assessment of the full calibrations; and

(e) The signature of the authorized medical physicist.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10, 12-26-13.

- 64E-5.642 Periodic Spot-Checks of Teletherapy Units.

(1) A licensee authorized to use teletherapy units for medical use shall perform output spot-checks on each teletherapy unit at least every month.

(2) Spot-checks shall include the determination of:

(a) Timer constancy and timer linearity over the range of use;

(b) On-off error;

(c) The coincidence of the radiation field and the field indicated by the light beam localizing device;

(d) The accuracy of all distance measuring and localization devices used for medical use;

(e) The output for one typical set of operating conditions; and

(f) The difference between the measurement made in paragraph 64E-5.642(2)(e), F.A.C., and the anticipated output, expressed as a percentage of the anticipated output, which is the value obtained at the last full calibration corrected mathematically for physical decay.

(3) A licensee shall use the dosimetry system described in Rule 64E-5.640, F.A.C., to make the spot-check required in paragraph 64E-5.642(2)(e), F.A.C.

(4) A licensee shall perform spot-checks required by subsection 64E-5.642(1), F.A.C., following procedures established by the authorized medical physicist.

(5) A licensee shall have the authorized medical physicist review the results of each output spot-check within 15 days and promptly notify the licensee in writing of the results of each output spot-check. The licensee shall keep a copy of each written notification for 3 years.

(6) A licensee authorized to use a teletherapy unit for medical use shall perform safety spot-checks of each teletherapy facility monthly and after each source installation.

(7) Safety spot-checks shall assure proper operation of:

(a) Electrical interlocks at each teletherapy room entrance;

(b) Electrical or mechanical stops installed to limit use of the primary beam of radiation, restriction of source housing angulation or elevation, carriage or stand travel, and operation of the beam on-off mechanism;

(c) Source exposure indicator lights on the teletherapy unit, on the control console, and in the facility;

(d) Viewing and intercom systems;

(e) Treatment room doors from inside and outside the treatment room; and

(f) Electrically assisted treatment room doors with the teletherapy unit electrical power turned off.

(8) If the results of the checks required in subsection 64E-5.642(7), F.A.C., of this section indicate the malfunction of any system, a licensee shall lock the control console in the off position and not use the unit.

(9) A licensee shall promptly repair any system identified in subsection 64E-5.642(7), F.A.C., that is not operating properly. The teletherapy unit shall not be used until all repairs are completed.

(10) A licensee shall maintain a record of each spot-check required by subsections 64E-5.642(1) and (6), F.A.C., for 3 years and a copy of the procedures required by subsection 64E-5.641(4), F.A.C., until the licensee no longer possesses the teletherapy unit. The record shall include:

(a) The date of the spot-check;

(b) The manufacturer’s name, model number, and serial number for both the teletherapy unit and source;

(c) The manufacturer’s name, model number and serial number of the instrument used to measure the output of the teletherapy unit;

(d) The timer linearity and constancy;

(e) The calculated on-off error;

(f) A determination of the coincidence of the radiation field and the field indicated by the light beam localizing device;

(g) The determined accuracy of each distance measuring or localization device;

(h) The difference between the anticipated output and the measured output;

(i) Notations indicating the operability of each entrance door electrical interlock, each electrical or mechanical stop, each beam condition indicator light, the viewing system and doors; and

(j) The name of the individual who performed the periodic spot-check and the signature of the authorized medical physicist who reviewed the record of the spot check.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.761, Amended 2-11-10.

- 64E-5.6421 Periodic Spot-Checks for Remote Afterloader Units.

(1) A licensee authorized to use a remote afterloader unit for medical use shall perform the following spot-checks:

(a) Before the first use of a high dose-rate, medium dose-rate, or pulsed dose-rate remote afterloader unit on a given day;

(b) Before each patient treatment with a low dose-rate remote afterloader unit; and

(c) After each source installation.

(2) Spot-checks shall include the determination of:

(a) Electrical interlocks at each remote afterloader unit room entrance;

(b) Source exposure indicator lights on the remote afterloader unit, on the control console, and in the facility;

(c) Viewing and intercom systems in each high dose-rate, medium dose-rate, and pulsed dose-rate remote afterloader facility;

(d) Emergency response equipment;

(e) Radiation monitors used to indicate the source position;

(f) Timer accuracy;

(g) Clock (date and time) in the unit’s computer; and

(h) Decayed source(s) activity in the unit’s computer.

(3) If the results of the checks required in subsection 64E-5.6421(2), F.A.C., of this section indicate the malfunction of any system, a licensee shall lock the control console in the off position and not use the unit except as may be necessary to repair, replace, or check the malfunctioning system.

(4) A licensee shall perform spot-checks required by subsection 64E-5.6421(2), F.A.C., following procedures established by the authorized medical physicist.

(5) A licensee shall have the authorized medical physicist review the results of each spot-check within 15 days and promptly notify the licensee in writing of the results of each spot-check. The licensee shall keep a copy of each written notification for 3 years.

(6) A licensee shall retain a copy of the procedures required by subsection 64E-5.6421(4), F.A.C., until the licensee no longer possesses the remote afterloader unit.

(7) A licensee shall maintain a record of each spot-check required by subsection 64E-5.6421(2), F.A.C., for 3 years and a copy of the procedures required by subsections 64E-5.6421(4) and (5), F.A.C., until the licensee no longer possesses the remote afterloader unit. The record shall include:

(a) The date of the spot-check;

(b) The manufacturer’s name, model number, and serial number for both the remote afterloader unit and source;

(c) An assessment of timer accuracy;

(d) Notations indicating the operability of each entrance door electrical interlock, radiation monitors, source exposure indicator lights, viewing and intercom systems, and clock and decayed source activity in the unit’s computer; and

(e) The name of the individual who performed the periodic spot-check and the signature of the authorized medical physicist who reviewed the record of the spot-check.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

- 64E-5.6422 Periodic Spot-Checks for Gamma Stereotactic Radiosurgery Units.

(1) A licensee authorized to use a gamma stereotactic radiosurgery unit for medical use shall perform the following spot-checks:

(a) Monthly:

(b) Before the first use of the unit on a given day; and

(c) After each source installation.

(2) To satisfy the requirements of paragraph 64E-5.6422(1)(a), F.A.C., spot checks shall include the determination of:

(a) Assure the proper operation of the:

1. Treatment table retraction mechanism, using backup battery power or hydraulic backups with the unit off;

2. Helmet microswitches;

3. Emergency timing circuits; and

4. Stereotactic frames and localizing devices (trunnions).

(b) Determine the following elements:

1. The output for one typical set of operating conditions measured with the dosimetry system described in subsection 64E-5.640(2), F.A.C.;

2. The difference between the measurement made in subparagraph 64E-5.6422(2)(b)1., F.A.C., and the anticipated output, expressed as a percentage of the anticipated output value obtained at last full calibration corrected mathematically for physical decay;

3. Source output against computer calculation;

4. Timer accuracy and linearity over the range of use;

5. On-off error; and

6. Trunnion centricity.

(3) A licensee shall perform spot-checks required by subsection 64E-5.6422(1), F.A.C., following procedures established by the authorized medical physicist.

(4) A licensee shall have the authorized medical physicist review the results of each spot-check within 15 days and promptly notify the licensee in writing of the results of each spot-check. The licensee shall keep a copy of each written notification for 3 years.

(5) To satisfy the requirements of paragraphs 64E-5.6422(1)(b) and (c), F.A.C., the licensee’s spot-checks must assure proper operation of the following:

(a) Electrical interlocks at each gamma stereotactic radiosurgery room entrance;

(b) Source exposure indicator lights on the gamma stereotactic radiosurgery unit, on the control console, and in the facility;

(c) Viewing and intercom systems;

(d) Timer termination;

(e) Radiation monitors used to indicate room exposures; and

(f) Emergency off buttons.

(6) If the results of the checks required in subsection 64E-5.6422(5), F.A.C., of this section indicate the malfunction of any system, a licensee shall lock the control console in the off position and not use the unit except as may be necessary to repair, replace, or check the malfunctioning system.

(7) A licensee shall arrange for the repair of any system identified in subsection 64E-5.6422(2), F.A.C., that is not operating properly as soon as possible.

(8) A licensee shall maintain a record of each spot-check required by subsections 64E-5.6422(2) and (5), F.A.C., for 3 years and a copy of the procedures required in subsections 64E-5.6422(2) and (3), F.A.C., until the licensee no longer possesses the gamma stereotactic radiosurgery unit. The record shall include:

(a) The date of the spot-check;

(b) The manufacturer’s name, model number, and serial number for the gamma stereotactic radiosurgery unit;

(c) The manufacturer’s name, model number and serial number of the instrument used to measure the output of the gamma stereotactic radiosurgery unit;

(d) The timer linearity and constancy;

(e) The calculated on-off error;

(f) A determination of trunnion centricity;

(g) The difference between the anticipated output and the measured output;

(h) An assessment of source output against computer calculations;

(i) Notations indicating the operability of radiation monitors, helmet microswitches, emergency timing circuits, emergency off buttons, electrical interlocks, source exposure indicator lights, viewing and intercom systems, timer termination, treatment table retraction mechanism, and stereotactic frames and localizing devices (trunnions); and

(j) The name of the individual who performed the periodic spot-check and the signature of the authorized medical physicist who reviewed the record of the spot-check.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10, Amended 12-26-13.

- 64E-5.6423 Additional Technical Requirements for Mobile Remote Afterloader Units.

(1) A licensee providing mobile remote afterloader service for medical use shall perform the following:

(a) Check survey instruments before medical use at each address of use or on each day of use, whichever is more frequent; and

(b) Account for all sources before departure from a client’s address of use.

(2) In addition to the periodic spot-checks required by Rule 64E-5.6421, F.A.C., a licensee authorized to use mobile afterloaders for medical use shall perform checks on each remote afterloader unit before use at each address of use. At a minimum, checks must be made to verify the operation of the following:

(a) Electrical interlocks on treatment area access points;

(b) Source exposure indicator lights on the remote afterloader unit, on the control console, and in the facility;

(c) Viewing and intercom systems;

(d) Applicators, source transfer tubes, and transfer tube-applicator interfaces;

(e) Radiation monitors used to indicate room exposures;

(f) Source positioning (accuracy); and

(g) Radiation monitors used to indicate whether the source has returned to a safe shielded position.

(3) In addition to the requirements for checks in subsection 64E-5.6423(2), F.A.C., a licensee shall ensure overall proper operation of the remote afterloader unit by conducting a simulated cycle of treatment before use at each address of use.

(4) If the results of the checks required in subsection 64E-5.6423(2), F.A.C., indicate the malfunction of any system, a licensee shall lock the control console in the off position and not use the unit except as may be necessary to repair, replace, or check the malfunctioning system.

(5) The licensee shall keep a copy of each check for mobile remote afterloader unit required by subsection 64E-5.6423(2), F.A.C., for three years. The records shall include:

(a) The date of the check;

(b) The manufacturer’s name, model number, and serial number of the remote afterloader unit;

(c) Notations accounting for all sources before the licensee departs from a facility;

(d) Notations indicating the operability of each entrance door electrical interlock, radiation monitors, source exposure indicator lights, viewing and intercom system, applicators, source transfer tubes, and transfer tube applicator interfaces, and source positioning accuracy; and

(e) The signature of the individual who performed the check.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

- 64E-5.643 Radiation Surveys for Teletherapy Facilities.

(1) The licensee shall perform radiation surveys with an operable radiation survey instrument calibrated as provided in Rule 64E-5.615, F.A.C., before medical use, after each installation of a teletherapy source; following repairs to the source(s) driving unit, or other electronic or mechanical component that could expose the source(s), reduce shielding around the source(s), or compromise the radiation safety of the unit or the source(s); and after making any change for which an amendment is required by Rule 64E-5.636, F.A.C.

(a) The maximum and average radiation levels from the surface of the main source(s) safe with the source(s) in the shielded position do not exceed the levels stated in the Sealed Source and Device Registry.

(b) With the teletherapy source in the on position with the largest clinically available treatment field and with a scattering phantom in the primary beam of radiation, radiation levels in restricted areas shall be unlikely to cause any occupationally exposed individuals to receive a dose in excess of the limits specified in Rule 64E-5.304, F.A.C.; and radiation dose rates of any individual member of the public in unrestricted areas shall not exceed the limits specified in paragraph 64E-5.312(1)(c), F.A.C.

(2) If the results of the surveys required in subsection 64E-5.643(1), F.A.C., indicate any radiation levels in excess of the limits specified, the licensee shall lock the control in the off position and shall not use the unit:

(a) Except to repair, replace, or test the teletherapy unit, the teletherapy unit shielding, or the treatment room shielding; or

(b) Until the licensee has received a specific exemption from the department.

(3) A licensee shall maintain a record of the radiation measurements made following installation of a source for the duration of the license. The record shall include:

(a) The date of the measurements;

(b) The reason the survey is required;

(c) The manufacturer’s name, model number and serial number of the teletherapy unit, the source, and the instrument used to measure radiation levels;

(d) Each dose rate measured around the teletherapy source while in the off position and the average of all measurements;

(e) A plan of the areas surrounding the treatment room that were surveyed;

(f) The measured dose rate at several points in each area expressed in millirems (microsieverts) per hour;

(g) The calculated maximum level of radiation over a period of 1 week for each restricted and unrestricted area; and

(h) The signature of the RSO or the authorized medical physicist.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.762, Amended 10-8-00, 2-11-10, 12-26-13.

- 64E-5.644 Radiation Surveys for Remote Afterloader and Gamma Stereotactic Radiosurgery Facilities.

(1) The licensee shall perform radiation surveys with an operable radiation survey instrument calibrated as provided in Rule 64E-5.615, F.A.C., to ensure that the maximum radiation levels and average radiation levels from the surface of the main source safe with the source(s) in the shielded position do not exceed the levels stated in the Sealed Source and Device Registry.

(2) The licensee shall make the survey specified in subsection 64E-5.644(1), F.A.C., at the installation of a new source and following repairs to the source(s) shielding, the source(s) driving unit, or other electronic or mechanical component that could expose the source, reduce the shielding around the source(s), or compromise the radiation safety of the unit or the source(s).

(3) A licensee shall retain a record of the radiation surveys required by subsection 64E-5.644(1), F.A.C., for the duration of the license. These records shall include:

(a) The date of the measurements;

(b) The manufacturer's name, model number and serial number of the treatment unit, source, and instrument used to measure radiation levels;

(c) Each dose rate measured around the source while the unit is in the off position and the average of all measurements; and

(d) The signature of the RSO or authorized medical physicist who performed the test.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.763, Amended 2-11-10.

- 64E-5.645 Remote Afterloader, Gamma Stereotactic, and Teletherapy Therapy-Related Computer Systems.

The licensee shall perform acceptance testing on the treatment planning system of high, medium, low, pulsed dose-rate remote afterloaders, gamma stereotactic, and teletherapy therapy-related computer systems in accordance with published protocols accepted by nationally recognized bodies. An example of a nationally recognized body is the American Association of Physicists in Medicine. At a minimum, the acceptance testing must include, as applicable, verification of the following:

(1) The source-specific input parameters required by the dose calculation algorithm;

(2) The accuracy of dose, dwell time, and treatment time calculations at representative points;

(3) The accuracy of isodose plots and graphic displays;

(4) The accuracy of the software used to determine sealed source positions from radiographic images; and

(5) The accuracy of electronic transfer of the treatment delivery parameters to the treatment delivery unit from the treatment planning system.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Amended 1-1-94, Formerly 10D-91.764, Amended 10-8-00, 2-11-10, 12-26-13.

- 64E-5.646 Reports of Teletherapy Surveys, Checks, Tests, and Measurements.

A licensee shall furnish a copy of the records required in Rules 64E-5.643, 64E-5.644, and 64E-5.645, F.A.C., and the output from the teletherapy source expressed as rads (grays) per hour at 1 meter from the source as determined during the full calibration required in Rule 64E-5.641, F.A.C., to the department within 30 days following completion of the action that initiated the record requirement.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–8-25-91, Formerly 10D-91.765.

- 64E-5.647 Five Year Inspection for Teletherapy and Gamma Stereotactic Radiosurgery Units.

(1) A licensee shall have each teletherapy unit and gamma stereotactic radiosurgery unit fully inspected and serviced during source replacement or at least every 5 years, whichever comes first, to assure proper functioning of the source exposure mechanism.

(2) This inspection and servicing shall only be performed by persons specifically licensed to do so by the department, an agreement state, or the U.S. Nuclear Regulatory Commission.

(3) A licensee shall maintain a record of the inspection and servicing for the duration of the license. The record shall contain:

(a) The inspector’s name;

(b) The inspector’s radioactive materials license number;

(c) The date of inspection;

(d) The manufacturer’s name and model number and serial number for both the treatment unit and source;

(e) A list of components inspected;

(f) A list of components serviced and the type of service;

(g) A list of components replaced; and

(h) The signature of the inspector.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.766, Amended 2-11-10.

SUBPART I

TRAINING AND EXPERIENCE REQUIREMENTS

- 64E-5.648 Training for a Radiation Safety Officer.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the RSO as provided in Rule 64E-5.605, F.A.C., to be an individual who:

(1) Is certified by a specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in subsections 64E-5.648(4) and (5), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/ med-use-toolkit/spec-board-cert.html.)  To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a)1. Hold a bachelor’s or graduate degree from an accredited college or university in physical science or engineering or biological science with a minimum of 20 college credits in physical science;

2. Have 5 or more years of professional experience in health physics (graduate training may be substituted for no more than 2 years of the required experience) including at least 3 years in applied health physics; and

3. Pass an examination administered by diplomates of the specialty board, which evaluates knowledge and competence in radiation physics and instrumentation, radiation protection, mathematics pertaining to the use and measurement of radioactivity, radiation biology, and radiation dosimetry; or

(b)1. Hold a master’s or doctor’s degree in physics, medical physics, other physical science, engineering, or applied mathematics from an accredited college or university;

2. Have 2 years of full-time practical training and/or supervised experience in medical physics either:

a. Under the supervision of a medical physicist who is certified in medical physics by a specialty board recognized by the NRC or an agreement state; or

b. In clinical nuclear medicine facilities providing diagnostic or therapeutic services under the direction of physicians who meet the requirements for authorized users in Rule 64E-5.657, 64E-5.650 or 64E-5.660, F.A.C.;

3. Pass an examination, administered by diplomates of the specialty board, that assesses knowledge and competence in clinical diagnostic radiological or nuclear medicine physics and in radiation safety; or

(2) Have completed a structured educational program consisting of both:

(a) 200 hours of classroom and laboratory training in the following areas:

1. Radiation physics and instrumentation;

2. Radiation protection;

3. Mathematics pertaining to the use and measurement of radioactivity;

4. Radiation biology; and

5. Radiation dosimetry.

(b) One year of full-time radiation safety experience under the supervision of the individual identified as the RSO on a NRC or agreement state license or permit issued by a NRC master material licensee that authorizes similar type(s) of use(s) of radioactive material involving the following:

1. Shipping, receiving, and performing related radiation surveys;

2. Using and performing checks for proper operation of instruments used to determine the activity of dosages, survey meters, and instruments used to measure radionuclides;

3. Securing and controlling radioactive material;

4. Using administrative controls to avoid mistakes in the administration of radioactive material;

5. Using procedures to prevent or minimize radioactive contamination and using proper decontamination procedures;

6. Using emergency procedures to control radioactive material; and

7. Disposing of radioactive material; or

(3)(a) Is a medical physicist who has been certified by a specialty board whose certification process has been recognized by the NRC or an agreement state under subsection 64E-5.656(1), F.A.C., and has experience in radiation safety for similar types of use of radioactive material for which the licensee is seeking the approval of the individual as RSO and who meets the requirements in subsections 64E-5.648(4) and (5), F.A.C., of this section; or

(b) Is an authorized user, authorized medical physicist, or authorized nuclear pharmacist identified on the licensee’s license and has experience with the radiation safety aspects of similar types of use of radioactive material for which the individual has RSO responsibilities; and

(4) Have obtained written attestation, signed by a preceptor RSO, or residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category as designated by the applicant seeking authorized status) who meets the requirements in subsection 64E-5.648(5) and in subparagraphs 64E-5.648(1)(a)1., and 64E-5.648(1)(a)2., or 64E-5.648(1)(b)1., and 64E-5.648(1)(b)2., or subsection 64E-5.648(2) or paragraph 64E-5.648(3)(a), F.A.C., of this section, and has demonstrated the ability to function independently as a RSO to fulfill the radiation safety related duties for a medical use licensee; and

(5) Has training in the radiation safety, regulatory issues, and emergency procedures for the types of use for which a licensee seeks approval.  This training requirement may be satisfied by completing training that is supervised by a RSO, authorized medical physicist, authorized nuclear pharmacist, or authorized user, as appropriate, who is authorized for the type(s) of use for which the licensee is seeking approval.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.767, Amended 2-11-10.

- 64E-5.649 Training for Uptake, Dilution, or Excretion Studies.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user of a radiopharmaceutical listed in subsection 64E-5.626(1), F.A.C., to:

(1) Be certified by a medical specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in paragraph 64E-5.649(3)(b), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/ materials/miau/med-use-toolkit/spec-board-cert.html.) To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a) Complete 60 hours of training and experience in basic radionuclide handling techniques and radiation safety applicable to the medical use of unsealed radioactive material for uptake, dilution, and excretion studies that includes the topics listed in paragraph 64E-5.649(3)(a) and subparagraph 64E-5.649(3)(a)2., F.A.C., of this section; and

(b) Pass an examination, administered by diplomates of the specialty board, that assesses knowledge and competence in radiation safety, radionuclide handling, and quality control; or

(2) Be an authorized user under Rule 64E-5.650 or 64E-5.660, F.A.C., NRC or equivalent Agreement State requirements; or

(3)(a) Have completed 60 hours of training and experience, including a minimum of 8 hours of classroom and laboratory training, in basic radionuclide handling techniques applicable to the medical use of unsealed radioactive material for uptake, dilution, and excretion studies. The training and experience must include the following:

1. Classroom and laboratory training in the following areas:

a. Radiation physics and instrumentation;

b. Radiation protection;

c. Mathematics pertaining to the use and measurement of radioactivity;

d. Chemistry of radioactive material for medical use;

e. Radiation biology; and

2. Work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657, 64E-5.649, 64E-5.650 or 64E-5.660, F.A.C., NRC or equivalent Agreement State requirements, involving the following:

a. Ordering, receiving, and unpacking radioactive materials safely and performing the related radiation surveys;

b. Performing quality control procedures on instruments used to determine the activity of dosages and performing checks for proper operation of survey meters;

c. Calculating, measuring, and safely preparing patient or human research subject dosages;

d. Using administrative controls to prevent a medical event involving the use of unsealed radioactive material;

e. Using procedures to contain spilled radioactive material safely and using proper decontamination procedures; and

f. Administering dosages of radioactive drugs to patients or human research subjects.

(b) Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657, 64E-5.649, 64E-5.650 or 64E-5.660, F.A.C., NRC or equivalent Agreement State requirements, that the individual has satisfactorily completed the requirements in paragraph 64E-5.649(1)(a) or 64E-5.649(3)(a), F.A.C., of this section and has demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for medical uses authorized under subsection 64E-5.626(1), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.769, Amended 2-11-10, 12-26-13.

- 64E-5.650 Training for Imaging and Localization Studies for Which a Written Directive Is Not Required.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user specified in subsection 64E-5.627(1), F.A.C., to:

(1) Be certified by a medical specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in paragraph 64E-5.650(3)(b), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/ materials/miau/med-use-toolkit/spec-board-cert.html.) To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a) Complete 700 hours of training and experience in basic radionuclide handling techniques and radiation safety applicable to the medical use of unsealed radioactive material for imaging and localization studies that includes the topics listed in subparagraphs 64E-5.650(3)(a)1. and 64E-5.650(3)(a)2., F.A.C., of this section; and

(b) Pass an examination, administered by diplomates of the specialty board, which assesses knowledge and competence in radiation safety, radionuclide handling, and quality control; or

(2) Be an authorized user under Rule 64E-5.660, F.A.C., and meet the requirements in sub-subparagraph 64E-5.650(3)(a)2.g., F.A.C., NRC or equivalent Agreement State requirements; or paragraph 64E-5.650(3)(a), F.A.C.; or

(3)(a) Have completed 700 hours of training and experience, including a minimum of 80 hours of classroom and laboratory training, in basic radionuclide handling techniques applicable to the medical use of unsealed radioactive material for imaging and localization studies. The training and experience must include, at a minimum the following:

1. Classroom and laboratory training in the following areas:

a. Radiation physics and instrumentation;

b. Radiation protection;

c. Mathematics pertaining to the use and measurement of radioactivity;

d. Chemistry of radioactive material for medical use;

e. Radiation biology; and

2. Work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657, 64E-5.650 64E-5.660, and sub-subparagraph 64E-5.650(3)(a)2.g., F.A.C., NRC or equivalent Agreement State requirements, involving the following:

a. Ordering, receiving, and unpacking radioactive materials safely and performing the related radiation surveys;

b. Performing quality control procedures on instruments used to determine the activity of dosages and performing checks for proper operation of survey meters;

c. Calculating, measuring, and safely preparing patient or human research subject dosages;

d. Using administrative controls to prevent a medical event involving the use of unsealed radioactive material;

e. Using procedures to safely contain spilled radioactive material and using proper decontamination procedures;

f. Administering dosages of radioactive drugs to patients or human research subjects; and

g. Eluting generator systems appropriate for preparation of radioactive drugs for imaging and localization studies, measuring and testing the eluate for radionuclide purity, and processing the eluate with reagent kits to prepare labeled radioactive drugs; and

(b) Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rules 64E-5.657, 64E-5.650, 64E-5.660, and sub-subparagraph 64E-5.650(3)(a)2.g., F.A.C., NRC or equivalent Agreement State requirements, that the individual has satisfactorily completed the requirements in paragraph 64E-5.650(1)(a) or 64E-5.650(3)(a), F.A.C., of this section and has demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for medical uses authorized under subsections 64E-5.626(1) and 64E-5.627(1), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.770, Amended 2-11-10, 12-26-13, 12-26-13.

- 64E-5.651 Training for Therapeutic Use of Radiopharmaceuticals.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.771, Repealed 2-11-10.

- 64E-5.652 Training for Use of Manual Brachytherapy Sources.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user of a brachytherapy source specified in Rule 64E-5.632, F.A.C., to:

(1) Be certified by a medical specialty board whose certification process has been recognized by the NRC or an agreement state, and who meets the requirements in paragraph 64E-5.652(2)(c), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/ miau/med-use-toolkit/spec-board-cert.html.) To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a) Successfully complete a minimum of 3 years of residency training in a radiation oncology program approved by the Residency Review Committee of the Accreditation Council for Graduate Medical Education or the Royal College of Physicians and Surgeons of Canada or the Committee on Post-Graduate Training of the American Osteopathic Association; and

(b) Pass an examination, administered by diplomates of the specialty board, that tests knowledge and competence in radiation safety, radionuclide handling, treatment planning, quality assurance, and clinical use of manual brachytherapy; or

(2)(a) Have completed a structured educational program in basic radionuclide handling techniques applicable to the use of manual brachytherapy sources that includes:

1. 200 hours of classroom and laboratory training in the following areas:

a. Radiation physics and instrumentation;

b. Radiation protection;

c. Mathematics pertaining to the use and measurement of radioactivity; and

d. Radiation biology; and

2. 500 hours of work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657 or 64E-5.652, F.A.C., NRC or equivalent Agreement State requirements at a medical institution, clinic, or private practice facility, involving the following:

a. Ordering, receiving, and unpacking radioactive materials safely and performing the related radiation surveys;

b. Checking survey meters for proper operation;

c. Preparing, implanting, and removing brachytherapy sources;

d. Maintaining running inventories of material on hand;

e. Using administrative controls to prevent a medical event involving the use of radioactive material;

f. Using emergency procedures to control radioactive material; and

(b) Have completed 3 years of supervised clinical experience in radiation oncology, under an authorized user who meets the requirements in Rule 64E-5.657 or 64E-5.652, F.A.C., NRC or equivalent Agreement State requirements, as part of a formal training program approved by the Residency Review Committee for Radiation Oncology of the Accreditation Council for Graduate Medical Education or the Royal College of Physicians and Surgeons of Canada or the Committee on Postdoctoral Training of the American Osteopathic Association. This experience may be obtained concurrently with the supervised work experience required by subparagraph 64E-5.652(2)(a)2., F.A.C., of this section; and

(c) Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657 or 64E-5.652, F.A.C., NRC or equivalent Agreement State requirements, that the individual has satisfactorily completed the requirements in paragraph 64E-5.652(1)(a) or 64E-5.652(2)(a) and 64E-5.652(2)(b), F.A.C., of this section and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for medical uses of manual brachytherapy sources authorized under Rule 64E-5.632, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.772, Amended 2-11-10, 12-26-13.

- 64E-5.653 Training for Ophthalmic Use of Strontium 90.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user of only strontium 90 for ophthalmic radiotherapy to:

(1) Be authorized user under Rule 64E-5.652, F.A.C., NRC or equivalent Agreement State requirements; or

(2)(a) Have completed 24 hours of classroom and laboratory training applicable to the medical use of strontium-90 for ophthalmic radiotherapy. The training must include the following:

1. Radiation physics and instrumentation:

2. Radiation protection;

3. Mathematics pertaining to the use and measurement of radioactivity; and

4. Radiation biology; and

(b) Have supervised clinical training in ophthalmic radiotherapy under the supervision of an authorized user at a medical institution, clinic, or private practice that includes the use of strontium-90 for the ophthalmic treatment of five individuals. This supervised clinical training must involve the following:

1. Examination of each individual to be treated;

2. Calculation of the dose to be administered;

3. Administration of the dose; and

4. Follow-up and review of each individual’s case history; and

(c) Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657, 64E-5.652 or 64E-5.653, F.A.C., NRC or equivalent Agreement State requirements, that the individual has satisfactorily completed the requirements in paragraphs 64E-5.653(2)(a) and 64E-5.653(2)(b), F.A.C., of this section and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for a medical use licensee authorized for strontium-90 for ophthalmic use.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.773, Amended 2-11-10, 12-26-13.

- 64E-5.654 Training for Use of Sealed Sources for Diagnosis.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user of a sealed source in a device specified in Rule 64E-5.631, F.A.C., to be a physician, dentist, or podiatrist:

(1) Be certified by a specialty board whose certification process includes all of the requirements in subsections 64E-5.654(2) and (3), F.A.C., of this section and whose certification has been recognized by the NRC or an agreement state.  (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.); or

(2) Have completed 8 hours of classroom and laboratory training in basic radionuclide handling techniques specifically applicable to the use of the device. The training must include the following:

(a) Radiation physics and instrumentation;

(b) Radiation protection;

(c) Mathematics pertaining to the use and measurement of radioactivity; and

(d) Radiation biology; and

(3) Have completed training in the use of the device for the uses requested.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.774, Amended 2-11-10, 12-26-13.

- 64E-5.655 Training for Use of Remote Afterloader Units, Teletherapy Units, and Gamma Stereotactic Radiosurgery Units.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user of a sealed source specified in Rule 64E-5.634, F.A.C., to:

(1) Be certified by a medical specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in paragraph 64E-5.655(2)(c) and subsection 64E-5.655(3), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.) To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a) Successfully complete a minimum of 3 years of residency training in a radiation therapy program approved by the Residency Review Committee of the Accreditation Council for Graduate Medical Education or the Royal College of Physicians and Surgeons of Canada or the Committee on Post-Graduate Training of the American Osteopathic Association; and

(b) Pass an examination, administered by diplomates of the specialty board, which tests knowledge and competence in radiation safety, radionuclide handling, treatment planning, quality assurance, and clinical use of stereotactic radiosurgery, remote afterloaders and external beam therapy; or

(2)(a) Have completed a structured educational program in basic radionuclide techniques applicable to the use of a sealed source in a therapeutic medical unit that includes the following:

1. 200 hours of classroom and laboratory training in the following areas:

a. Radiation physics and instrumentation;

b. Radiation protection;

c. Mathematics pertaining to the use and measurement of radioactivity; and

d. Radiation biology; and

2. 500 hours of work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657 or 64E-5.655, F.A.C., NRC or equivalent Agreement State requirements at a medical institution, clinic, or private practice facility, involving the following:

a. Reviewing full calibration measurements and periodic spot-checks;

b. Preparing treatment plans and calculating treatment doses and times;

c. Using administrative controls to prevent a medical event involving the use of radioactive material;

d. Implementing emergency procedures to be followed in the event of the abnormal operation of the medical unit or console;

e. Checking and using survey meters;

f. Selecting the proper dose and how it is to be administered; and

(b) Have completed 3 years of supervised clinical experience in radiation therapy, under an authorized user who meets the requirements in Rule 64E-5.657 or 64E-5.655, F.A.C., NRC or equivalent Agreement State requirements as part of a formal training program approved by the Residency Review Committee for Radiation Oncology of the Accreditation Council for Graduate Medical Education or the Royal College of Physicians and Surgeons of Canada or the Committee on Postdoctoral Training of the American Osteopathic Association. This experience may be obtained concurrently with the supervised work experience required by subparagraph 64E-5.655(2)(a)2., F.A.C., of this section; and

(c) Have obtained written attestation that the individual has satisfactorily completed the requirements in paragraph 64E-5.655(1)(a) or 64E-5.655(2)(a) and 64E-5.655(2)(b) and subsection 64E-5.655(3), F.A.C., of this section, and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for a medical use licensee for each type of therapeutic medical unit for which the individual is requesting authorized user status. The written attestation must be signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657 or 64E-5.655, F.A.C., NRC or equivalent Agreement State requirements for an authorized user for each type of therapeutic medical unit for which the individual is requesting authorized user status; and

(3) Have received training in device operation, safety procedures, and clinical use for the type(s) of use for which authorization is sought. This training requirement may be satisfied by satisfactory completion of a training program provided by the vendor for new users or by receiving training supervised by an authorized user or authorized medical physicist, as appropriate, who is authorized for the type(s) of use for which the individual is seeking authorization.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.775, Amended 2-11-10, 12-26-13.

- 64E-5.656 Training for an Authorized Medical Physicist.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized medical physicist to:

(1) Be certified by a specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in paragraph 64E-5.656(2)(b) and subsection 64E-5.656(3), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.) To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a) Hold a master’s or doctor’s degree in physics, medical physics, other physical science, engineering, or applied mathematics from an accredited college or university;

(b) Have 2 years of full-time practical training and/or supervised experience in medical physics:

1. Under the supervision of a medical physicist who is certified in medical physics by a specialty board recognized by the NRC or an agreement state; or

2. In clinical radiation facilities providing high-energy, external beam therapy (photons and electrons with energies greater than or equal to 1 million electron volts) and brachytherapy services under the direction of physicians who meet the requirements for authorized users in Rule 64E-5.657, 64E-5.652 or 64E-5.655, F.A.C.; and

(c) Pass an examination, administered by diplomates of the specialty board, that assesses knowledge and competence in clinical radiation therapy, radiation safety, calibration, quality assurance, and treatment planning for external beam therapy, brachytherapy, and stereotactic radiosurgery; or

(2)(a) Hold a master’s or doctor’s degree in physics, medical physics, other physical science, engineering, or applied mathematics from an accredited college or university; and has completed 1 year of full-time training in medical physics and an additional year of full-time work experience under the supervision of an individual who meets the requirements for an authorized medical physicist for the type(s) of use for which the individual is seeking authorization. This training and work experience must be conducted in clinical radiation facilities that provide high-energy, external beam therapy (photons and electrons with energies greater than or equal to 1 million electron volts) and brachytherapy services and must include:

1. Performing sealed source leak tests and inventories;

2. Performing decay corrections;

3. Performing full calibration and periodic spot checks of external beam treatment units, stereotactic radiosurgery units, and remote afterloading units as applicable; and

4. Conducting radiation surveys around external beam treatment units, stereotactic radiosurgery units, and remote afterloading units as applicable; and

(b) Have obtained written attestation that the individual has satisfactorily completed the requirements in subsection 64E-5.656(3) and paragraphs 64E-5.656(1)(a) and (b) or 64E-5.656(2)(a) and subsection 64E-5.656(3), F.A.C., of this section, and have demonstrated the ability to function independently as an authorized medical physicist to fulfill the radiation safety related duties for each type of therapeutic medical unit for which the individual is requesting authorized medical physicist status. The written attestation must be signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.656 or 64E-5.657, F.A.C., NRC or equivalent Agreement State requirements, for an authorized medical physicist for each type of therapeutic medical unit for which the individual is requesting authorized medical physicist status; and

(3) Have training for the type(s) of use for which authorization is sought that includes hands-on device operation, safety procedures, clinical use, and the operation of a treatment planning system. This training requirement may be satisfied by satisfactorily completing either a training program provided by the vendor or by training supervised by an authorized medical physicist authorized for the type(s) of use for which the individual is seeking authorization.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.776, Amended 2-11-10, 12-26-13.

- 64E-5.657 Training for Experienced RSO, Teletherapy or Medical Physicist, Authorized Medical Physicist, Authorized User, Nuclear Pharmacist, and Authorized Nuclear Pharmacist.

(1)(a) An individual identified as a RSO, a teletherapy or medical physicist, or a nuclear pharmacist on a NRC or agreement state license or a permit issued by a NRC or agreement state broad scope licensee or master material license permit or by a master material license permittee of broad scope, need not comply with the training requirements of Rule 64E-5.648, 64E-5.656 or 64E-5.659, F.A.C., respectively.

(b) An individual identified as a RSO, an authorized medical physicist, or an authorized nuclear pharmacist on a NRC or agreement state license or a permit issued by a NRC or agreement state broad scope licensee or master material license permit or by a master material license permittee of broad scope, need not comply with the training requirements of Rule 64E-5.648, 64E-5.656 or 64E-5.659, F.A.C., respectively.

(2) Physicians, dentists, or podiatrists identified as authorized users for the medical use of radioactive material on a license issued by the NRC or agreement state, a permit issued by a NRC master material licensee, a permit issued by a NRC or agreement state broad scope licensee, or a permit issued by a NRC master material license broad scope permittee who perform only those medical uses for which they were authorized, need not comply with the training requirements of Rule 64E-5.649, 64E-5.650, 64E-5.660, 64E-5.661, 64E-5.662, 64E-5.663, 64E-5.652, 64E-5.653, 64E-5.654 or 64E-5.655, F.A.C.

(3) Individuals who need not comply with training requirements as described in this section may serve as preceptors for, and supervisors of, applicants seeking authorization on department radioactive materials licenses for the same uses for which these individuals are authorized.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.022, 404.051(1), (4), (10), (11), 404.061(2), (3), 404.071(3), 404.141 FS. History–New 8-25-91, Amended 5-15-96, Formerly 10D-91.777, Amended 2-11-10.

- 64E-5.658 Recentness of Training.

The training and experience specified in Rules 64E-5.648, 64E-5.649, 64E-5.650, 64E-5.652, 64E-5.653, 64E-5.654, 64E-5.655, 64E-5.656, 64E-5.657, 64E-5.659, 64E-5.660, 64E-5.661, 64E-5.662 and 64E-5.663, F.A.C., shall have been obtained within the 7 years preceding the date of application or the individual shall have had related continuing education or experience since the required training and experience was completed and within the 7 years preceding the date of application.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 8-25-91, Formerly 10D-91.779, Amended 2-11-10.

- 64E-5.659 Training for an Authorized Nuclear Pharmacist.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized nuclear pharmacist to:

(1) Be certified by a specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in paragraph 64E-5.659(2)(b), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.) To have its certification process recognized, a specialty board shall require all candidates for certification to:

(a) Have graduated from a pharmacy program accredited by the American Council on Pharmaceutical Education (ACPE) or have passed the Foreign Pharmacy Graduate Examination Committee (FPGEC) examination;

(b) Hold a current, active license to practice pharmacy;

(c) Provide evidence of having acquired at least 4000 hours of training/experience in nuclear pharmacy practice. Academic training may be substituted for no more than 2000 hours of the required training and experience; and

(d) Pass an examination in nuclear pharmacy administered by diplomats of the specialty board, that assess knowledge and competency in procurement, compounding, quality assurance, dispensing, distribution, health and safety, radiation safety, provision of information and consultation, monitoring patient outcomes, research and development; or

(2)(a) Have completed 700 hours in a structured educational program consisting of both:

1. 200 hours of classroom and laboratory training in the following areas:

a. Radiation physics and instrumentation;

b. Radiation protection;

c. Mathematics pertaining to the use and measurement of radioactivity;

d. Chemistry of radioactive material for medical use; and

e. Radiation biology; and

2. Supervised practical experience in a nuclear pharmacy involving:

a. Shipping, receiving, and performing related radiation surveys;

b. Using and performing checks for proper operation of instruments used to determine the activity of dosages, survey meters, and, if appropriate, instruments used to measure alpha or beta-emitting radionuclides;

c. Calculating, assaying, and safely preparing dosages for patients or human research subjects;

d. Using administrative controls to avoid medical events in the administration of radioactive material; and

e. Using procedures to prevent or minimize radioactive contamination and using proper decontamination procedures; and

(b) Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in paragraphs 64E-5.659(1)(a), 64E-5.659(1)(b) and 64E-5.659(1)(c) or 64E-5.659(2)(a), F.A.C., of this section and have demonstrated the ability to function independently as an authorized nuclear pharmacist to fulfill the radiation safety related duties for a medical use licensee.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

- 64E-5.660 Training for Use of Unsealed Radioactive Material for Which a Written Directive Is Required in Rule 64E-5.626, 64E-5.627 or 64E-5.630, F.A.C.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require the authorized user of unsealed radioactive materials specified in Rule 64E-5.626, 64E-5.627 or 64E-5.630, F.A.C., which require a written directive to:

(1) Be certified by a medical specialty board whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in sub-subparagraphs 64E-5.660(2)(a)2.g. and paragraph 64E-5.660(2)(b), F.A.C., of this section. (Specialty boards whose certification processes have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.) To be recognized, a specialty board shall require all candidates for certification to:

(a) Successfully complete residency training in a radiation therapy or nuclear medicine training program or a program in a related medical specialty. These residency training programs must include 700 hours of training and experience as described in subparagraph 64E-5.660(2)(a)1. through sub-subparagraph 64E-5.660(2)(a)2.e., F.A.C., of this section. Eligible training programs must be approved by the Residency Review Committee of the Accreditation Council for Graduate Medical Education, the Royal College of Physicians and Surgeons of Canada, or the Committee on Post-Graduate Training of the American Osteopathic Association; and

(b) Pass an examination, administered by diplomats of the specialty board, which tests knowledge and competence in radiation safety, radionuclide handling, quality assurance, and clinical use of unsealed radioactive material for which a written directive is required; or

(2)(a) Have completed 700 hours of training and experience, including a minimum of 200 hours of classroom and laboratory training, in basic radionuclide handling techniques applicable to the medical use of unsealed radioactive material requiring a written directive. The training and experience must include the following:

1. Classroom and laboratory training in the following areas:

a. Radiation physics and instrumentation;

b. Radiation protection;

c. Mathematics pertaining to the use and measurement of radioactivity;

d. Chemistry of radioactive material for medical use; and

e. Radiation biology; and

2. Work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657 or 4E-5.660, F.A.C., NRC or equivalent Agreement State requirements. A supervising authorized user, who meets the requirements in subsection 64E-5.660(2), F.A.C., must also have experience in administering dosages in the same dosage category or categories (i.e., sub-subparagraph 64E-5.660(2)(a)2.g., F.A.C.,) as the individual requesting authorized user status. The work experience must involve the following:

a. Ordering, receiving, and unpacking radioactive materials safely and performing the related radiation surveys;

b. Performing quality control procedures on instruments used to determine the activity of dosages, and performing checks for proper operation of survey meters;

c. Calculating, measuring, and safely preparing patient or human research subject dosages;

d. Using administrative controls to prevent a medical event involving the use of unsealed radioactive material;

e. Using procedures to contain spilled radioactive material safely and using proper decontamination procedures;

f. Performing checks for proper operation of survey meters; and

g. Administering dosages of radioactive drugs to patients or human research subjects involving a minimum of three cases in each of the following categories for which the individual is requesting authorized user status as listed below:

(I) Oral administration of less than or equal to 1.22 gigabecquerels (33 millicuries) of sodium iodide I-131, for which a written directive is required or sub-sub-subparagraph 64E-5.660(2)(a)2.g.(II), F.A.C.;

(II) Oral administration of greater than 1.22 gigabecquerels (33 millicuries) of sodium iodide I-131;

(III) Parenteral administration of any beta emitter, or a photon-emitting radionuclide with a photon energy less than 150 keV, for which a written directive is required; and/or

(IV) Parenteral administration of any other radionuclide, for which a written directive is required; and

(b) Have obtained written attestation that the individual has satisfactorily completed the requirements in paragraph 64E-5.660(1)(a), sub-subparagraph 64E-5.660(2)(a)2.g. or paragraph 64E-5.660(2)(a), F.A.C., of this section, and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for a medical use licensee authorized under Rule 64E-5.626, 64E-5.627 or 64E-5.630, F.A.C., that require a written directive. Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657 or 64E-5.660, F.A.C., NRC or equivalent Agreement State requirements. The preceptor authorized user, who meets the requirements in subsection 64E-5.660(2), F.A.C., must have experience in administering dosages in the same dosage category or categories specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(II), F.A.C., as the individual requesting authorized user status.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10, Amended 12-26-13.

- 64E-5.661 Training for the Oral Administration of Sodium Iodide I-131 Requiring a Written Directive in Quantities Less Than or Equal to 1.22 Gigabecquerels (33 Millicuries).

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require an authorized user for the oral administration of sodium iodide I-131 requiring a written directive in quantities less than or equal to 1.22 Gigabecquerels (33 millicuries), to:

(1) Be certified by a medical specialty board whose certification process includes all of the requirements in paragraphs 64E-5.661(3)(a) and 64E-5.661(3)(b), F.A.C., of this section and whose certification process has been recognized by the NRC or an agreement state and who meets the requirements in paragraph 64E-5.661(3)(c), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC’s Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.); or

(2) Be an authorized user under Rule 64E-5.660, F.A.C., or uses listed in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(I) or 64E-5.660(2)(a)2.g.(II), Rule 64E-5.662, F.A.C., or equivalent agreement state requirements; or

(3)(a) Has successfully completed 80 hours of classroom and laboratory training, applicable to the medical use of sodium iodide I-131 for procedures requiring a written directive. The training must include the following:

1. Radiation physics and instrumentation;

2. Radiation protection;

3. Mathematics pertaining to the use and measurement of radioactivity;

4. Chemistry of radioactive material for medical use; and

5. Radiation biology; and

(b) Have work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657, 64E-5.660, 64E-5.661 or 64E-5.662, F.A.C., NRC or equivalent Agreement State requirements. A supervising authorized user who meets the requirements in subsection 64E-5.660(2), F.A.C., must also have experience in administering dosages as specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(I) or 64E-5.660(2)(a)2.g.(II), F.A.C. The work experience must involve the following:

1. Ordering, receiving, and unpacking radioactive materials safely and performing the related radiation surveys;

2. Performing quality control procedures on instruments used to determine the activity of dosages and performing checks for proper operation of survey meters;

3. Calculating, measuring, and safely preparing patient or human research subject dosages;

4. Using administrative controls to prevent a medical event involving the use of radioactive material;

5. Using procedures to contain spilled radioactive material safely and using proper decontamination procedures; and

6. Administering dosages to patients or human research subjects, that includes at least 3 cases involving the oral administration of less than or equal to 1.22 gigabecquerels (33 millicuries) of sodium iodide I-131; and

(c) Have obtained written attestation that the individual has satisfactorily completed the requirements in paragraphs 64E-5.661(3)(a) and 64E-5.661(3)(b), F.A.C., of this section, and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for a medical use licensee that required a written directive under Rule 64E-5.626, 64E-5.627 or 64E-5.630, F.A.C. Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657, 64E-5.660, 64E-5.661 or 64E-5.662, F.A.C., NRC or equivalent Agreement State requirements. A preceptor authorized user, who meets the requirement in subsection 64E-5.660(2), F.A.C., must also have experience in administering dosages as specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(I) or 64E-5.660(2)(a)2.g.(II), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10, Amended 12-26-13.

- 64E-5.662 Training for the Oral Administration of Sodium Iodide I-131 Requiring a Written Directive in Quantities Greater Than 1.22 Gigabecquerels (33 Millicuries).

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require an authorized user for the oral administration of sodium iodide I-131 requiring a written directive in quantities greater than 1.22 Gigabecquerels (33 millicuries), to:

(1) Be certified by a medical specialty board whose certification process includes all of the requirements in paragraphs 64E-5.662(3)(a) and 64E-5.662(3)(b), F.A.C., of this section, and whose certification has been recognized by the NRC or an agreement state, and who meets the requirements in paragraph 64E-5.662(3)(c), F.A.C., of this section. (The names of board certifications which have been recognized by the NRC or an agreement state will be posted on the NRC's Web page at http://www.nrc.gov/materials/miau/med-use-toolkit/spec-board-cert.html.); or

(2) Be an authorized user under Rule 64E-5.660, F.A.C., for uses listed in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(II), F.A.C., NRC or equivalent Agreement State requirements; or

(3)(a) Have successfully completed 80 hours of classroom and laboratory training, applicable to the medical use of sodium iodide I-131 for procedures requiring a written directive. The training must include:

1. Radiation physics and instrumentation;

2. Radiation protection;

3. Mathematics pertaining to the use and measurement of radioactivity;

4. Chemistry of radioactive material for medical use; and

5. Radiation biology; and

(b) Have work experience, under the supervision of an authorized user who meets the requirements in Rules 64E-5.657, 64E-5.660 or 64E-5.662, F.A.C., NRC or equivalent Agreement State requirements. A supervising authorized user, who meets the requirements in subsection 64E-5.660(2), F.A.C., must also have experience in administering dosages as specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(II), F.A.C. The work experience must involve the following:

1. Ordering, receiving, and unpacking radioactive materials safely and performing the related radiation surveys;

2. Performing quality control procedures on instruments used to determine the activity of dosages and performing checks for proper operation of survey meters;

3. Calculating, measuring, and safely preparing patient or human research subject dosages;

4. Using administrative controls to prevent a medical event involving the use of radioactive material;

5. Using procedures to contain spilled radioactive material safely and using proper decontamination procedures; and

6. Administering dosages to patients or human research subjects, that includes at least 3 cases involving the oral administration of greater than 1.22 gigabecquerels (33 millicuries) of sodium iodide I-131; and

(c) Have obtained written attestation that the individual has satisfactorily completed the requirements in paragraphs 64E-5.662(3)(a) and 64E-5.662(3)(b), F.A.C., of this section, and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for a medical use licensee authorized under Rule 64E-5.626, 64E-5.627 or 64E-5.630, F.A.C., that require written directives. Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657, 64E-5.660 or 64E-5.662, F.A.C., NRC or equivalent Agreement State requirements. A preceptor authorized user, who meets the requirements in subsection 64E-5.660(2), F.A.C., must also have experience in administering dosages as specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(II), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10, Amended 12-26-13.

- 64E-5.663 Training for the Parenteral Administration of Unsealed Radioactive Material Requiring a Written Directive.

Except as provided in Rule 64E-5.657, F.A.C., the licensee shall require an authorized user for the parenteral administration requiring a written directive, to:

(1) Be an authorized user under Rule 64E-5.660, F.A.C., for uses listed in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(III) or 64E-5.660(2)(a)2.g.(IV), F.A.C., NRC or equivalent Agreement State requirements; or

(2) Be an authorized user under Rule 64E-5.652 or 64E-5.655, F.A.C., NRC or equivalent Agreement State requirements and who meets the requirements in subsection 64E-5.663(4), F.A.C. of this section; or

(3) Be certified by a medical specialty board whose certification process has been recognized by the NRC or an agreement state under Rule 64E-5.652 or 64E-5.655, F.A.C., and who meets the requirements in subsection 64E-5.663(4), F.A.C., of this section.

(4)(a) Have successfully completed 80 hours of classroom and laboratory training, applicable to parenteral administrations, for which a written directive is required, of any beta emitter, or any photon-emitting radionuclide with a photon energy less than 150 keV, and/or parenteral administration of any other radionuclide for which a written directive is required. The training must include the following:

1. Radiation physics and instrumentation;

2. Radiation protection;

3. Mathematics pertaining to the use and measurement of radioactivity;

4. Chemistry of radioactive material for medical use; and

5. Radiation biology; and

(b) Have work experience, under the supervision of an authorized user who meets the requirements in Rule 64E-5.657, 64E-5.660 or 64E-5.663, F.A.C., NRC or equivalent Agreement State requirements, in the parenteral administration, for which a written directive is required, of any beta emitter, or any photon-emitting radionuclide with a photon energy less than 150 keV, and/or parenteral administration of any other radionuclide for which a written directive is required. A supervising authorized user who meets the requirements in Rule 64E-5.660, F.A.C., NRC or equivalent Agreement State requirements, must have experience in administering dosages as specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(III) or 64E-5.660(2)(a)2.g.(IV), F.A.C., NRC or equivalent Agreement State requirements. The work experience must involve the following:

1. Ordering, receiving, and unpacking radioactive materials safely, and performing the related radiation surveys;

2. Performing quality control procedures on instruments used to determine the activity of dosages, and performing checks for proper operation of survey meters;

3. Calculating, measuring, and safely preparing patient or human research subject dosages;

4. Using administrative controls to prevent a medical event involving the use of unsealed radioactive material;

5. Using procedures to contain spilled radioactive material safely, and using proper decontamination procedures; and

6. Administering dosages to patients or human research subjects, that include at least 3 cases involving the parenteral administration, for which a written directive is required, of any beta emitter, or any photon-emitting radionuclide with a photon energy less than 150 keV and/or at least 3 cases involving the parenteral administration of any other radionuclide, for which a written directive is required; and

(c) Have obtained written attestation that the individual has satisfactorily completed the requirements in subsection 64E-5.663(2) or 64E-5.663(3), F.A.C., of this section, and have demonstrated the ability to function independently as an authorized user to fulfill the radiation safety related duties for a medical use licensee authorized for the parenteral administration of unsealed radioactive material requiring a written directive. Have obtained written attestation, signed by a preceptor authorized user or a residency program director who represents a consensus of residency program faculties (as long as at least one member of the residency program faculty is an authorized individual in the same category designated by the applicant seeking authorized status) who meets the requirements in Rule 64E-5.657, 64E-5.660 or 64E-5.663, F.A.C., NRC or equivalent Agreement State requirements. A preceptor authorized user, who meets the requirements in Rule 64E-5.660, F.A.C., must have experience in administering dosages as specified in sub-sub-subparagraph 64E-5.660(2)(a)2.g.(III) or 64E-5.660(2)(a)2.g.(IV), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10, Amended 12-26-13.

SUBPART J

OTHER MEDICAL USES OF RADIOACTIVE MATERIAL OR RADIATION FROM RADIOACTIVE MATERIAL

- 64E-5.664 Other Medical Uses of Radioactive Material or Radiation From Radioactive Material.

A licensee may use radioactive materials or a radiation source from radioactive materials approved for medical use which is not specifically addressed in Rule 64E-5.626, 64E-5.627, 64E-5.630, 64E-5.631, 64E-5.632 or 64E-5.634, F.A.C., provided the following are satisfied:

(1) The applicant or licensee has received written approval from the department in a license or license amendment and uses the material in accordance with the regulations and specific license conditions the department considers necessary for the medical use of the material;

(2) The applicant or licensee has submitted the information required by Rules 64E-5.207 and 64E-5.208, F.A.C.; and

(3) The licensee shall provide specific information on the following:

(a) Radiation safety precautions and instruction;

(b) Methodology for measuring dosages or doses to be administered to patients or human research subjects;

(c) Calibration, maintenance, and repair of instruments and equipment necessary for radiation safety; and

(d) Security of radioactive materials, training or experience of individuals involved in these uses or other information not specified in paragraph 64E-5.664(3)(a), (b) or (c), F.A.C.

Rulemaking Authority 404.022, 404.051, 404.061, 404.071, 404.081, 404.141 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), (8), (9), (10), (11), 404.061(2), (3), 404.071(1), 404.081, 404.141 FS. History–New 2-11-10.

- 64E-5.701 Equipment Requirements.

(1) Safety Device. A device which prevents the entry of any portion of an individual’s body into the primary x-ray beam path or which causes the beam to be shut off upon entry into its path shall be provided on all open-beam configurations. A registrant may apply to the Department for an exemption from the requirement of a safety device. Such application shall include:

(a) A description of the various safety devices that have been evaluated;

(b) The reason each of these devices cannot be used; and

(c) A description of the alternative methods that will be employed to minimize the possibility of an accidental exposure, including procedures to assure that the operator and others in the area will be informed of the absence of safety devices.

(2) Warning Devices.

(a) Open-beam configurations shall be provided with a readily discernible indication of:

1. X-ray tube on-off status located near the radiation source housing, if the primary beam is controlled in this manner; or

2. Shutter open-closed status located near each port on the radiation source housing, if the primary beam is controlled in this manner.

(b) Warning devices shall be labeled so that their purpose is easily identified.

(c) Warning devices shall have fail-safe characteristics.

(3) Ports. Unused ports on radiation source housings shall be secured in the closed position in a manner which will prevent inadvertent opening.

(4) Labeling. All analytical x-ray equipment shall be labeled with a readily discernible sign or signs bearing the radiation symbol and the words:

(a) “CAUTION – HIGH INTENSITY X-RAY BEAM”, or words having a similar intent, on the x-ray source housing; and

(b) “CAUTION RADIATION – THIS EQUIPMENT PRODUCES RADIATION WHEN ENERGIZED”, or words having a similar intent, near any switch that energizes an x-ray tube if the radiation source is an x-ray tube.

(5) Shutters. On open-beam configurations, each port on the radiation source housing shall be equipped with a shutter that cannot be opened unless a collimator or a coupling has been connected to the port.

(6) Warning Lights. An easily visible warning light labeled with the words “X-RAY ON”, or words having a similar intent, having fail-safe characteristics, shall be located near any switch that energizes an x-ray tube and shall be illuminated only when the tube is energized.

(7) Radiation Source Housing. Each x-ray tube housing installed after the effective date of these regulations shall be so constructed that, with all shutters closed, the leakage radiation measured at a distance of 5 cm from its surface is not capable of producing a dose in excess of 2.5 mrem (0.025 mSv) in any given hour at any specified tube rating.

(8) Each x-ray generator shall be supplied with a protective cabinet which limits leakage radiation measured at a distance of 5 cm from its surface such that it is not capable of producing a dose in excess of 0.25 mrem (2.5 µSv) in any given hour.

Rulemaking Authority 404.051, 404.141, 404.22 FS. Law Implemented 404.051(1), (4), (6), 404.141, 404.22 FS. History–New 7-17-85, Formerly 10D-91.803.

- 64E-5.702 Area Requirements.

(1) Radiation Levels. The local components of an analytical x-ray system shall be located and arranged and shall include sufficient shielding or have access control such that no radiation levels exist in any area surrounding the local component group which could result in a dose to an individual present there in excess of the dose limits given in Rule 64E-5.312, F.A.C. For systems utilizing x-ray tubes, these levels shall be met at any specified tube rating.

(2) Surveys.

(a) Radiation surveys, as required by Rule 64E-5.314, F.A.C., of all analytical x-ray systems sufficient to show compliance with (1), above, shall be performed:

1. Upon installation of the equipment and at least once every 12 months thereafter;

2. Following any change in the initial arrangement, number or type of local components in the system;

3. Following any maintenance requiring the disassembly or removal of a local component of the system;

4. During the performance of maintenance and alignment procedures if the procedures require the presence of a primary x-ray beam when any local component in the system is disassembled or removed;

5. Any time a visual inspection of the local components of the system reveals an abnormal condition; and

6. Whenever personnel monitoring devices show a significant increase over the previous monitoring period or the readings are approaching the limits specified in Rule 64E-5.312, F.A.C.

(b) Radiation survey measurements shall not be required if a registrant can demonstrate compliance with (1), above, to the satisfaction of the Department.

(3) Posting. Each area or room containing analytical x-ray equipment shall be conspicuously posted with a sign or signs bearing the radiation symbol and the words “CAUTION – X-RAY EQUIPMENT”, or words having a similar intent.

Rulemaking Authority 404.051, 404.141, 404.22 FS. Law Implemented 404.051(1), (4), (6), 404.141, 404.22 FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.804.

- 64E-5.703 Operating Requirements.

(1) Procedures. Normal operating procedures shall be written and available to all analytical x-ray equipment workers. No individual shall be permitted to operate analytical x-ray equipment in any manner other than that specified in the procedures unless such individual has obtained the written approval of the radiation safety officer.

(2) Bypassing. No individual shall bypass a safety device or interlock unless such individual has obtained the written approval of the radiation safety officer. Such approval shall be for a specified period of time. When a safety device or interlock has been bypassed, a readily discernible sign bearing the words “SAFETY DEVICE NOT WORKING”, or words having a similar intent, shall be placed on the radiation source housing.

(3) Repair or Modification of X-ray Tube Systems. Except as specified in (2), above, no operation involving removal of covers, shielding materials or tube housings or modifications to shutters, collimators or beam stops shall be performed without first ascertaining that the tube is off and will remain off until safe conditions have been restored. The main switch, rather than interlocks, shall be used for routine shutdown in preparation for repairs.

Rulemaking Authority 404.051, 404.081, 404.22 FS. Law Implemented 404.051(1), (4), (6), 404.081(1), 404.22 FS. History–New 7-17-85, Formerly 10D-91.805.

- 64E-5.704 Personnel Requirements.

(1) Instruction. No person shall be permitted to operate or maintain analytical x-ray equipment unless such person has received instruction in and demonstrated competence as to:

(a) Identification of radiation hazards associated with the use of the equipment;

(b) Significance of the various radiation warning and safety devices incorporated into the equipment, or the reasons they have not been installed on certain pieces of equipment and the extra precautions required in such cases;

(c) Proper operating procedures for the equipment;

(d) Symptoms of an acute localized overexposure; and

(e) Proper procedures for reporting an actual or suspected overexposure.

(2) Personnel Monitoring.

(a) Finger or wrist dosimetric devices shall be provided to and shall be used by:

1. Analytical x-ray equipment workers using systems having an open-beam configuration and not equipped with a safety device; and

2. Personnel maintaining analytical x-ray equipment, if the maintenance procedures require the presence of a primary x-ray beam when any local component in the analytical x-ray system is disassembled or removed.

(b) Reported dose values shall not be used for the purpose of determining compliance with Rule 64E-5.304, F.A.C., unless evaluated by a qualified person, as defined in subsection 64E-5.501(61), F.A.C.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.051(1), (4), 404.22 FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.806.

- 64E-5.801 Registration Requirements.

(1) No person shall receive, possess, use, transfer, or acquire a particle accelerator facility or a particle accelerator except as authorized by a registration certificate issued by the Department pursuant to these rules.

(2) Application for registration shall be made on DH Form 1107, 10/15, “Radiation Machine Facility Registration,” (see Rule 64E-5.511, F.A.C.) and shall contain all information required by the form and accompanying instructions. Rules 64E-5.502 and 64E-5.511, F.A.C., contrain requirements concerning registration and the payment of registration fees.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (9), 404.22(1) FS. History–New 7-17-85, Amended 5-15-96, Formerly 10D-91.902, Amended 6-3-15, 3-21-16.

- 64E-5.802 General Requirements for the Issuance of a Registration Certificate for Particle Accelerators.

A registration application for acquisition and use of a particle accelerator or particle accelerator facility will be approved only if the Department determines that:

(1) The applicant is qualified by reason of training and experience to use the accelerator in question for the purpose requested in accordance with this part and Parts III and IX in such a manner as to minimize danger to public health and safety or property;

(2) The applicant’s proposed or existing equipment, facilities, and operating and emergency procedures are adequate to protect health and minimize danger to public health and safety or property;

(3) The issuance of the registration certificate will not be inimical to the health and safety of the public, and the applicant satisfies any applicable special requirement in Rule 64E-5.803, F.A.C.;

(4) The applicant has appointed a radiation safety officer;

(5) The applicant or the applicant’s staff have substantial experience in the use of particle accelerators and training sufficient to properly use the accelerator for accomplishment of the intended objectives; and

(6) The applicant has a radiation safety training program for operators of particle accelerators.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (8), (9), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.903.

- 64E-5.803 Particle Accelerators for Therapeutic Use on Humans.

In addition to the general registration requirements set forth in Rule 64E-5.802, F.A.C., accelerators used for treatment of humans will be registered only if the Department determines that:

(1) The applicant agrees to appoint a medical committee of at least two physicians, one of whom is expert in radiation therapy, plus a person experienced in depth dose calculations and radiation protection, for the purpose of evaluating and approving all proposed uses involving exposure of human beings;

(2) Persons designated on the application as the authorized users have had training and experience in treatment of humans utilizing radiations of the type and at energies near those produced by the accelerator to be employed;

(3) Individuals designated on the registration application as authorized users are physicians, as defined in Rule 64E-5.101, F.A.C.; and

(4) The applicable provisions of Rule 64E-5.508, F.A.C., are met.

Rulemaking Authority 404.051, 404.081, 404.141, 404.22 FS. Law Implemented 404.022, 404.081(1), 404.141, 404.051(1), (4), (8), (9), 404.22(1) FS. History–New 7-17-85, Amended 8-29-94, Formerly 10D-91.904.

- 64E-5.804 Limitations.

(1) The registrant shall not permit any person to act as a particle accelerator operator until such person:

(a) Has been instructed in accelerator radiation safety and has demonstrated an understanding thereof;

(b) Has received copies of and instructions in this part and the applicable requirements of Parts III and IX, pertinent registration conditions and the registrant’s operating and emergency procedures, and has demonstrated an understanding thereof; and

(c) Has demonstrated competence to use the particle accelerator, related equipment and survey instruments which will be employed in his assignment.

(2) The radiation safety officer shall have the authority to terminate the operations at a particular accelerator facility if such action is deemed necessary to protect health and minimize danger to public health and safety or property.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.906.

- 64E-5.805 Shielding and Safety Design Requirements.

(1) A radiological physicist, as defined in Rule 64E-5.501, F.A.C., shall be consulted in the design of each particle accelerator installation and shall be responsible for specification of barrier materials and thicknesses. Plans and specifications, along with assumptions and calculations on which the shielding design is based, shall be filed with the Department.

(2) Each particle accelerator installation shall be provided with such primary and secondary barriers as are necessary to assure compliance with Rules 64E-5.304 and 64E-5.312, F.A.C.

(3) A qualified person, as defined in Rule 64E-5.501, F.A.C., not necessarily the consultant involved in the design, shall be engaged to perform a radiation survey of the facility when the accelerator is first capable of producing radiation, and such survey shall include measurements of all types of radiation produced under all modes of operation at maximum operating potential.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (5), (6), 404.22(1) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.907.

- 64E-5.806 Particle Accelerator Controls and Interlock Systems.

(1) Instrumentation, readouts and controls on the particle accelerator control console shall be clearly identified and easily discernible.

(2) All entrances into a target room or other high radiation area shall be provided with interlocks that shut down the machine under conditions of barrier penetration.

(3) When an interlock system has been tripped, it shall only be possible to resume operation of the accelerator by manually resetting controls at the position where the interlock has been tripped, and lastly at the main control console. The approval of such plans shall not preclude the requirement of additional modifications should a subsequent analysis of operating conditions indicate the possibility of an individual receiving a dose in excess of the limits prescribed in Rules 64E-5.304, 64E-5.310 and 64E-5.312, F.A.C.

(4) Each safety interlock shall be on a circuit which shall allow its operation independently of all other safety interlocks.

(5) All safety interlocks shall function to the extent that any defect or component failure in an interlock system will either prevent operation of the accelerator or will produce a conspicuous audible signal or flashing warning light both at the control and at the barrier involved. Warning signals, where employed, shall not be disconnected or otherwise disabled, but shall continue to produce the warning signal during accelerator operation until the interlock has been restored to proper operation. Any failure of the warning system shall prevent operation of the accelerator.

(6) A scram button or other emergency power cutoff switch shall be located and easily identifiable in all high radiation areas. Such a cutoff switch shall include a manual reset so that the accelerator cannot be restarted from the accelerator control console without resetting the cutoff switch.

Rulemaking Authority 404.051, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.22(1) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.908.

- 64E-5.807 Warning Devices.

(1) All locations designated as high radiation areas, and all entrances to such locations shall be equipped with easily observable warning lights that operate when and only when radiation is being produced.

(2) Except in facilities designed for human exposure, each high radiation area shall have an audible warning device which shall be activated for 15 seconds prior to the possible creation of such high radiation area. Such warning device shall be clearly discernible in all high radiation areas and in any adjacent radiation areas.

(3) Barriers, temporary or otherwise, and pathways leading to high radiation areas shall be identified in accordance with Rules 64E-5.322 and 64E-5.323, F.A.C.

Rulemaking Authority 404.022, 404.051 FS. Law Implemented 404.022(1), 404.051(1), (4) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.909.

- 64E-5.808 Operating Procedures.

(1) Particle accelerators, when not in operation, shall be secured to prevent unauthorized use.

(2) The safety interlock system shall not be used to turn off the accelerator beam except in an emergency or test situation.

(3) All safety and warning devices, including interlocks, shall be checked for proper operability at intervals not to exceed 1 month. Results of such tests shall be maintained for inspection by the Department at the accelerator facility.

(4) Electrical circuit diagrams of the accelerator, and the associated interlock systems, shall be kept current and available to the operator at each accelerator facility and maintained for inspection by the Department.

(5) If, for any reason, it is necessary to intentionally bypass a safety interlock or interlocks, such action shall be:

(a) Authorized by the radiation safety committee or radiation safety officer;

(b) Recorded in a permanent log and notice posted at the accelerator control console; and

(c) Terminated as soon as possible.

(6) A copy of the current operating and the emergency procedures shall be maintained at the accelerator control panel.

Rulemaking Authority 404.051, 404.081, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), 404.081(1), 404.22(1) FS. History–New 7-17-85, Formerly 10D-91.910.

- 64E-5.809 Radiation Monitoring Requirements.

(1) At each particle accelerator facility capable of producing radioactive materials by activation, the registrant shall provide appropriate portable monitoring equipment which is operable and has been calibrated for the radiations being produced at the facility. Such equipment shall be tested for proper operation daily and calibrated at intervals not to exceed 12 months and after each servicing or repair.

(2) A radiation survey shall be performed and documented by a qualified person, as defined in subsection 64E-5.501(61), F.A.C., when changes have been made in shielding, operation, or equipment within the facility or in the occupancy of adjacent areas.

(3) Radiation levels in all high radiation areas shall be continuously monitored. The monitoring device shall be electrically independent of the accelerator control and interlock systems and capable of providing a visual or audible alarm at the entrance to high radiation areas, and other appropriate locations, so that people entering or present become aware of the existence of the hazard. A remote readout shall be located at the control panel when the production of radioactive materials by activation could cause a high radiation area.

(4) Area monitors designed and intended to display the exposure rate shall be calibrated at intervals not to exceed 1 year and after each servicing and repair.

(5) Whenever applicable, surveys shall be made to determine the amount of airborne particulate radioactivity present in areas of airborne hazards.

(6) Whenever applicable, smear surveys shall be made to determine the degree of contamination in target and other pertinent areas.

(7) All area surveys shall be made in accordance with the written procedures established by a qualified person, as defined in subsection 64E-5.501(61), F.A.C., or the radiation safety officer of the particle accelerator facility.

(8) Records of all radiation protection surveys, calibration results, instrumentation tests, and smear results shall be kept current and on file at each accelerator facility for inspection by the Department for 3 years.

Rulemaking Authority 404.051, 404.061, 404.081, 404.22 FS. Law Implemented 404.022, 404.051(1), (4), (6), 404.061(1), 404.081(1), 404.22 FS. History–New 7-17-85, Amended 4-4-89, Formerly 10D-91.911.

- 64E-5.810 Ventilation Systems.

(1) Means shall be provided to ensure that personnel entering any area where airborne radioactivity may be produced will not be exposed to concentrations in excess of the limits specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table I, Column 3.

(2) A registrant shall not vent, release or otherwise discharge airborne radioactive material to an uncontrolled area in concentrations which exceed the limits specified in State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations, June 2012, (see 64E-5.101, F.A.C.) Table II, Column 1, except as authorized pursuant to Rule 64E-5.329, F.A.C. For purposes of this paragraph, concentrations may be averaged over a period not greater than 1 year. Every reasonable effort should be made to maintain releases of radioactive material to uncontrolled areas as far below these limits as reasonably achievable.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(1), 404.081(1) FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.912, Amended 12-26-13.

- 64E-5.901 Posting of Notices to Workers.

(1) Each licensee or registrant shall post current copies of the following documents:

(a) The regulations in this part and in Part III;

(b) The license, conditions or documents incorporated into the license by reference and amendments thereto;

(c) The operating procedures applicable to activities under the license or registration;

(d) Any notice of violation involving radiological working conditions, proposed imposition of administrative penalties, or order issued pursuant to Part I, and any response from the licensee or registrant;

(e) The certificate of registration; and

(f) The emergency procedures applicable to activities conducted under the license or registration.

(2) If posting of a document specified in paragraphs (a), (b), and (c) above is not physically practical, the licensee or registrant may post a notice which describes the document and states where it may be examined. The documents specified in paragraphs (d), (e), and (f) above must be posted in their entirety.

(3) “Notice to Employees 3/01”, which is herein incorporated by reference and which is available from the department, shall be posted by each licensee or registrant as required by these regulations.

(4) Department documents posted pursuant to paragraph 64E-5.901(1)(d), F.A.C., shall be posted within 5 working days after receipt of the documents from the department; the licensee’s or registrant’s response, if any, shall be posted within 5 working days after dispatch from the licensee or registrant. Such documents shall remain posted for a minimum of 5 working days or until action correcting the violation has been completed, whichever is later.

(5) Documents, notices, or forms shall be posted to permit workers to observe them on the way to or from work, shall be conspicuous, and shall be replaced if defaced or altered.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081(1), (2) FS. History–New 7-17-85, Amended 4-4-89, 5-12-93, 5-15-96, Formerly 10D-91.1002, Amended 12-19-01.

- 64E-5.902 Instructions to Workers.

(1) All individuals who in the course of employment are likely to receive an occupational dose in excess of 100 millirem (1mSv) in a year:

(a) Shall be kept informed of the storage, transfer, or use of sources of radiation in the licensee’s or registrant’s facility;

(b) Shall be instructed in the health protection problems associated with exposure to radiation or radioactive material, in precautions or procedures to minimize exposures, and in the purposes and functions of protective devices employed;

(c) Shall be instructed in, and instructed to observe, to the extent within the worker’s control, the applicable provisions of these regulations and licenses for the protection of personnel from exposures to radiation or radioactive material;

(d) Shall be instructed of their responsibility to report promptly to the licensee or registrant any condition which may constitute, lead to, or cause a violation of the Act, these regulations, and licenses or unnecessary exposure to radiation or radioactive material;

(e) Shall be instructed in the appropriate response to warnings made in the event of any unusual occurrence or malfunction that may involve exposure to radiation or radioactive material; and

(f) Shall be advised as to the radiation exposure reports which workers shall be furnished pursuant to Rule 64E-5.903, F.A.C.

(2) In determining those individuals subject to the requirements of (1), above, licensees or registrants shall consider assigned activities during normal and abnormal situations involving exposure to sources of radiation or radioactive material that reasonably can be expected to occur during the life of the licensee’s or registrant’s facility. The extent of these instructions shall be commensurate with potential radiological health protection problems present in the workplace.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.051(1), (4), 404.061(2), 404.081 FS. History–New 7-17-85, Amended 1-1-94, Formerly 10D-91.1003, Amended 5-18-98.

- 64E-5.903 Notification and Reports to Individuals.

(1) Licensees or registrants for which personnel monitoring is required shall prepare a report as specified in this section of the radiation exposure data for each affected individual and the results of any measurements, analyses and calculations of radioactive material deposited or retained in the body of the individual. The information reported shall include data and results obtained pursuant to these regulations, orders, or license conditions, as shown in records maintained by the licensee or registrant pursuant to Part III. Each notification and report shall:

(a) Be in writing;

(b) Include appropriate identifying data such as the name of the licensee or registrant and the name of the individual;

(c) Include the individual’s exposure information; and

(d) Contain the following statement: “This report is furnished to you under the provisions of the Florida Department of Health regulation entitled Chapter 64E-5, F.A.C., Control of Radiation Hazards. You should preserve this report for future reference.”

(2) Each licensee or registrant for which personnel monitoring is required shall furnish each worker annually a written copy of the report specified in (1), above, of the worker’s exposure to radiation or radioactive material as shown in records maintained by the licensee or registrant pursuant to Part III. The licensee or registrant shall maintain records that the report was furnished for 3 years.

(3) Each licensee or registrant shall furnish to the worker upon termination of employment a written report as specified in (1), above, of the worker’s exposure to radiation received by that worker from operations of the licensee or registrant. Such report shall be furnished within 30 days from the time of termination of employment or within 30 days after the exposure of the individual has been determined by the licensee or registrant, whichever is later. The report shall cover each calendar quarter in which the worker’s activities involved exposure to sources of radiation and shall include the dates and locations of work under the license or registration in which the worker participated. The licensee or registrant shall maintain records that the report was furnished for 3 years.

(4) When a licensee or registrant is required pursuant to Part III to report to the Department any exposure of an individual to radiation or radioactive material, the licensee or the registrant shall also provide the individual a report on the exposure data included therein. Such reports shall be transmitted at a time not later than the transmittal to the Department.

(5) At the request of a worker who is terminating employment in a given calendar quarter with the licensee or registrant in work involving radiation dose, or of a worker who, while employed by another person, is terminating assignment to work involving radiation dose in the licensee’s or registrant’s facility in that calendar quarter, each licensee or registrant shall provide to each such worker, or to the worker’s designee, at termination, a written report regarding the radiation dose received by that worker from operations of the licensee or registrant during that specifically identified calendar quarter or fraction thereof, or provide a written estimate of that dose if the finally determined personnel monitoring results are not available at that time. Estimated doses shall be clearly indicated as such.

Rulemaking Authority 404.051, 404.061, 404.081 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.081 FS. History–New 7-17-85, Amended 5-12-93, Formerly 10D-91.1004.

- 64E-5.904 Presence of Representatives of Licensees or Registrants and Workers During Inspection.

(1) Each licensee or registrant shall afford to the Department at all reasonable times opportunity to inspect materials, machines, activities, facilities, premises and records pursuant to these regulations.

(2) During an inspection, Department inspectors may consult privately with workers as specified in this part. The licensee or registrant may accompany Department inspectors during other phases of an inspection.

(3) If, at the time of inspection, an individual has been authorized by the workers to represent them during Department inspections, the licensee or registrant shall notify the inspectors of such authorization and shall give the workers’ representative an opportunity to accompany the inspectors during the inspection of physical working conditions.

(4) Each workers’ representative shall be routinely engaged in work under control of the licensee or registrant and shall have received instructions as specified in this part.

(5) Different representatives of licensees or registrants and workers may accompany the inspectors during different phases of an inspection if there is no resulting interference with the conduct of the inspection; however, only one workers’ representative at a time may accompany the inspectors.

(6) With approval of the licensee or registrant or the workers’ representative, an individual who is not routinely engaged in work under control of the licensee or registrant, such as a consultant to the licensee or registrant or to the workers’ representative, shall be afforded the opportunity to accompany Department inspectors during the inspection of physical working conditions.

(7) Notwithstanding the other provisions of this section, Department inspectors are authorized to refuse to permit accompaniment by an individual who deliberately interferes with a fair and orderly inspection. With regard to areas containing information classified by an agency of the U.S. Government in the interest of national security, an individual who accompanies an inspector may have access to such information only if authorized to do so. With regard to any area containing proprietary information, the workers’ representative for that area shall be an individual previously authorized by the licensee or registrant to enter that area.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1) FS. History–New 7-17-85, Formerly 10D-91.1005.

- 64E-5.905 Consultation with Workers During Inspections.

(1) Department inspectors may consult privately with workers concerning matters of occupational radiation protection and other matters related to applicable provisions of these regulations and licenses to the extent the inspectors deem necessary for the conduct of an effective and thorough inspection.

(2) During the course of an inspection, any worker may bring privately to the attention of the inspectors, either orally or in writing, any past or present condition which the worker has reason to believe may have contributed to or caused a violation of the Act, these regulations, or license condition, or any unnecessary exposure of an individual to sources of radiation under the licensee’s or registrant’s control. Any such notice in writing shall comply with the requirements of subsection 64E-5.906(1), F.A.C.

(3) The provisions of (2), above, shall not be interpreted as authorization to disregard instructions pursuant to Rule 64E-5.902, F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1) FS. History–New 7-17-85, Formerly 10D-91.1006.

- 64E-5.906 Request by Workers for Inspections.

(1) Any worker or representative of workers believing that a violation of the Act, these regulations or license conditions exists or has occurred in work under a license or registration with regard to radiological working conditions in which the worker is engaged may request an inspection by giving notice of the alleged violation to the Department. Any such notice shall be in writing, shall set forth the specific grounds for the notice, and shall be signed by the worker or representative of the workers. A copy shall be provided to the licensee or registrant by the Department no later than at the time of inspection.

(2) If, upon receipt of such notice, the Department determines that the complaint meets the requirements set forth in this section and that there are reasonable grounds to believe that the alleged violation exists or has occurred, an inspection shall be made as soon as practicable to determine if such alleged violation exists or has occurred. Inspections pursuant to this section need not be limited to matters referred to in the complaint.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1) FS. History–New 7-17-85, Formerly 10D-91.1007.

- 64E-5.907 Inspections Not Warranted; Informal Review.

(1) If the Department determines, with respect to a complaint under this part, that an inspection is not warranted because there are no reasonable grounds to believe that a violation exists or has occurred, the Department shall notify the complainant in writing of such determination. The complainant may obtain review of such determination by submitting a written statement of position with the Department. The Department will provide the licensee or registrant with a copy of such statement by certified mail. The licensee or registrant may submit an opposing written statement of position with the Department. The Department will provide the complainant with a copy of such statement by certified mail.

(2) Upon the request of the complainant, the Department may hold an informal conference in which the complainant and the licensee or registrant may orally present their views. An informal conference may also be held at the request of the licensee or registrant, but disclosure of the identity of the complainant will be made only following receipt of written authorization from the complainant. After considering all written and oral views presented, the Department shall affirm, modify, or reverse the previous determination of the Department and furnish the complainant and the licensee or registrant a written notification of the decision and the reason therefor.

(3) If the Department determines that an inspection is not warranted because the requirements of subsection 64E-5.906(1), F.A.C., have not been met, the complainant shall be notified in writing of such determination. Such determination shall be without prejudice to the filing of a new complaint meeting the requirements of subsection 64E-5.906(1), F.A.C.

Rulemaking Authority 404.051, 404.061, 404.071 FS. Law Implemented 404.022, 404.051(1), (4), 404.061(2), 404.071(1) FS. History–New 7-17-85, Formerly 10D-91.1008.