## 25-6.055 Portable Standards.

(1) Each utility shall have one or more watthour meters to be used as portable standards, which shall have adequate capacity and voltage range to test all watthour meters used by the utility for billing purposes.

(a) All portable standard watthour meters shall be compared with a basic reference standard once a year.

(b) Each portable standard watthour meter shall be adjusted, if necessary, so that its accuracy will be within plus or minus 0.10 percent at 1.00 power factor and within plus or minus 0.20 percent at 0.50 power factor.

(2) If error exceeding that referenced in paragraph 25-6.055(1)(b), F.A.C., in the percent registration of a watthour meter used as a portable standard is observed in the comparisons in subsection 25-6.055(1), F.A.C., the utility shall investigate the source of the error. If the cause of the error cannot be corrected, use of the watthour meter as a basic reference standard shall be discontinued.

(3) The calibration history of each standard shall be made available to the Commission upon request.

Rulemaking Authority 366.05(1) FS. Law Implemented 366.05(1), (3) FS. History–New 7-29-69, Amended 5-13-85, Formerly 25-6.55, Amended 5-19-97.